## Sheet1
Method name | N (no. labs) | Mean | Bias | CV | Sample | Reference Value | Year | Source (CAP/EurA1c) | Type (Fresh/frozen) | Manufacturer | Total Error
Vitros 5.1 FS. 4600. 5600 | 183 | 6.14 | -0.16 | 2.3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Vitros | 4.348
Vitros 5.1 FS. 4600. 5600 | 183 | 7.59 | -0.01 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Vitros | 4.89
Vitros 5.1 FS. 4600. 5600 | 183 | 9.42 | 0.22 | 3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Vitros | 6.1
Vitros 5.1 FS. 4600. 5600 | 210 | 5.42 | -0.18 | 2.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Vitros | 4.328
Vitros 5.1 FS. 4600. 5600 | 210 | 9.34 | -0.06 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Vitros | 5.036
Vitros 5.1 FS. 4600. 5600 | 210 | 7.01 | -0.19 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Vitros | 4.514
Vitros 5.1 FS. 4600. 5600 | 151 | 8.33 | -0.3 | 2.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Vitros | 4.6
Vitros 5.1 FS. 4600. 5600 | 151 | 5.17 | -0.15 | 2.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Vitros | 5.338
Vitros 5.1 FS. 4600. 5600 | 209 | 6.41 | -0.17 | 2.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Vitros | 4.142
Vitros 5.1 FS. 4600. 5600 | 209 | 8.32 | -0.07 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Vitros | 4.438
Vitros 5.1 FS. 4600. 5600 | 209 | 5.51 | -0.14 | 2.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Vitros | 4.564
Vitros 5.1 FS. 4600. 5600 | 208 | 6.37 | -0.12 | 2.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Vitros | 5.172
Vitros 5.1 FS. 4600. 5600 | 208 | 6.85 | -0.12 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Vitros | 4.976
Vitros 5.1 FS. 4600. 5600 | 208 | 9.54 | -0.11 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Vitros | 5.378
Vitros 5.1 FS. 4600. 5600 | 191 | 6.46 | -0.33 | 2.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Vitros | 4.57
Vitros 5.1 FS. 4600. 5600 | 191 | 9.9 | -0.38 | 2.7 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Vitros | 4.912
Vitros 5.1 FS. 4600. 5600 | 191 | 6.47 | -0.35 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Vitros | 4.55
Vitros 5.1 FS. 4600. 5600 | 154 | 5.57 | -0.27 | 3.5 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Vitros | 6.59
Vitros 5.1 FS. 4600. 5600 | 154 | 11.31 | -0.4 | 2.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Vitros | 5.088
Vitros 5.1 FS. 4600. 5600 | 154 | 9.23 | -0.3 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Vitros | 4.992
Vitros 5.1 FS. 4600. 5600 | 154 | 4.95 | -0.09 | 3.2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Vitros | 6.182
Vitros 5.1 FS. 4600. 5600 | 154 | 7.01 | -0.37 | 3.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Vitros | 5.706
Vitros 5.1 FS. 4600. 5600 | 156-192 | 11.64 | -0.05 | 3.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Vitros | 7.202
Vitros 5.1 FS. 4600. 5600 | 156-192 | 5.82 | -0.11 | 4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Vitros | 7.73
Vitros 5.1 FS. 4600. 5600 | 156-192 | 5.14 | -0.03 | 3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Vitros | 5.85
Vitros 5.1 FS. 4600. 5600 | 156-192 | 8.02 | -0.12 | 3.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Vitros | 6.544
Vitros 5.1 FS. 4600. 5600 | 156-192 | 9.27 | -0.03 | 3.2 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Vitros | 6.242
Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.35 | 0.03 | 2.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Vitros | 4.146
Vitros 5.1 FS. 4600. 5600 | 157-187 | 9.37 | 0.21 | 3.3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Vitros | 6.678
Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.34 | 0.03 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Vitros | 4.342
Vitros 5.1 FS. 4600. 5600 | 157-187 | 12.35 | 0.32 | 4.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Vitros | 8.356
Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.89 | -0.05 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Vitros | 5.046
Vitros 5.1 FS. 4600. 5600 | 161 | 5.28 | 0.01 | 2.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Vitros | 4.126
Vitros 5.1 FS. 4600. 5600 | 161 | 10.87 | 0.28 | 3.3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Vitros | 6.748
Vitros 5.1 FS. 4600. 5600 | 161 | 6.06 | -0.14 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Vitros | 4.172
Vitros 5.1 FS. 4600. 5600 | 161 | 12.58 | 0.35 | 3.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Vitros | 6.622
Vitros 5.1 FS. 4600. 5600 | 161 | 7.43 | -0.08 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Vitros | 4.428
Vitros 5.1 FS. 4600. 5600 | 155-190 | 9.25 | 0.14 | 2.7 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Vitros | 5.432
Vitros 5.1 FS. 4600. 5600 | 155-190 | 5.94 | -0.07 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Vitros | 4.438
Vitros 5.1 FS. 4600. 5600 | 155-190 | 12.11 | 0.4 | 2.9 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Vitros | 6.084
Vitros 5.1 FS. 4600. 5600 | 155-190 | 5.04 | 0.02 | 2.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Vitros | 4.724
Vitros 5.1 FS. 4600. 5600 | 155-190 | 7.56 | -0.02 | 2.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Vitros | 5.468
Vitros 5.1 FS. 4600. 5600 | 170-191 | 6.5 | 0.09 | 2.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Vitros | 5.382
Vitros 5.1 FS. 4600. 5600 | 170-191 | 9.71 | 0.18 | 3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Vitros | 6.06
Vitros 5.1 FS. 4600. 5600 | 170-191 | 5.4 | 0.06 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Vitros | 4.176
Vitros 5.1 FS. 4600. 5600 | 170-191 | 8.67 | 0.16 | 3 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Vitros | 6.04
Vitros 5.1 FS. 4600. 5600 | 170-191 | 7.42 | 0.17 | 3.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Vitros | 6.246
Vitros 5.1 FS. 4600. 5600 | 167 | 7.47 | 0.05 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Vitros | 5.146
Vitros 5.1 FS. 4600. 5600 | 167 | 5.21 | 0.01 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Vitros | 3.93
Vitros 5.1 FS. 4600. 5600 | 167 | 8.34 | 0.03 | 2.8 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Vitros | 5.518
Vitros 5.1 FS. 4600. 5600 | 167 | 10.51 | 0.18 | 2.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Vitros | 5.864
Vitros 5.1 FS. 4600. 5600 | 167 | 5.92 | 0.05 | 2.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Vitros | 4.754
Vitros 5.1 FS. 4600. 5600 | 194 | 7.28 | 0.03 | 2.6 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Vitros | 5.126
Vitros 5.1 FS. 4600. 5600 | 194 | 6.83 | 0.02 | 2.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Vitros | 4.724
Vitros 5.1 FS. 4600. 5600 | 194 | 8.66 | 0.01 | 2.7 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Vitros | 5.302
Vitros 5.1 FS. 4600. 5600 | 194 | 9.6 | 0.1 | 3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Vitros | 5.98
Vitros 5.1 FS. 4600. 5600 | 194 | 5.42 | -0.03 | 2.1 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Vitros | 4.086
Vitros 5.1 FS. 4600. 5600 | 190 | 7.16 | 0.01 | 3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Vitros | 5.89
Vitros 5.1 FS. 4600. 5600 | 190 | 5.25 | 0.06 | 2.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Vitros | 4.764
Vitros 5.1 FS. 4600. 5600 | 190 | 8.4 | -0.02 | 2.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Vitros | 5.468
Vitros 5.1 FS. 4600. 5600 | 190 | 9.71 | -0.08 | 3.1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Vitros | 5.996
Vitros 5.1 FS. 4600. 5600 | 190 | 6.1 | -0.02 | 2.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Vitros | 5.272
Vitros 5.1 FS. 4600. 5600 | 171 | 6.28 | -0.03 | 2.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Vitros | 4.87
Vitros 5.1 FS. 4600. 5600 | 171 | 9.06 | -0.05 | 3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Vitros | 5.83
Vitros 5.1 FS. 4600. 5600 | 171 | 5.27 | -0.04 | 2.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Vitros | 4.272
Vitros 5.1 FS. 4600. 5600 | 171 | 8.4 | 0.05 | 2.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Vitros | 5.342
Vitros 5.1 FS. 4600. 5600 | 171 | 7.6 | -0.02 | 2.3 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Vitros | 4.488
Vitros 5.1 FS. 4600. 5600 | 176 | 6.92 | -0.19 | 2.4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Vitros | 4.514
Vitros 5.1 FS. 4600. 5600 | 176 | 9.22 | -0.1 | 2.3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Vitros | 4.408
Vitros 5.1 FS. 4600. 5600 | 176 | 5.83 | -0.24 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Vitros | 4.856
Vitros 5.1 FS. 4600. 5600 | 193 | 7.69 | -0.36 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Vitros | 4.344
Vitros 5.1 FS. 4600. 5600 | 193 | 5.08 | -0.22 | 3.3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Vitros | 6.248
Vitros 5.1 FS. 4600. 5600 | 193 | 5.81 | -0.33 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Vitros | 5.746
Roche Diagnostics Cobas c513 | 10 | 8.921 | 3.9 | 2.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Roche Diagnostics | 9.192
Roche Diagnostics Cobas c513 | 10 | 5.8649 | 2.4 | 4.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Roche Diagnostics | 11.416
Abbott Alinity | 17 | 5.261 | -2.5 | 8.2 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 13.572
Abbott Alinity | 15 | 5.2793 | -2.3 | 8.2 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 13.772
Abbott Alinity | 12 | 5.4531 | -0.4 | 2.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 3.716
Abbott Alinity | 12 | 5.4531 | -0.4 | 1.9 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 3.324
Abbott Alinity ci series | 18 | 5.13 | -0.12 | 1.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Abbott | 3.212
Abbott Alinity ci series | 18 | 6.34 | -0.07 | 1.5 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Abbott | 2.87
Abbott Alinity ci series | 18 | 8.32 | 0.11 | 0.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Abbott | 1.678
Abbott Alinity ci series | 18 | 5.41 | -0.06 | 1.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Abbott | 2.684
Abbott Alinity ci series | 18 | 9.62 | 0.14 | 0.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Abbott | 1.904
Abbott Alinity ci series | 19 | 8.03 | 0.06 | 1.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Abbott | 2.412
Abbott Alinity ci series | 19 | 5.84 | -0.05 | 1.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Abbott | 2.302
Abbott Alinity ci series | 19 | 5.05 | -0.08 | 1.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Abbott | 2.664
Abbott Alinity ci series | 19 | 7.39 | -0.01 | 1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Abbott | 1.95
Abbott Alinity ci series | 19 | 9.28 | 0.11 | 1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Abbott | 2.07
Abbott Alinity ci series | 36 | 7.79 | -0.02 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Abbott | 2.724
Abbott Alinity ci series | 36 | 5.78 | -0.13 | 1.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Abbott | 3.006
Abbott Alinity ci series | 36 | 8.63 | -0.06 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Abbott | 3.272
Abbott Alinity ci series | 36 | 5.76 | -0.1 | 1.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Abbott | 2.644
Abbott Alinity ci series | 36 | 9.53 | -0.03 | 1.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Abbott | 3.106
Abbott Alinity ci series | 53 | 5.34 | -0.14 | 3.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Abbott | 5.936
Abbott Alinity ci series | 53 | 8.28 | 0.02 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Abbott | 4.332
Abbott Alinity ci series | 53 | 5.01 | -0.14 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Abbott | 6.328
Abbott Alinity ci series | 53 | 10.17 | 0.12 | 1 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Abbott | 2.08
Abbott Alinity ci series | 53 | 5.87 | -0.12 | 2.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Abbott | 4.192
Abbott Alinity ci series | 59 | 8.93 | 0.07 | 1.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Abbott | 2.422
Abbott Alinity ci series | 59 | 6.77 | -0.07 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Abbott | 3.654
Abbott Alinity ci series | 59 | 9.49 | 0.1 | 1.3 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Abbott | 2.648
Abbott Alinity ci series | 59 | 6.02 | -0.11 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Abbott | 3.418
Abbott Alinity ci series | 59 | 7.51 | -0.01 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Abbott | 3.126
Abbott Alinity ci series | 67 | 5.62 | -0.18 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Abbott | 3.152
Abbott Alinity ci series | 67 | 9.21 | 0.09 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Abbott | 3.226
Abbott Alinity ci series | 67 | 5.35 | -0.12 | 1.4 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Abbott | 2.624
Abbott Alinity ci series | 67 | 7.29 | 0.04 | 1.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Abbott | 2.784
Abbott Alinity ci series | 67 | 9.32 | 0.06 | 1.2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Abbott | 2.412
Abbott Alinity ci series | 87 | 5.1 | -0.11 | 2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Abbott | 3.81
Abbott Alinity ci series | 87 | 6.59 | -0.09 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Abbott | 3.046
Abbott Alinity ci series | 87 | 8.27 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Abbott | 3.528
Abbott Alinity ci series | 87 | 5.92 | -0.06 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Abbott | 3.664
Abbott Alinity ci series | 87 | 10.4 | 0.28 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Abbott | 3.612
Abbott Architect | 24 | 10.9 |  | 6.4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Abbott | 12.544
Abbott Architect | 24 | 4.8 |  | 8.8 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Abbott | 17.248
Abbott Architect | 24 | 8.3 |  | 5.3 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Abbott | 10.388
Abbott Architect | 28 | 6.5 |  | 7.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Abbott | 14.504
Abbott Architect | 28 | 6.6 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Abbott | 15.092
Abbott Architect | 28 | 11.6 |  | 5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Abbott | 9.8
Abbott Architect | 28 | 6.5 |  | 7.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Abbott | 14.504
Abbott Architect | 28 | 6.6 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Abbott | 15.092
Abbott Architect | 28 | 11.6 |  | 5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Abbott | 9.8
Abbott Architect | 47 | 5 |  | 5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Abbott | 9.8
Abbott Architect | 47 | 11.2 |  | 3.9 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Abbott | 7.644
Abbott Architect | 47 | 7.4 |  | 4.1 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Abbott | 8.036
Abbott Architect | 46 | 8.9 |  | 5.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Abbott | 11.172
Abbott Architect | 46 | 10.9 |  | 4.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Abbott | 8.036
Abbott Architect | 46 | 6 |  | 4.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Abbott | 8.232
Abbott Architect | 50 | 6.73 | 0.13 | 4.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Abbott | 8.95
Abbott Architect | 50 | 9.62 | 0.12 | 4.2 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Abbott | 8.352
Abbott Architect | 50 | 7.7 | 0.3 | 5.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Abbott | 10.884
Abbott Architect | 52 | 5.94 | 0.04 | 5 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Abbott | 9.84
Abbott Architect | 52 | 9.98 | 0.18 | 3.9 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Abbott | 7.824
Abbott Architect | 52 | 7.56 | 0.16 | 5.2 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Abbott | 10.352
Abbott Architect | 53 | 5.02 | -0.18 | 5.9 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Abbott | 11.384
Abbott Architect | 53 | 8.74 | 0.04 | 3.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Abbott | 6.9
Abbott Architect | 53 | 6.22 | -0.08 | 4.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Abbott | 8.152
Abbott Architect | 59 | 8.73 | 0.23 | 3.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Abbott | 7.09
Abbott Architect | 59 | 5.38 | -0.02 | 3.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Abbott | 6.448
Abbott Architect | 59 | 6.56 | 0.16 | 3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Abbott | 6.04
Abbott Architect | 33 | 5.9564 | -2.6 | 6.8 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Abbott | 10.728
Abbott Architect | 34 | 7.0727 | -4.2 | 6.6 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Abbott | 8.736
Abbott Architect | 35 | 6.2035 | 0.1 | 3.9 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Abbott | 7.744
Abbott Architect | 35 | 7.4936 | 0.4 | 3.1 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Abbott | 6.476
Abbott Architect | 42 | 7.1459 | -3.5 | 6.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Abbott | 8.456
Abbott Architect | 42 | 5.6819 | -3.6 | 8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Abbott | 12.08
Abbott Architect | 75 | 7.3746 | -1 | 3.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Abbott | 5.076
Abbott Architect | 75 | 5.9381 | -0.8 | 3 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Abbott | 5.08
Abbott Architect (enzymatic) | 39 | 8.4909 | -0.8 | 4 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Abbott | 7.04
Abbott Architect (enzymatic) | 39 | 5.508 | -1.5 | 6.2 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Abbott | 10.652
Abbott Architect (enzymatic) | 24 | 8.7288 | 1.8 | 1.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Abbott | 5.524
Abbott Architect (enzymatic) | 24 | 5.6361 | -0.1 | 3.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Abbott | 5.976
Abbott Architect (enzymatic) | 28 | 5.2702 | -2.4 | 9.3 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 15.828
Abbott Architect (enzymatic) | 22 | 5.3434 | -1.6 | 8 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 14.08
Abbott Architect (enzymatic) | 6 | 5.5172 | 0.3 | 7.4 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 14.804
Abbott Architect (enzymatic) | 6 | 5.508 | 0.2 | 7.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 14.116
Abbott Architect (enzymatic) | 24 | 5.6727 | -3.8 | 5.2 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Abbott | 6.392
Abbott Architect (enzymatic) | 24 | 7.0544 | -4.3 | 6.8 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Abbott | 9.028
Abbott Architect (enzymatic) | 21 | 5.9656 | -0.6 | 1.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Abbott | 2.34
Abbott Architect (enzymatic) | 21 | 7.4844 | 0.4 | 1.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Abbott | 3.732
Abbott Architect (immunoassay) | 31 | 5.7551 | 2.9 | 9.8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 22.108
Abbott Architect (immunoassay) | 31 | 5.7185 | 2.5 | 8.8 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 19.748
Abbott Architect (immunoassay) | 21 | 5.6361 | 1.6 | 8.2 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 17.672
Abbott Architect (immunoassay) | 21 | 5.6361 | 1.6 | 8.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 17.476
Abbott Architect (immunoassay) | 6 | 5.929 | -1 | 2 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Abbott | 2.92
Abbott Architect (immunoassay) | 6 | 7.21 | -2.6 | 6.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Abbott | 9.356
Abbott Architect c System | 66 | 6.14 | -0.16 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Abbott | 7.484
Abbott Architect c System | 66 | 7.45 | -0.15 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Abbott | 7.69
Abbott Architect c System | 66 | 8.97 | -0.23 | 5.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Abbott | 10.158
Abbott Architect c System | 74 | 5.46 | -0.14 | 3.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Abbott | 6.916
Abbott Architect c System | 74 | 9.36 | -0.04 | 4.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Abbott | 8.388
Abbott Architect c System | 74 | 7.1 | -0.1 | 3.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Abbott | 5.976
Abbott Architect c System | 83 | 5.4 | 0 | 4.4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Abbott | 8.624
Abbott Architect c System | 83 | 8.28 | -0.02 | 3.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Abbott | 7.428
Abbott Architect c System | 83 | 6.8 | 1.15 | 3.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Abbott | 7.814
Abbott Architect c System | 47 | 6.76 | 0.18 | 3.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 7.628
Abbott Architect c System | 47 | 8.68 | 0.29 | 4.4 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 8.914
Abbott Architect c System | 47 | 5.7 | 0.05 | 4.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 8.086
Abbott Architect c System | 58 | 6.69 | 0.2 | 4.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 9.412
Abbott Architect c System | 58 | 7.2 | 0.23 | 3.7 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 7.482
Abbott Architect c System | 58 | 9.96 | 0.31 | 3.9 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 7.954
Abbott Architect c System | 42 | 10.35 | 0.07 | 3.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 7.518
Abbott Architect c System | 42 | 6.92 | 0.1 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 5.784
Abbott Architect c System | 26 | 8.74 | 0.11 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 5.402
Abbott Architect c System | 26 | 5.41 | 0.09 | 3 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 5.97
Abbott Architect c System | 24 | 5.93 | 0.09 | 2.6 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 5.186
Abbott Architect c System | 24 | 12.27 | 0.56 | 3.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 8.008
Abbott Architect c System | 24 | 9.77 | 0.24 | 4.3 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 8.668
Abbott Architect c System | 24 | 5.1 | 0.06 | 3.9 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 7.704
Abbott Architect c System | 24 | 7.49 | 0.11 | 2.8 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 5.598
Abbott Architect c System | 21-32 | 12.2 | 0.51 | 3 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 6.39
Abbott Architect c System | 21-32 | 6.22 | 0.29 | 3.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 6.758
Abbott Architect c System | 21-32 | 5.45 | 0.28 | 3.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 6.748
Abbott Architect c System | 21-32 | 8.43 | 0.29 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 3.426
Abbott Architect c System | 21-32 | 9.7 | 0.4 | 2.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 5.692
Abbott Architect c System | 5-14 | 5.56 | 0.24 | 5.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 10.628
Abbott Architect c System | 5-14 | 9.36 | 0.2 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 4.512
Abbott Architect c System | 5-14 | 5.51 | 0.2 | 4.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 8.236
Abbott Architect c System | 5-14 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect c System | 5-14 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect c (enzymatic) | 26 | 6.53 | -0.05 | 1.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 3.086
Abbott Architect c (enzymatic) | 26 | 8.39 | 0 | 1.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 3.332
Abbott Architect c (enzymatic) | 26 | 5.53 | -0.12 | 2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 3.8
Abbott Architect c (enzymatic) | 53 | 6.44 | -0.05 | 4.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 9.162
Abbott Architect c (enzymatic) | 53 | 6.98 | 0.01 | 4.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 9.614
Abbott Architect c (enzymatic) | 53 | 9.39 | -0.26 | 5.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 10.716
Abbott Architect c (enzymatic) | 42 | 10.15 | -0.13 | 1.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 2.418
Abbott Architect c (enzymatic) | 42 | 6.66 | -0.16 | 1.4 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 2.584
Abbott Architect c (enzymatic) | 32 | 8.55 | -0.08 | 1.1 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 2.076
Abbott Architect c (enzymatic) | 32 | 5.18 | -0.14 | 1.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 3.192
Abbott Architect c (enzymatic) | 47 | 5.67 | -0.17 | 2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 3.75
Abbott Architect c (enzymatic) | 47 | 11.76 | 0.05 | 2.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 4.166
Abbott Architect c (enzymatic) | 47 | 9.52 | -0.01 | 2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 3.91
Abbott Architect c (enzymatic) | 47 | 4.9 | -0.14 | 2.4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 4.564
Abbott Architect c (enzymatic) | 47 | 7.22 | -0.16 | 1.7 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 3.172
Abbott Architect c (enzymatic) | 59-80 | 11.89 | 0.2 | 1.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 3.532
Abbott Architect c (enzymatic) | 59-80 | 5.78 | -0.15 | 1.4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 2.594
Abbott Architect c (enzymatic) | 59-80 | 5.07 | -0.1 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 3.232
Abbott Architect c (enzymatic) | 59-80 | 8.1 | -0.04 | 1.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 2.704
Abbott Architect c (enzymatic) | 59-80 | 9.3 | 0 | 1.1 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 2.156
Abbott Architect c (enzymatic) | 90-103 | 5.22 | -0.1 | 1.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 2.84
Abbott Architect c (enzymatic) | 90-103 | 9.24 | 0.08 | 1.3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 2.628
Abbott Architect c (enzymatic) | 90-103 | 5.24 | -0.07 | 1.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 2.87
Abbott Architect c (enzymatic) | 90-103 | 12.24 | 0.21 | 1.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 2.758
Abbott Architect c (enzymatic) | 90-103 | 5.83 | -0.11 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 3.026
Abbott Architect c (enzymatic) | 97 | 5.17 | -0.1 | 2.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Abbott | 4.016
Abbott Architect c (enzymatic) | 97 | 10.6 | 0.01 | 2.5 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Abbott | 4.91
Abbott Architect c (enzymatic) | 97 | 6.04 | -0.16 | 2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Abbott | 3.76
Abbott Architect c (enzymatic) | 97 | 12.31 | 0.08 | 2.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Abbott | 4.196
Abbott Architect c (enzymatic) | 97 | 7.42 | -0.09 | 1.9 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Abbott | 3.634
Abbott Architect c (enzymatic) | 109-132 | 9.04 | -0.07 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Abbott | 3.066
Abbott Architect c (enzymatic) | 109-132 | 5.88 | -0.13 | 1.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Abbott | 3.006
Abbott Architect c (enzymatic) | 109-132 | 11.75 | 0.04 | 1.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Abbott | 2.588
Abbott Architect c (enzymatic) | 109-132 | 4.91 | -0.11 | 1.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Abbott | 2.83
Abbott Architect c (enzymatic) | 109-132 | 7.48 | -0.1 | 1.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Abbott | 2.644
Abbott Architect c (enzymatic) | 146-170 | 6.31 | -0.1 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Abbott | 3.036
Abbott Architect c (enzymatic) | 146-170 | 9.43 | -0.1 | 1.6 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Abbott | 3.036
Abbott Architect c (enzymatic) | 146-170 | 5.25 | -0.09 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Abbott | 3.242
Abbott Architect c (enzymatic) | 146-170 | 8.42 | -0.09 | 1.5 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Abbott | 2.85
Abbott Architect c (enzymatic) | 146-170 | 7.23 | -0.02 | 1.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Abbott | 2.92
Abbott Architect c (enzymatic) | 158 | 7.32 | -0.1 | 1.3 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Abbott | 2.448
Abbott Architect c (enzymatic) | 158 | 5.05 | -0.15 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Abbott | 3.182
Abbott Architect c (enzymatic) | 158 | 8.27 | -0.04 | 1.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Abbott | 2.116
Abbott Architect c (enzymatic) | 158 | 10.41 | 0.08 | 1.1 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Abbott | 2.236
Abbott Architect c (enzymatic) | 158 | 5.77 | -0.1 | 1.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Abbott | 2.448
Abbott Architect c (enzymatic) | 189 | 7.14 | -0.11 | 1.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Abbott | 2.438
Abbott Architect c (enzymatic) | 189 | 6.67 | -0.14 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Abbott | 2.604
Abbott Architect c (enzymatic) | 189 | 8.6 | -0.05 | 1.2 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Abbott | 2.302
Abbott Architect c (enzymatic) | 189 | 9.47 | -0.03 | 1.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Abbott | 2.322
Abbott Architect c (enzymatic) | 189 | 5.29 | -0.16 | 1.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Abbott | 2.388
Abbott Architect c (enzymatic) | 200 | 6.99 | -0.16 | 1.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Abbott | 2.388
Abbott Architect c (enzymatic) | 200 | 5.05 | -0.14 | 1.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Abbott | 2.8
Abbott Architect c (enzymatic) | 200 | 8.21 | -0.21 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Abbott | 2.338
Abbott Architect c (enzymatic) | 200 | 9.7 | -0.09 | 1.2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Abbott | 2.262
Abbott Architect c (enzymatic) | 200 | 5.91 | -0.21 | 1.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Abbott | 2.534
Abbott Architect c (enzymatic) | 174 | 6.18 | -0.13 | 1.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Abbott | 2.026
Abbott Architect c (enzymatic) | 174 | 9.09 | -0.02 | 0.9 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Abbott | 1.744
Abbott Architect c (enzymatic) | 174 | 5.17 | -0.14 | 1.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Abbott | 2.016
Abbott Architect c (enzymatic) | 174 | 8.3 | -0.05 | 1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Abbott | 1.91
Abbott Architect c (enzymatic) | 174 | 7.53 | -0.09 | 1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Abbott | 1.87
Abbott Architect c System | 205 | 6.39 | -0.02 | 1.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Abbott | 2.528
Abbott Architect c System | 205 | 8.7 | 0.1 | 1.1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Abbott | 2.256
Abbott Architect c System | 205 | 5.38 | -0.04 | 1.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Abbott | 2.704
Abbott Architect c System | 205 | 7.41 | 0.03 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Abbott | 2.382
Abbott Architect c System | 205 | 9.93 | 0.18 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Abbott | 2.532
Abbott Architect c System | 234 | 5.14 | -0.11 | 1.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Abbott | 3.026
Abbott Architect c System | 234 | 6.37 | -0.04 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Abbott | 3.096
Abbott Architect c System | 234 | 8.35 | 0.14 | 1.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Abbott | 2.688
Abbott Architect c System | 234 | 5.4 | -0.07 | 1.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Abbott | 2.674
Abbott Architect c System | 234 | 9.66 | 0.18 | 1.2 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Abbott | 2.532
Abbott Architect c System | 235 | 8.1 | 0.13 | 1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Abbott | 2.09
Abbott Architect c System | 235 | 5.9 | 0.01 | 1 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Abbott | 1.97
Abbott Architect c System | 235 | 5.1 | -0.03 | 1.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Abbott | 2.126
Abbott Architect c System | 235 | 7.46 | 0.06 | 1.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Abbott | 2.216
Abbott Architect c System | 235 | 9.35 | 0.18 | 1.2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Abbott | 2.532
Abbott Architect c System | 214 | 7.86 | 0.05 | 1.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Abbott | 2.206
Abbott Architect c System | 214 | 5.82 | -0.09 | 1.1 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Abbott | 2.066
Abbott Architect c System | 214 | 8.76 | 0.07 | 0.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Abbott | 1.834
Abbott Architect c System | 214 | 5.79 | -0.07 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Abbott | 2.478
Abbott Architect c System | 214 | 9.6 | 0.04 | 1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Abbott | 2
Abbott Architect c System | 236 | 5.36 | -0.12 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Abbott | 3.212
Abbott Architect c System | 236 | 8.31 | 0.05 | 1.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Abbott | 2.598
Abbott Architect c System | 236 | 5.04 | -0.11 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Abbott | 3.222
Abbott Architect c System | 236 | 10.22 | 0.17 | 0.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Abbott | 1.934
Abbott Architect c System | 236 | 5.91 | -0.08 | 1.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Abbott | 2.468
Abbott Architect c System | 78 | 7.21 | 0.1 | 4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Abbott | 7.94
Abbott Architect c System | 78 | 9.53 | 0.21 | 3.5 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Abbott | 7.07
Abbott Architect c System | 78 | 6.1 | 0.03 | 3.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Abbott | 6.302
Abbott Architect c System | 57 | 8.06 | 0.01 | 3.2 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Abbott | 6.282
Abbott Architect c System | 57 | 5.28 | -0.02 | 3.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Abbott | 6.84
Abbott Architect c System | 57 | 6.15 | 0.01 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Abbott | 6.086
Abbott Architect c System | 210 | 9.28 | -0.09 | 1.5 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Abbott | 2.85
Abbott Architect c System | 210 | 4.87 | -0.05 | 1.8 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Abbott | 3.478
Abbott Architect c System | 210 | 5.96 | -0.08 | 1.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Abbott | 3.056
Abbott Architect c System | 210 | 5.14 | -0.14 | 1.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Abbott | 2.996
Abbott Architect c System | 210 | 7.95 | -0.13 | 1.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Abbott | 2.026
Abbott Architect c System | 233 | 5.42 | -0.04 | 1.6 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Abbott | 3.096
Abbott Architect c System | 233 | 5.62 | -0.04 | 1.5 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Abbott | 2.9
Abbott Architect c System | 233 | 9.39 | 0.08 | 1.8 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Abbott | 3.608
Abbott Architect c System | 233 | 5.22 | -0.06 | 1.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Abbott | 3.076
Abbott Architect c System | 233 | 7.4 | -0.01 | 1.6 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Abbott | 3.126
Abbott Architect c System | 240 | 8.94 | 0.08 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Abbott | 2.824
Abbott Architect c System | 240 | 6.74 | -0.1 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Abbott | 2.84
Abbott Architect c System | 240 | 9.49 | 0.1 | 1.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Abbott | 2.844
Abbott Architect c System | 240 | 5.98 | -0.15 | 1.6 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Abbott | 2.986
Abbott Architect c System | 240 | 7.49 | -0.03 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Abbott | 3.106
Abbott Architect c System | 217 | 5.63 | -0.17 | 1.8 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Abbott | 3.358
Abbott Architect c System | 217 | 9.16 | 0.04 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Abbott | 2.98
Abbott Architect c System | 217 | 5.34 | -0.13 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Abbott | 3.006
Abbott Architect c System | 217 | 7.25 | 0 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Abbott | 2.94
Abbott Architect c System | 217 | 9.31 | 0.05 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Abbott | 2.99
Abbott Architect c System | 225 | 5.1 | -0.11 | 1.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Abbott | 2.83
Abbott Architect c System | 225 | 6.6 | -0.08 | 1.3 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Abbott | 2.468
Abbott Architect c System | 225 | 8.26 |  | 1.7 | GH5-13 |  | 2021 | CAP | Fresh | Abbott | 3.332
Abbott Architect c System | 225 | 5.93 | -0.05 | 1.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Abbott | 2.302
Abbott Architect c System | 225 | 10.35 | 0.23 | 1 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Abbott | 2.19
Abbott Architect I | 53 | 7.07 | 0.49 | 5.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 11.466
Abbott Architect I | 53 | 9.26 | 0.87 | 6.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 12.826
Abbott Architect I | 53 | 6.02 | 0.37 | 4.2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 8.602
Abbott Architect I | 58 | 6.47 | -0.02 | 3.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 7.232
Abbott Architect I | 58 | 6.85 | -0.12 | 3.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 6.152
Abbott Architect I | 58 | 9.42 | -0.23 | 3.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 6.042
Abbott Architect I | 35 | 10.28 | 0 | 8.4 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 16.464
Abbott Architect I | 35 | 6.86 | 0.04 | 8 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 15.72
Abbott Architect I | 28 | 8.79 | 0.16 | 8.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 16.82
Abbott Architect I | 28 | 5.47 | 0.15 | 8.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 16.026
Abbott Architect I | 21 | 5.87 | 0.03 | 6.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 12.182
Abbott Architect I | 21 | 11.5 | -0.21 | 6.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 12.53
Abbott Architect I | 21 | 9.65 | 0.12 | 7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 13.84
Abbott Architect I | 21 | 5.2 | 0.16 | 8.6 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 17.016
Abbott Architect I | 21 | 7.49 | 0.11 | 6.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 13.046
Abbott Architect I | 16-23 | 11.33 | -0.36 | 4.8 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 9.048
Abbott Architect I | 16-23 | 5.78 | -0.15 | 3.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 7.494
Abbott Architect I | 16-23 | 5.09 | -0.08 | 5.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 10.896
Abbott Architect I | 16-23 | 8.01 | -0.13 | 5.1 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 9.866
Abbott Architect I | 16-23 | 9.17 | -0.13 | 5.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 10.65
Abbott Architect I | 12-17 | 5.24 | -0.08 | 4.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 8.348
Abbott Architect I | 12-17 | 8.91 | -0.25 | 5.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 10.334
Abbott Architect I | 12-17 | 5.26 | -0.05 | 5.9 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 11.514
Abbott Architect I | 12-17 | 11.36 | -0.67 | 4.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 8.542
Abbott Architect I | 12-17 | 5.74 | -0.2 | 4.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 8.62
Abbott Architect I | 11 | 5.32 | 0.05 | 5.4 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Abbott | 10.634
Abbott Architect I | 11 | 10.28 | -0.31 | 5.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Abbott | 9.882
Abbott Architect I | 11 | 5.99 | -0.21 | 4.9 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Abbott | 9.394
Abbott Architect I | 11 | 11.73 | -0.5 | 5.9 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Abbott | 11.064
Abbott Architect I | 11 | 7.29 | -0.22 | 5.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Abbott | 10.756
Abbott Architect I | 8-11 | 8.93 | -0.18 | 5.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Abbott | 10.404
Abbott Architect I | 8-11 | 5.87 | -0.14 | 5.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Abbott | 10.64
Abbott Architect I | 8-11 | 11.09 | -0.62 | 6.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Abbott | 11.728
Abbott Architect I | 8-11 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect I | 8-11 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect I | 0-12 | 6.34 | -0.07 | 5.4 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Abbott | 10.514
Abbott Architect I | 0-12 | 9.36 | -0.17 | 5.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Abbott | 11.002
Abbott Architect I | 0-12 | 5.42 | 0.08 | 2.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Abbott | 5.372
Abbott Architect I | 0-12 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Abbott | 0
Abbott Architect I | 0-12 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Abbott | 0
Abbott Architect I | 31 | 8.02 | -0.03 | 4.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Abbott | 8.986
Abbott Architect I | 31 | 5.37 | 0.07 | 5.2 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Abbott | 10.262
Abbott Architect I | 31 | 6.14 | 0 | 5.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Abbott | 10.78
Abbott Architect | 17 | 8.0152 | -6 | 7.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Abbott | 8.308
Abbott Architect | 17 | 5.1512 | -5.4 | 10.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Abbott | 14.984
Abbott Architect | 24 | 8.6648 | 1.1 | 3.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Abbott | 8.744
Abbott Architect | 24 | 5.5904 | -0.6 | 4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Abbott | 7.24
Abbott Architect | 29 | 5.1695 | -3.5 | 6.9 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 10.024
Abbott Architect | 29 | 5.1878 | -3.3 | 6.4 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Abbott | 9.244
Abbott Architect | 39 | 5.5263 | 0.4 | 2.7 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 5.692
Abbott Architect | 39 | 5.5263 | 0.4 | 2.9 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 6.084
Abbott Architect/Aeroset | 49 | 12.5 |  | 5.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Abbott | 9.996
Abbott Architect/Aeroset | 49 | 5.1 |  | 4.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Abbott | 8.232
Abbott Architect/Aeroset | 49 | 9.2 |  | 5.1 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Abbott | 9.996
Abbott Architect/Aeroset | 47 | 11.7 |  | 4.8 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Abbott | 9.408
Abbott Architect/Aeroset | 47 | 6.9 |  | 4.3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Abbott | 8.428
Abbott Architect/Aeroset | 47 | 5 |  | 5.3 | GH2-06 | 5 | 2008 | CAP | Fresh | Abbott | 10.388
Abbott Architect/Aeroset | 47 | 4.9 |  | 5.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Abbott | 11.368
Abbott Architect/Aeroset | 47 | 8.2 |  | 4.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Abbott | 8.428
Abbott Architect/Aeroset | 47 | 5.8 |  | 8 | GH2-03 | 6 | 2009 | CAP | Fresh | Abbott | 15.68
Abbott not specified/other | 72 | 5.4897 | 0 | 4.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 8.036
Abbott not specified/other | 71 | 5.4897 | 0 | 4.5 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 8.82
Abbott not specified/other | 6 | 6.1394 | 1.3 | 5.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Abbott | 11.296
Abbott not specified/other | 6 | 7.6766 | 2.3 | 4 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Abbott | 10.14
Alere Afinion | 142 | 8.4086 | -1.7 | 2.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Abbott | 3.592
Alere Afinion | 141 | 5.4989 | -1.6 | 5.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Abbott | 8.396
Alere Afinion | 150 | 5.4074 | -0.9 | 4.2 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 7.332
Alere Afinion | 147 | 5.4165 | -0.8 | 4.5 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Abbott | 8.02
Alere Afinion | 32-119 | 6.3 | -0.11 | 1.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Alere | 3.614
Alere Afinion | 32-119 | 9.49 | -0.04 | 1.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Alere | 3.488
Alere Afinion | 32-119 | 5.26 | -0.08 | 1.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Alere | 3.644
Alere Afinion | 32-119 | 8.49 | -0.02 | 1.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Alere | 3.312
Alere Afinion | 32-119 | 7.35 | 0.1 | 1.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Alere | 3.432
Alere Afinion | 22 | 7.18 | -0.24 | 1.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Alere | 3.484
Alere Afinion | 22 | 5.14 | -0.06 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Alere | 3.272
Alere Afinion | 22 | 8.14 | -0.17 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Alere | 3.162
Alere Afinion | 22 | 10.25 | -0.08 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Alere | 3.644
Alere Afinion | 22 | 5.79 | -0.08 | 1.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Alere | 3.252
Alere Afinion | 113 | 7.15 | -0.1 | 2 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Alere | 3.82
Alere Afinion | 113 | 6.74 | -0.07 | 2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Alere | 3.85
Alere Afinion | 113 | 8.78 | 0.13 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Alere | 4.246
Alere Afinion | 113 | 9.23 | -0.27 | 2.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Alere | 4.042
Alere Afinion | 113 | 5.34 | -0.11 | 1.5 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Alere | 2.83
Alere Afinion | 153 | 7.07 | -0.08 | 1.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Alere | 3.056
Alere Afinion | 153 | 5.24 | 0.05 | 2.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Alere | 4.95
Alere Afinion | 153 | 8.28 | -0.14 | 1.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Alere | 2.996
Alere Afinion | 153 | 9.73 | -0.06 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Alere | 3.468
Alere Afinion | 153 | 6.1 | -0.02 | 2.3 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Alere | 4.488
Alere Afinion | 37 | 6.13 | -0.18 | 2.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Alere | 3.936
Alere Afinion | 37 | 8.98 | -0.13 | 1.4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Alere | 2.614
Alere Afinion | 37 | 5.3 | -0.01 | 2.5 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Alere | 4.89
Alere Afinion | 37 | 8.21 | -0.14 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Alere | 3.388
Alere Afinion | 37 | 7.48 | -0.14 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Alere | 2.8
Alere Afinion | 76 | 5.9107 | -1.2 | 3.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Alere | 5.66
Alere Afinion | 74 | 7.4295 | -0.2 | 3.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Alere | 6.072
Alere Afinion | 165 | 6.1852 | -0.1 | 3.9 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Alere | 7.544
Alere Afinion | 168 | 7.4387 | -0.2 | 3.6 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Alere | 6.856
Alere Afinion | 274 | 7.4021 | -0.7 | 3.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Alere | 5.376
Alere Afinion | 269 | 5.9564 | -0.6 | 3.6 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Alere | 6.456
Alere Afinion 2 | 12 |  |  |  | GH5-01 | 5.46 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 |  |  |  | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 | 9.15 | -0.16 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Alere | 3.76
Alere Afinion 2 | 12 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 22 | 5.2 | -0.05 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Alere | 5.046
Alere Afinion 2 | 22 | 6.38 | -0.03 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Alere | 5.262
Alere Afinion 2 | 22 | 8.15 | -0.06 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Alere | 4.056
Alere Afinion 2 | 22 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 22 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 24 | 7.85 | -0.12 | 1.9 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Alere | 3.604
Alere Afinion 2 | 24 | 5.88 | -0.01 | 2.1 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Alere | 4.106
Alere Afinion 2 | 24 | 5.11 | -0.02 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Alere | 4.684
Alere Afinion 2 | 24 | 7.4 | 0 | 2.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Alere | 4.116
Alere Afinion 2 | 24 | 8.96 | -0.21 | 1.8 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Alere | 3.318
Alere Afinion 2 | 29 | 7.7 | -0.11 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Alere | 2.634
Alere Afinion 2 | 29 | 5.83 | -0.08 | 1.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Alere | 3.252
Alere Afinion 2 | 29 | 8.64 | -0.05 | 1.5 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Alere | 2.89
Alere Afinion 2 | 29 | 5.79 | -0.07 | 1.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Alere | 3.458
Alere Afinion 2 | 29 | 9.62 | 0.06 | 1.7 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Alere | 3.392
Alere Afinion 2 | 56 | 5.42 | -0.06 | 2.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Alere | 4.056
Alere Afinion 2 | 56 | 8.09 | -0.17 | 1.6 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Alere | 2.966
Alere Afinion 2 | 56 | 5.11 | -0.04 | 1.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Alere | 3.684
Alere Afinion 2 | 56 | 9.83 | -0.22 | 1.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Alere | 2.72
Alere Afinion 2 | 56 | 5.86 | -0.13 | 1.8 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Alere | 3.398
Alere Afinion 2 | 100 | 8.81 | -0.05 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Alere | 3.086
Alere Afinion 2 | 100 | 6.85 | 0.01 | 2.1 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Alere | 4.126
Alere Afinion 2 | 100 | 9.3 | -0.09 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Alere | 3.83
Alere Afinion 2 | 100 | 6.1 | -0.03 | 2.1 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Alere | 4.086
Alere Afinion 2 | 100 | 7.61 | 0.09 | 1.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Alere | 2.834
Alere Afinion 2 | 38 | 5.74 | -0.06 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Alere | 4.252
Alere Afinion 2 | 38 | 9.19 | 0.07 | 1.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Alere | 3.794
Alere Afinion 2 | 38 | 5.44 | -0.03 | 1.8 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Alere | 3.498
Alere Afinion 2 | 38 | 7.34 | 0.09 | 1.8 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Alere | 3.618
Alere Afinion 2 | 38 | 9.28 | 0.02 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Alere | 3.94
Alere Afinion 2 | 114 | 5.18 | -0.03 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Alere | 4.282
Alere Afinion 2 | 114 | 6.59 | -0.09 | 1.8 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Alere | 3.438
Alere Afinion 2 | 114 | 9.12 |  | 2.5 | GH5-13 |  | 2021 | CAP | Fresh | Alere | 4.9
Alere Afinion 2 | 114 | 5.84 | -0.14 | 2.3 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Alere | 4.368
Alere Afinion 2 | 114 | 10.11 | -0.01 | 2.3 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Alere | 4.498
Alere Afinion AS100 | 97 |  |  |  | GH5-01 | 5.46 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion AS100 | 97 |  |  |  | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion AS100 | 97 | 9.24 | -0.07 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Alere | 2.478
Alere Afinion AS100 | 97 | 5.28 | 0 | 2.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Alere | 4.312
Alere Afinion AS100 | 97 | 7.28 | -0.13 | 1.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Alere | 2.81
Alere Afinion AS100 | 150 | 9.22 | -0.15 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Alere | 3.378
Alere Afinion AS100 | 150 | 4.96 | 0.04 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Alere | 6.116
Alere Afinion AS100 | 150 | 5.96 | -0.08 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Alere | 4.624
Alere Afinion AS100 | 150 | 5.23 | -0.05 | 2.7 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Alere | 5.242
Alere Afinion AS100 | 150 | 8.02 | -0.06 | 1.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Alere | 3.076
Alere Afinion AS100 | 28 | 6.16 | -0.25 | 1.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Alere | 3.278
Alere Afinion AS100 | 28 | 8.54 | -0.06 | 1.9 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Alere | 3.664
Alere Afinion AS100 | 28 | 5.36 | -0.06 | 1.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Alere | 3.272
Alere Afinion AS100 | 28 | 7.3 | -0.08 | 1.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Alere | 3.448
Alere Afinion AS100 | 28 | 9.93 | 0.18 | 1.3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Alere | 2.728
Alere Afinion AS100 | 136 | 5.11 | -0.14 | 1.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Alere | 3.192
Alere Afinion AS100 | 136 | 6.23 | -0.18 | 2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Alere | 3.74
Alere Afinion AS100 | 136 | 8.14 | -0.07 | 1.7 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Alere | 3.262
Alere Afinion AS100 | 136 | 5.3 | -0.17 | 2.1 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Alere | 3.946
Alere Afinion AS100 | 136 | 9.38 | -0.1 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Alere | 3.036
Alere Afinion AS100 | 79 | 7.86 | -0.11 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Alere | 4.202
Alere Afinion AS100 | 79 | 5.85 | -0.04 | 2.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Alere | 5.252
Alere Afinion AS100 | 79 | 5.09 | -0.04 | 2.7 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Alere | 5.252
Alere Afinion AS100 | 79 | 7.39 | -0.01 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Alere | 5.87
Alere Afinion AS100 | 79 | 9.01 | -0.16 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Alere | 5.132
Alere Afinion AS100 | 26 | 7.65 | -0.16 | 1.7 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Alere | 3.172
Alere Afinion AS100 | 26 | 5.8 | -0.11 | 2.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Alere | 4.986
Alere Afinion AS100 | 26 | 8.6 | -0.09 | 2.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Alere | 4.026
Alere Afinion AS100 | 26 | 5.73 | -0.13 | 3.1 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Alere | 5.946
Alere Afinion AS100 | 26 | 9.6 | 0.04 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Alere | 3.96
Alere Afinion AS100 | 112 | 5.3 | -0.18 | 2.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Alere | 3.936
Alere Afinion AS100 | 112 | 8.05 | -0.21 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Alere | 3.514
Alere Afinion AS100 | 112 | 5.06 | -0.09 | 2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Alere | 3.83
Alere Afinion AS100 | 112 | 9.85 | -0.2 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Alere | 3.132
Alere Afinion AS100 | 112 | 5.85 | -0.14 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Alere | 4.76
Alere Afinion AS100 | 105 | 8.64 | -0.22 | 1.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Alere | 3.308
Alere Afinion AS100 | 105 | 6.72 | -0.12 | 2.1 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Alere | 3.996
Alere Afinion AS100 | 105 | 9.21 | -0.18 | 1.9 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Alere | 3.544
Alere Afinion AS100 | 105 | 6 | -0.13 | 1.7 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Alere | 3.202
Alere Afinion AS100 | 105 | 7.59 | 0.07 | 1.9 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Alere | 3.794
Alere Afinion AS100 | 20 | 5.57 | -0.23 | 2.3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Alere | 4.278
Alere Afinion AS100 | 20 | 9.04 | -0.08 | 1.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Alere | 3.448
Alere Afinion AS100 | 20 | 5.3 | -0.17 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Alere | 3.554
Alere Afinion AS100 | 20 | 7.22 | -0.03 | 1.7 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Alere | 3.302
Alere Afinion AS100 | 20 | 9.2 | -0.06 | 1.8 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Alere | 3.468
Alere Afinion AS100 | 88 | 5.09 | -0.12 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Alere | 4.192
Alere Afinion AS100 | 88 | 6.5 | -0.18 | 2.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Alere | 4.132
Alere Afinion AS100 | 88 | 9.07 |  | 3 | GH5-13 |  | 2021 | CAP | Fresh | Alere | 5.88
Alere Afinion AS100 | 88 | 5.74 | -0.24 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Alere | 3.288
Alere Afinion AS100 | 88 | 9.9 | -0.22 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Alere | 3.504
Arkray Adams HA-8180 series | 16 | 5.39 | 0.07 | 1.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Arkray | 2.226
Arkray Adams HA-8180 series | 16 | 9.25 | 0.09 | 0.6 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Arkray | 1.266
Arkray Adams HA-8180 series | 16 | 5.37 | 0.06 | 1.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Arkray | 2.608
Arkray Adams HA-8180 series | 16 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 16 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 0-17 | 9.18 | 0.07 | 0.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Arkray | 1.05
Arkray Adams HA-8180 series | 0-17 | 6.03 | 0.02 | 1.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Arkray | 2.96
Arkray Adams HA-8180 series | 0-17 | 11.81 | 0.1 | 1.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Arkray | 2.256
Arkray Adams HA-8180 series | 0-17 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 0-17 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 19 | 7.35 | 0.1 | 1.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Arkray | 2.256
Arkray Adams HA-8180 series | 19 | 6.82 | 0.01 | 1.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Arkray | 3.146
Arkray Adams HA-8180 series | 19 | 8.7 | 0.05 | 1.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Arkray | 2.794
Arkray Adams HA-8180 series | 19 |  |  |  | GH5-14 | 9.50 | 2017 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 19 |  |  |  | GH5-15 | 5.45 | 2017 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8180 series | 17 | 7.16 | 0.01 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Arkray | 2.754
Arkray Adams HA-8180 series | 17 | 5.19 | 0 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Arkray | 3.92
Arkray Adams HA-8180 series | 17 | 8.46 | 0.04 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Arkray | 2.784
Arkray Adams HA-8180 series | 17 | 9.83 | 0.04 | 1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Arkray | 2
Arkray Adams HA-8180 series | 17 | 6.09 | -0.03 | 1.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Arkray | 2.91
Arkray Adams HA-8180 series | 11 | 6.26 | -0.05 | 0.8 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Arkray | 1.518
Arkray Adams HA-8180 series | 11 | 9.09 | -0.02 | 1.8 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Arkray | 3.508
Arkray Adams HA-8180 series | 11 | 5.23 | -0.08 | 0.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Arkray | 1.684
Arkray Adams HA-8180 series | 11 | 8.38 | 0.03 | 1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Arkray | 1.99
Arkray Adams HA-8180 series | 11 | 7.64 | 0.02 | 0.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Arkray | 1.392
Arkray Adams HA-8180 series | 21 | 9.34 | -0.03 | 1.5 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Arkray | 2.91
Arkray Adams HA-8180 series | 21 | 4.94 | 0.02 | 1.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Arkray | 2.96
Arkray Adams HA-8180 series | 21 | 6.07 | 0.03 | 1.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Arkray | 2.774
Arkray Adams HA-8180 series | 21 | 5.28 | 0 | 1.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Arkray | 2.744
Arkray Adams HA-8180 series | 21 | 8.13 | 0.05 | 1.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Arkray | 2.598
Arkray Adams HA-8180 series | 17 | 5.42 | -0.04 | 1.8 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Arkray | 3.488
Arkray Adams HA-8180 series | 17 | 5.54 | -0.12 | 2.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Arkray | 5.564
Arkray Adams HA-8180 series | 17 | 9.41 | 0.1 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Arkray | 3.432
Arkray Adams HA-8180 series | 17 | 5.2 | -0.08 | 1.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Arkray | 2.272
Arkray Adams HA-8180 series | 17 | 7.43 | 0.02 | 0.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Arkray | 1.784
Arkray Adams HA-8180 series | 22 | 8.87 | 0.01 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Arkray | 2.754
Arkray Adams HA-8180 series | 22 | 6.81 | -0.03 | 1.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Arkray | 3.106
Arkray Adams HA-8180 series | 22 | 9.36 | -0.03 | 1.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Arkray | 2.91
Arkray Adams HA-8180 series | 22 | 6.05 | -0.08 | 1.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Arkray | 2.664
Arkray Adams HA-8180 series | 22 | 7.42 | -0.1 | 1.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Arkray | 2.056
Arkray Adams HA-8180 series | 14 | 6.42 | 0.01 | 1.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Arkray | 2.362
Arkray Adams HA-8180 series | 14 | 8.67 | 0.07 | 1.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Arkray | 3.01
Arkray Adams HA-8180 series | 14 | 5.33 | -0.09 | 1.1 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Arkray | 2.066
Arkray Adams HA-8180 series | 14 | 7.44 | 0.06 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Arkray | 2.412
Arkray Adams HA-8180 series | 14 | 9.89 | 0.14 | 2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Arkray | 4.06
Arkray Adams HA-8180 series | 22 | 5.26 | 0.01 | 1.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Arkray | 2.754
Arkray Adams HA-8180 series | 22 | 6.47 | 0.06 | 1.2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Arkray | 2.412
Arkray Adams HA-8180 series | 22 | 8.25 | 0.04 | 1.5 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Arkray | 2.98
Arkray Adams HA-8180 series | 22 | 5.53 | 0.06 | 1.3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Arkray | 2.608
Arkray Adams HA-8180 series | 22 | 9.54 | 0.06 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Arkray | 3.196
Arkray Adams HA-8180 series | 16 | 8.03 | 0.06 | 1.6 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Arkray | 3.196
Arkray Adams HA-8180 series | 16 | 5.92 | 0.03 | 1.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Arkray | 3.362
Arkray Adams HA-8180 series | 16 | 5.14 | 0.01 | 1.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Arkray | 2.362
Arkray Adams HA-8180 series | 16 | 7.3 | -0.1 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Arkray | 3.624
Arkray Adams HA-8180 series | 16 | 9.18 | 0.01 | 1.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Arkray | 2.558
Arkray Adams HA-8180 series | 12 | 7.86 | 0.05 | 1.3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Arkray | 2.598
Arkray Adams HA-8180 series | 12 | 5.9 | -0.01 | 0.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Arkray | 1.362
Arkray Adams HA-8180 series | 12 | 8.73 | 0.04 | 1.3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Arkray | 2.588
Arkray Adams HA-8180 series | 12 | 5.93 | 0.07 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Arkray | 2.618
Arkray Adams HA-8180 series | 12 | 9.57 | 0.01 | 1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Arkray | 1.97
Arkray Adams HA-8180 series | 24 | 5.54 | 0.06 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 3.588
Arkray Adams HA-8180 series | 24 | 8.32 | 0.06 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 2.804
Arkray Adams HA-8180 series | 24 | 5.18 | 0.03 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 3.362
Arkray Adams HA-8180 series | 24 | 10.08 | 0.03 | 1.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray | 2.774
Arkray Adams HA-8180 series | 24 | 6.01 | 0.02 | 1.7 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray | 3.352
Arkray Adams HA-8180 series | 17 | 5.74 | -0.06 | 1.4 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Arkray | 2.684
Arkray Adams HA-8180 series | 17 | 9.04 | -0.08 | 1.1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Arkray | 2.076
Arkray Adams HA-8180 series | 17 | 5.37 | -0.1 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Arkray | 3.036
Arkray Adams HA-8180 series | 17 | 7.15 | -0.1 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Arkray | 2.448
Arkray Adams HA-8180 series | 17 | 9.21 | -0.05 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Arkray | 2.89
Arkray Adams HA-8180 series | 26 | 5.2 | -0.01 | 1.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 2.734
Arkray Adams HA-8180 series | 26 | 6.69 | 0.01 | 1.3 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 2.558
Arkray Adams HA-8180 series | 26 | 8.28 |  | 1.4 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 2.744
Arkray Adams HA-8180 series | 26 | 5.96 | -0.02 | 0.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray | 1.744
Arkray Adams HA-8180 series | 26 | 10.06 | -0.06 | 0.8 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray | 1.508
Arkray Adams HA-8190V | 25 | 5.51 | 0.03 | 1.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 2.578
Arkray Adams HA-8190V | 25 | 8.34 | 0.08 | 1.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 2.432
Arkray Adams HA-8190V | 25 | 5.14 | -0.01 | 1.2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 2.342
Arkray Adams HA-8190V | 25 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8190V | 25 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8190V | 25 | 5.15 | -0.06 | 1.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 2.488
Arkray Adams HA-8190V | 25 | 6.68 | 0 | 1.1 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 2.156
Arkray Adams HA-8190V | 25 | 8.32 |  | 1.3 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 2.548
Arkray Adams HA-8190V | 25 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray | 0
Arkray Adams HA-8190V | 25 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray | 0
Arkray The Lab 001 | 17 | 5.49 | 0.01 | 2.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 4.322
Arkray The Lab 001 | 17 | 8.15 | -0.11 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 2.634
Arkray The Lab 001 | 17 | 5.13 | -0.02 | 0.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 1.744
Arkray The Lab 001 | 17 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray | 0
Arkray The Lab 001 | 17 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray | 0
Arkray The Lab 001 | 13 | 5.12 | -0.09 | 0.7 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 1.282
Arkray The Lab 001 | 13 | 6.62 | -0.06 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 2.684
Arkray The Lab 001 | 13 | 8.29 |  | 2.5 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 4.9
Arkray The Lab 001 | 13 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray | 0
Arkray The Lab 001 | 13 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray | 0
Axis-Shield Afinion | 16 | 6.19 | 0.29 | 3.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Axis-Shield | 7.346
Axis-Shield Afinion | 16 | 9.87 | 0.07 | 3.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Axis-Shield | 6.538
Axis-Shield Afinion | 16 | 7.49 | 0.09 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Axis-Shield | 6.95
Axis-Shield Afinion | 11 | 5.57 | 0.37 | 2.8 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Axis-Shield | 5.858
Axis-Shield Afinion | 11 | 8.75 | 0.05 | 1.9 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Axis-Shield | 3.774
Axis-Shield Afinion | 11 | 6.61 | 0.31 | 1.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Axis-Shield | 3.838
Axis-Shield Afinion | 18 | 8.54 | 0.04 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Axis-Shield | 4.352
Axis-Shield Afinion | 18 | 5.6 | 0.2 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Axis-Shield | 5.492
Axis-Shield Afinion | 18 | 6.56 | 0.16 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Axis-Shield | 6.628
Axis-Shield Afinion | 28 | 5.59 | -0.01 | 3.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Axis-Shield | 6.066
Axis-Shield Afinion | 28 | 8.98 | -0.42 | 2.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Axis-Shield | 5.264
Axis-Shield Afinion | 28 | 7.17 | -0.03 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Axis-Shield | 4.87
Axis-Shield Afinion | 23 | 5.54 | 0.14 | 4.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Axis-Shield | 8.176
Axis-Shield Afinion | 23 | 8.1 | -0.2 | 3.4 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Axis-Shield | 6.464
Axis-Shield Afinion | 23 | 5.65 | 0 | 3.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Axis-Shield | 6.86
Axis-Shield Afinion | 24 | 7.14 | 0.03 | 3.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Axis-Shield | 6.498
Axis-Shield Afinion | 24 | 9.02 | -0.3 | 3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Axis-Shield | 5.58
Axis-Shield Afinion | 24 | 6.11 | 0.04 | 3.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Axis-Shield | 6.116
Axis-Shield Afinion | 34 | 7.64 | -0.41 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Axis-Shield | 4.294
Axis-Shield Afinion | 34 | 5.25 | -0.05 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Axis-Shield | 5.438
Axis-Shield Afinion | 34 | 6.04 | -0.1 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Axis-Shield | 5.976
Axis-Shield Afinion | 64 | 6.55 | -0.03 | 3.4 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Axis-Shield | 6.634
Axis-Shield Afinion | 64 | 8.37 | -0.02 | 3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Axis-Shield | 5.86
Axis-Shield Afinion | 64 | 5.73 | 0.08 | 2.5 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Axis-Shield | 4.98
Axis-Shield Afinion | 55 | 6.28 | -0.21 | 3.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Axis-Shield | 7.238
Axis-Shield Afinion | 55 | 6.82 | -0.15 | 4.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Axis-Shield | 8.474
Axis-Shield Afinion | 55 | 9.45 | -0.2 | 4.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Axis-Shield | 9.012
Axis-Shield Afinion | 54 | 6.79 | 0 | 2.4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Axis-Shield | 4.704
Axis-Shield Afinion | 54 | 10.16 | -0.12 | 2.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Axis-Shield | 5.368
Axis-Shield Afinion | 54 | 6.86 | 0.04 | 2.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Axis-Shield | 4.548
Axis-Shield Afinion | 15 | 8.43 | -0.2 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Axis-Shield | 4.308
Axis-Shield Afinion | 15 | 5.39 | 0.07 | 3.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Axis-Shield | 6.93
Axis-Shield Afinion | 19 | 5.95 | 0.11 | 2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Axis-Shield | 4.03
Axis-Shield Afinion | 19 | 11.59 | -0.12 | 3.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Axis-Shield | 6.152
Axis-Shield Afinion | 19 | 9.44 | -0.09 | 1.9 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Axis-Shield | 3.634
Axis-Shield Afinion | 19 | 5.22 | 0.18 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Axis-Shield | 5.668
Axis-Shield Afinion | 19 | 7.47 | 0.09 | 2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Axis-Shield | 4.01
Axis-Shield Afinion | 18-76 | 11.68 | -0.01 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Axis-Shield | 6.066
Axis-Shield Afinion | 18-76 | 5.95 | 0.02 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Axis-Shield | 5.312
Axis-Shield Afinion | 18-76 | 5.32 | 0.15 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Axis-Shield | 6.422
Axis-Shield Afinion | 18-76 | 8.13 | -0.01 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Axis-Shield | 6.85
Axis-Shield Afinion | 18-76 | 9.31 | 0.01 | 4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Axis-Shield | 7.85
Axis-Shield Afinion | 13-57 | 5.26 | -0.06 | 3.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Axis-Shield | 6.016
Axis-Shield Afinion | 13-57 | 9.01 | -0.15 | 2.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Axis-Shield | 5.142
Axis-Shield Afinion | 13-57 | 5.28 | -0.03 | 3.6 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Axis-Shield | 7.026
Axis-Shield Afinion | 13-57 | 12.08 | 0.05 | 2.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Axis-Shield | 4.95
Axis-Shield Afinion | 13-57 | 6.05 | 0.11 | 4.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Axis-Shield | 8.93
Axis-Shield Afinion | 13 | 5.27 | 0 | 2.2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Axis-Shield | 4.312
Axis-Shield Afinion | 13 | 10.47 | -0.12 | 2.1 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Axis-Shield | 3.996
Axis-Shield Afinion | 13 | 6.25 | 0.05 | 2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Axis-Shield | 3.97
Axis-Shield Afinion | 13 | 12.32 | 0.09 | 3.5 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Axis-Shield | 6.95
Axis-Shield Afinion | 13 | 7.43 | -0.08 | 2.7 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Axis-Shield | 5.212
Axis-Shield Afinion | 15-82 | 8.97 | -0.14 | 1.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Axis-Shield | 2.604
Axis-Shield Afinion | 15-82 | 5.99 | -0.02 | 1.9 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Axis-Shield | 3.704
Axis-Shield Afinion | 15-82 | 11.65 | -0.06 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Axis-Shield | 3.86
Axis-Shield Afinion | 15-82 | 5.03 | 0.01 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Axis-Shield | 6.086
Axis-Shield Afinion | 15-82 | 7.5 | -0.08 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Axis-Shield | 4.232
A1cNOW | 23 | 5.05 | -0.35 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bayer/Metrika | 7.49
A1cNOW | 23 | 7.49 | -0.81 | 6.2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bayer/Metrika | 11.342
A1cNOW | 23 | 6.31 | 0.66 | 4.9 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bayer/Metrika | 10.264
A1cNOW | 19 | 7.96 | -0.54 | 5.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bayer/Metrika | 10.24
A1cNOW | 19 | 5.14 | -0.26 | 4.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bayer/Metrika | 9.344
A1cNOW | 19 | 6.02 | -0.38 | 5.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bayer/Metrika | 9.812
A1cNOW | 19 | 6.34 | 0.04 | 3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bayer/Metrika | 5.92
A1cNOW | 19 | 7.58 | -0.02 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bayer/Metrika | 5.076
A1cNOW | 19 | 9.21 | 0.01 | 4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bayer/Metrika | 7.85
A1cNOW | 25 | 4.86 | -0.34 | 4.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bayer/Metrika | 8.872
A1cNOW | 25 | 8.02 | -0.68 | 6 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bayer/Metrika | 11.08
A1cNOW | 25 | 5.85 | -0.45 | 4.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bayer/Metrika | 9.154
A1cNOW | 28 | 5.14 | -0.46 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bayer/Metrika | 6.988
A1cNOW | 28 | 8.49 | -0.91 | 6.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bayer/Metrika | 11.438
A1cNOW | 28 | 6.59 | -0.61 | 5.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bayer/Metrika | 9.974
A1cNOW | 16 | 6.37 | -0.74 | 5.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bayer | 9.648
A1cNOW | 16 | 8.27 | -1.05 | 4.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bayer | 6.986
A1cNOW | 16 | 5.4 | -0.67 | 7.3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bayer | 13.638
A1cNOW | 13 | 7.43 | -0.62 | 5.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bayer | 10.552
A1cNOW | 13 | 4.95 | -0.35 | 4.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bayer | 8.666
A1cNOW | 13 | 5.7 | -0.44 | 5.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bayer | 10.34
Bayer Advia | 24 | 9.9 |  | 6.9 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bayer | 13.524
Bayer Advia | 24 | 5.2 |  | 3.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bayer | 6.86
Bayer Advia | 24 | 8 |  | 5.8 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bayer | 11.368
Bayer Advia | 30 | 6.9 |  | 5.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 11.172
Bayer Advia | 30 | 6.9 |  | 5.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 10.976
Bayer Advia | 30 | 11 |  | 6.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 13.132
Bayer Advia | 30 | 6.9 |  | 5.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 11.172
Bayer Advia | 30 | 6.9 |  | 5.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 10.976
Bayer Advia | 30 | 11 |  | 6.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 13.132
Bayer Advia | 33 | 5.5 |  | 4.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bayer | 8.232
Bayer Advia | 33 | 11.3 |  | 4.9 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bayer | 9.604
Bayer Advia | 33 | 7.7 |  | 5.2 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bayer | 10.192
Bayer Advia | 31 | 9 |  | 4.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bayer | 8.82
Bayer Advia | 31 | 10.9 |  | 8.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bayer | 15.876
Bayer Advia | 31 | 6.4 |  | 8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bayer | 15.68
Bayer DCA 2000 | 156 | 10.4 |  | 3 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bayer | 5.88
Bayer DCA 2000 | 156 | 4.9 |  | 4.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bayer | 9.016
Bayer DCA 2000 | 156 | 8.1 |  | 3.2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bayer | 6.272
Bayer DCA 2000 | 159 | 6.7 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 5.684
Bayer DCA 2000 | 159 | 6.7 |  | 2.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 5.292
Bayer DCA 2000 | 159 | 11.4 |  | 3.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 6.86
Bayer DCA 2000 | 159 | 6.7 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 5.684
Bayer DCA 2000 | 159 | 6.7 |  | 2.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 5.292
Bayer DCA 2000 | 159 | 11.4 |  | 3.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 6.86
Bayer DCA 2000 | 227 | 5.2 |  | 2.9 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bayer | 5.684
Bayer DCA 2000 | 227 | 11.5 |  | 4.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bayer | 9.408
Bayer DCA 2000 | 227 | 7.4 |  | 2.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bayer | 5.292
Bayer DCA 2000 | 171 | 8.8 |  | 3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bayer | 5.88
Bayer DCA 2000 | 171 | 11 |  | 4.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bayer | 8.82
Bayer DCA 2000 | 171 | 6.1 |  | 2.9 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bayer | 5.684
Beckman AU Systems | 67 | 6.45 | -0.13 | 5.4 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Beckman Coulter | 10.454
Beckman AU Systems | 67 | 8.26 | -0.13 | 5.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Beckman Coulter | 11.042
Beckman AU Systems | 67 | 5.5 | -0.15 | 3.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Beckman Coulter | 7.102
Beckman AU Systems | 65 | 6.67 | -0.12 | 4.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Beckman Coulter | 8.308
Beckman AU Systems | 65 | 10 | -0.28 | 3.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Beckman Coulter | 7.364
Beckman AU Systems | 65 | 6.68 | -0.14 | 3.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Beckman Coulter | 7.504
Beckman AU Systems | 40 | 8.29 | -0.34 | 3.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Beckman Coulter | 6.716
Beckman AU Systems | 40 | 5.17 | -0.15 | 3.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Beckman Coulter | 5.926
Beckman AU Systems | 42 | 5.81 | -0.03 | 4.5 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Beckman Coulter | 8.79
Beckman AU Systems | 42 | 11.62 | -0.09 | 3.6 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Beckman Coulter | 6.966
Beckman AU Systems | 42 | 9.42 | -0.11 | 3.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Beckman Coulter | 6.75
Beckman AU Systems | 42 | 5.09 | 0.05 | 3.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Beckman Coulter | 7.302
Beckman AU Systems | 42 | 7.32 | -0.06 | 3.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Beckman Coulter | 6.996
Beckman AU Systems | 42-67 | 11.71 | 0.02 | 3.2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Beckman Coulter | 6.292
Beckman AU Systems | 42-67 | 5.98 | 0.05 | 2.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Beckman Coulter | 5.734
Beckman AU Systems | 42-67 | 5.21 | 0.04 | 3.5 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Beckman Coulter | 6.9
Beckman AU Systems | 42-67 | 8.06 | -0.08 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Beckman Coulter | 4.82
Beckman AU Systems | 42-67 | 9.21 | -0.09 | 2.4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Beckman Coulter | 4.614
Beckman AU Systems | 50-83 | 5.29 | -0.03 | 3.4 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Beckman Coulter | 6.634
Beckman AU Systems | 50-83 | 9.17 | 0.01 | 3.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Beckman Coulter | 7.262
Beckman AU Systems | 50-83 | 5.29 | -0.02 | 3.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Beckman Coulter | 6.252
Beckman AU Systems | 50-83 | 12.1 | 0.07 | 3.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Beckman Coulter | 7.322
Beckman AU Systems | 50-83 | 5.9 | -0.04 | 3.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Beckman Coulter | 6.428
Beckman AU Systems | 55 | 5.3 | 0.03 | 3.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Beckman Coulter | 6.106
Beckman AU Systems | 55 | 10.59 | 0 | 4.1 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Beckman Coulter | 8.036
Beckman AU Systems | 55 | 6.21 | 0.01 | 4.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Beckman Coulter | 8.83
Beckman AU Systems | 55 | 12.32 | 0.09 | 4.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Beckman Coulter | 8.126
Beckman AU Systems | 55 | 7.47 | -0.04 | 2.8 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Beckman Coulter | 5.448
Beckman AU Systems | 57-89 | 8.98 | -0.13 | 4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Beckman Coulter | 7.71
Beckman AU Systems | 57-89 | 5.87 | -0.14 | 3.9 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Beckman Coulter | 7.504
Beckman AU Systems | 57-89 | 11.6 | -0.11 | 4.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Beckman Coulter | 8.906
Beckman AU Systems | 57-89 | 4.92 | -0.1 | 4.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Beckman Coulter | 8.72
Beckman AU Systems | 57-89 | 7.42 | -0.16 | 4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Beckman Coulter | 7.68
Beckman AU Systems | 59-92 | 6.58 | 0.17 | 3.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Beckman Coulter | 7.814
Beckman AU Systems | 59-92 | 9.6 | 0.07 | 3.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Beckman Coulter | 6.734
Beckman AU Systems | 59-92 | 5.44 | 0.1 | 3.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Beckman Coulter | 6.764
Beckman AU Systems | 59-92 | 8.65 | 0.14 | 4.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Beckman Coulter | 8.176
Beckman AU Systems | 59-92 | 7.46 | 0.21 | 4.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 8.246
Beckman AU Systems | 87 | 7.16 | 0.01 | 3.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Beckman Coulter | 7.654
Beckman AU Systems | 87 | 5.27 | 0.08 | 4.3 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Beckman Coulter | 8.508
Beckman AU Systems | 87 | 8.32 | -0.1 | 4.7 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Beckman Coulter | 9.112
Beckman AU Systems | 87 | 9.66 | -0.13 | 4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Beckman Coulter | 7.71
Beckman AU Systems | 87 | 6.16 | 0.04 | 3.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Beckman Coulter | 6.704
Beckman AU Systems | 62 | 6.33 | 0.02 | 3.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Beckman Coulter | 7.272
Beckman AU Systems | 62 | 9.04 | -0.07 | 4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Beckman Coulter | 7.77
Beckman AU Systems | 62 | 5.37 | 0.06 | 3.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Beckman Coulter | 6.136
Beckman AU Systems | 62 | 8.34 | -0.01 | 4.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Beckman Coulter | 8.614
Beckman AU Systems | 62 | 7.59 | -0.03 | 3.8 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Beckman Coulter | 7.418
Beckman AU Systems | 63 | 7.51 | 0.09 | 3.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Beckman Coulter | 6.362
Beckman AU Systems | 63 | 5.18 | -0.02 | 2.8 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Beckman Coulter | 5.468
Beckman AU Systems | 63 | 8.3 | -0.01 | 2.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Beckman Coulter | 5.674
Beckman AU Systems | 63 | 10.35 | 0.02 | 2.8 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Beckman Coulter | 5.508
Beckman AU Systems | 63 | 5.9 | 0.03 | 3.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Beckman Coulter | 6.498
Beckman AU Systems | 91 | 7.26 | 0.01 | 3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 5.89
Beckman AU Systems | 91 | 6.83 | 0.02 | 3.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Beckman Coulter | 7.076
Beckman AU Systems | 91 | 8.55 | -0.1 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Beckman Coulter | 4.408
Beckman AU Systems | 91 | 9.43 | -0.07 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Beckman Coulter | 4.438
Beckman AU Systems | 91 | 5.43 | -0.02 | 2.9 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Beckman Coulter | 5.664
Beckman AU Systems | 10 | 5.23 | -0.02 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 5.076
Beckman AU Systems | 10 | 6.4 | -0.01 | 3.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 7.242
Beckman AU Systems | 10 |  |  |  | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU Systems | 10 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU Systems | 10 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU Systems | 15 | 7.84 | -0.13 | 1.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 3.398
Beckman AU Systems | 15 | 5.78 | -0.11 | 2.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | 4.202
Beckman AU Systems | 15 | 5.03 | -0.1 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 4.604
Beckman AU Systems | 15 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 0
Beckman AU Systems | 15 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 0
Beckman AU Systems | 13 | 7.67 | -0.14 | 3.7 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 7.112
Beckman AU Systems | 13 | 5.87 | -0.04 | 3.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 6.232
Beckman AU Systems | 13 | 8.48 | -0.21 | 5.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 9.786
Beckman AU Systems | 13 | 5.82 | -0.04 | 4.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 8.192
Beckman AU Systems | 13 | 9.33 | -0.23 | 4.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | 8.786
Beckman AU Systems | 24 | 5.4 | -0.08 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 5.408
Beckman AU Systems | 24 | 8.15 | -0.11 | 2.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 5.574
Beckman AU Systems | 24 | 5.1 | -0.05 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 5.634
Beckman AU Systems | 24 | 9.88 | -0.17 | 3.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 6.69
Beckman AU Systems | 24 | 5.92 | -0.07 | 4.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 7.966
Beckman AU Systems | 28 | 8.84 | -0.02 | 3.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 7.036
Beckman AU Systems | 28 | 6.8 | -0.04 | 2.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 5.448
Beckman AU Systems | 28 | 9.17 | -0.22 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 4.68
Beckman AU Systems | 28 | 6.07 | -0.06 | 3.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 6.8
Beckman AU Systems | 28 | 7.45 | -0.07 | 2.9 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 5.614
Beckman AU Systems | 19 | 5.71 | -0.09 | 2.3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 4.418
Beckman AU Systems | 19 | 8.87 | -0.25 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 4.258
Beckman AU Systems | 19 | 5.42 | -0.05 | 2.4 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 4.654
Beckman AU Systems | 19 | 7.19 | -0.06 | 2.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 4.644
Beckman AU Systems | 19 | 9.03 | -0.23 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 4.474
Beckman AU Systems | 27 | 5.11 | -0.1 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 5.584
Beckman AU Systems | 27 | 6.61 | -0.07 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 4.634
Beckman AU Systems | 27 | 8.06 |  | 3.8 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 7.448
Beckman AU Systems | 27 | 5.96 | -0.02 | 2.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 4.292
Beckman AU Systems | 27 | 9.89 | -0.23 | 1.8 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 3.298
Beckman AU Systems | 17 | 5.26 | 0.06 | 5.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 10.448
Beckman AU Systems | 17 | 8.65 | -0.05 | 5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 9.75
Beckman AU Systems | 17 | 6.21 | -0.09 | 4.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 9.318
Beckman AU Systems | 19 | 8.41 | -0.09 | 3.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 6.574
Beckman AU Systems | 19 | 5.43 | 0.03 | 3.2 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 6.302
Beckman AU Systems | 19 | 6.43 | 0.03 | 3.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 6.106
Beckman AU Systems | 37 | 5.34 | -0.06 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Beckman Coulter | 6.408
Beckman AU Systems | 37 | 8.15 | -0.15 | 3.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Beckman Coulter | 7.494
Beckman AU Systems | 37 | 6.66 | 1.01 | 3.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Beckman Coulter | 8.066
Beckman AU Systems | 24 | 5.81 | -0.49 | 4.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 8.134
Beckman AU Systems | 25 | 6.3 | 0 | 5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 9.8
Beckman AU Systems | 24 | 6.94 | -0.66 | 4.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 8.944
Beckman AU Systems | 25 | 7.64 | 0.04 | 4.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 8.86
Beckman AU Systems | 24 | 8.22 | -0.98 | 6.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 11.172
Beckman AU Systems | 25 | 9.22 | 0.02 | 5.5 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 10.8
Beckman AU Systems | 35 | 5.59 | -0.01 | 4.2 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Beckman Coulter | 8.222
Beckman AU Systems | 35 | 9.31 | -0.09 | 3.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Beckman Coulter | 6.966
Beckman AU Systems | 35 | 7.2 | 0 | 3.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Beckman Coulter | 6.664
Beckman AU Systems | 37 | 6.92 | -0.19 | 5.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 10.59
Beckman AU Systems | 37 | 9.16 | -0.16 | 4.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 8.856
Beckman AU Systems | 37 | 5.89 | -0.18 | 5 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 9.62
Beckman AU Systems | 45 | 7.76 | -0.29 | 3.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Beckman Coulter | 6.766
Beckman AU Systems | 45 | 5.15 | -0.15 | 3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Beckman Coulter | 5.73
Beckman AU Systems | 45 | 5.89 | -0.25 | 4.4 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Beckman Coulter | 8.374
Beckman AU Systems - Beckman reagent | 70 | 8.91 | 0.05 | 2.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 5.342
Beckman AU Systems - Beckman reagent | 70 | 6.92 | 0.08 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 4.588
Beckman AU Systems - Beckman reagent | 70 | 9.33 | -0.06 | 3.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 6.016
Beckman AU Systems - Beckman reagent | 70 | 6.16 | 0.03 | 2.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 4.734
Beckman AU Systems - Beckman reagent | 70 | 7.56 | 0.04 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 4.744
Beckman AU Systems - Beckman reagent | 89 | 9.04 | -0.33 | 3.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Beckman Coulter | 6.922
Beckman AU Systems - Beckman reagent | 89 | 5.03 | 0.11 | 3.4 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Beckman Coulter | 6.774
Beckman AU Systems - Beckman reagent | 89 | 6.07 | 0.03 | 3.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Beckman Coulter | 7.086
Beckman AU Systems - Beckman reagent | 89 | 5.36 | 0.08 | 4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Beckman Coulter | 7.92
Beckman AU Systems - Beckman reagent | 89 | 7.88 | -0.2 | 4.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Beckman Coulter | 8.032
Beckman AU Systems - Beckman reagent | 87 | 5.52 | 0.06 | 4.3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Beckman Coulter | 8.488
Beckman AU Systems - Beckman reagent | 87 | 5.73 | 0.07 | 3.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Beckman Coulter | 7.322
Beckman AU Systems - Beckman reagent | 87 | 9.01 | -0.3 | 3.5 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Beckman Coulter | 6.56
Beckman AU Systems - Beckman reagent | 87 | 5.29 | 0.01 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Beckman Coulter | 6.478
Beckman AU Systems - Beckman reagent | 87 | 7.34 | -0.07 | 3.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Beckman Coulter | 7.378
Beckman AU Systems - Beckman reagent | 53 | 5.81 | 0.01 | 2.8 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 5.498
Beckman AU Systems - Beckman reagent | 53 | 9.13 | 0.01 | 3.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 7.654
Beckman AU Systems - Beckman reagent | 53 | 5.5 | 0.03 | 2.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 4.93
Beckman AU Systems - Beckman reagent | 53 | 7.3 | 0.05 | 3.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 6.126
Beckman AU Systems - Beckman reagent | 53 | 9.28 | 0.02 | 3.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 7.664
Beckman AU Systems - Beckman reagent | 66 | 5.19 | -0.02 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 5.664
Beckman AU Systems - Beckman reagent | 66 | 6.59 | -0.09 | 3.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 6.182
Beckman AU Systems - Beckman reagent | 66 | 8.11 |  | 3.1 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 6.076
Beckman AU Systems - Beckman reagent | 66 | 5.93 | -0.05 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 4.654
Beckman AU Systems - Beckman reagent | 66 | 9.78 | -0.34 | 3.4 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 6.324
Beckman AU Systems - Beckman reagent | 64 | 6.33 | -0.08 | 3.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 7.564
Beckman AU Systems - Beckman reagent | 64 | 8.42 | -0.18 | 4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Beckman Coulter | 7.66
Beckman AU Systems - Beckman reagent | 64 | 5.32 | -0.1 | 3.8 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Beckman Coulter | 7.348
Beckman AU Systems - Beckman reagent | 64 | 7.29 | -0.09 | 3.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Beckman Coulter | 7.358
Beckman AU Systems - Beckman reagent | 64 | 9.42 | -0.33 | 5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Beckman Coulter | 9.47
Beckman AU Systems - Beckman reagent | 78 | 5.2 | -0.05 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 5.83
Beckman AU Systems - Beckman reagent | 78 | 6.41 | 0 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 5.292
Beckman AU Systems - Beckman reagent | 78 | 8.07 | -0.14 | 3.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 7.308
Beckman AU Systems - Beckman reagent | 78 | 5.42 | -0.05 | 3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 5.83
Beckman AU Systems - Beckman reagent | 78 | 9.3 | -0.18 | 4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 7.66
Beckman AU Systems - Beckman reagent | 72 | 8.08 | 0.11 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 5.402
Beckman AU Systems - Beckman reagent | 72 | 5.99 | 0.1 | 2.8 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | 5.588
Beckman AU Systems - Beckman reagent | 72 | 5.2 | 0.07 | 2.7 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 5.362
Beckman AU Systems - Beckman reagent | 72 | 7.51 | 0.11 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 5.99
Beckman AU Systems - Beckman reagent | 72 | 9.13 | -0.04 | 2.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 5.056
Beckman AU Systems - Beckman reagent | 58 | 7.84 | 0.03 | 3.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 6.106
Beckman AU Systems - Beckman reagent | 58 | 5.98 | 0.07 | 3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 5.95
Beckman AU Systems - Beckman reagent | 58 | 8.68 | -0.01 | 2.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 5.674
Beckman AU Systems - Beckman reagent | 58 | 5.96 | 0.1 | 3.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 7.548
Beckman AU Systems - Beckman reagent | 58 | 9.47 | -0.09 | 3 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | 5.79
Beckman AU Systems - Beckman reagent | 72 | 5.52 | 0.04 | 4.4 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 8.664
Beckman AU Systems - Beckman reagent | 72 | 8.23 | -0.03 | 3.1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 6.046
Beckman AU Systems - Beckman reagent | 72 | 5.21 | 0.06 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 6.528
Beckman AU Systems - Beckman reagent | 72 | 9.97 | -0.08 | 2.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 5.212
Beckman AU Systems - Beckman reagent | 72 | 6.01 | 0.02 | 3.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 6.488
Beckman Coulter AU Systems | 26 | 5.9381 | -0.9 | 5.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Beckman Coulter | 10.076
Beckman Coulter AU Systems | 26 | 7.4204 | -0.3 | 5.5 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Beckman Coulter | 10.48
Beckman Coulter AU Systems | 32 | 6.2492 | 0.6 | 7.2 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Beckman Coulter | 14.712
Beckman Coulter AU Systems | 33 | 7.4021 | -0.6 | 6.1 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Beckman Coulter | 11.356
Beckman Coulter AU Systems | 7 | 6.1486 | 1.4 | 7.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Beckman Coulter | 16.884
Beckman Coulter AU Systems | 6 | 7.6125 | 1.8 | 5 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Beckman Coulter | 11.6
Beckman Coulter AU Systems | 9 | 6.6518 | 5 | 7.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Beckman Coulter | 18.916
Beckman Coulter AU Systems | 9 | 7.7681 | 3.4 | 4.8 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Beckman Coulter | 12.808
Beckman Coulter AU Systems | 13 | 7.7132 | 2.7 | 7.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Beckman Coulter | 17.988
Beckman Coulter AU Systems | 11 | 6.2126 | 2.2 | 7.1 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Beckman Coulter | 16.116
Beckman Coulter AU Systems | 58 | 7.4204 | -0.5 | 6.7 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Beckman Coulter | 12.632
Beckman Coulter AU Systems | 59 | 6.0662 | 0.6 | 6.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Beckman Coulter | 14.124
Beckman Coulter AU Systems | 13 | 9.1589 | 6.5 | 5.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Beckman Coulter | 17.476
Beckman Coulter AU Systems | 14 | 5.9656 | 3.5 | 6.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Beckman Coulter | 16.436
Beckman Coulter AU Systems | 57 | 8.4635 | -1.1 | 6.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Beckman Coulter | 11.248
Beckman Coulter AU Systems | 57 | 5.6544 | 0.1 | 6.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Beckman Coulter | 13.036
Beckman Coulter AU Systems | 46 | 5.7459 | 2.8 | 8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Beckman Coulter | 18.48
Beckman Coulter AU Systems | 46 | 5.7276 | 2.6 | 7.9 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Beckman Coulter | 18.084
Beckman Coulter AU Systems | 66 | 5.5812 | 1 | 6 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 12.76
Beckman Coulter AU Systems | 66 | 5.6178 | 1.4 | 6.7 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 14.532
Beckman Coulter not specified/other | 14 | 5.6087 | 1.3 | 5.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 12.08
Beckman Coulter not specified/other | 14 | 5.5904 | 1.1 | 6.5 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 13.84
Beckman Coulter not specified/other | 6 | 7.4295 | -0.4 | 2.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Beckman Coulter | 5.088
Beckman Coulter not specified/other | 6 | 6.0662 | 0.6 | 1.5 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Beckman Coulter | 3.54
Beckman Coulter UniCel DxC | 15 | 5.8741 | -1.6 | 3.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Beckman Coulter | 6.044
Beckman Coulter UniCel DxC | 15 | 7.4021 | -0.5 | 3.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Beckman Coulter | 5.576
Beckman Coulter UniCel DxC | 17 | 6.2675 | 0.8 | 5.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Beckman Coulter | 10.796
Beckman Coulter UniCel DxC | 17 | 7.5668 | 1.2 | 5 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Beckman Coulter | 11
Beckman Coulter UniCel DxC | 27 | 7.457 | -0.1 | 3.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Beckman Coulter | 6.76
Beckman Coulter UniCel DxC | 27 | 6.0022 | -0.1 | 3.6 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Beckman Coulter | 6.956
Beckman Coulter UniCel DxC | 19 | 8.6099 | 0.5 | 3.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Beckman Coulter | 6.968
Beckman Coulter UniCel DxC | 19 | 5.5995 | -0.5 | 3.2 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Beckman Coulter | 5.772
Beckman Coulter UniCel DxC | 8 | 5.5446 | 0.6 | 3.8 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 8.048
Beckman Coulter UniCel DxC | 8 | 5.5172 | 0.3 | 3.8 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Beckman Coulter | 7.748
Beckman Synchron CX System | 32 | 5 |  | 3.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Beckman Coulter | 6.664
Beckman Synchron CX System | 32 | 7.9 |  | 4.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Beckman Coulter | 8.624
Beckman Synchron CX System | 32 | 5.8 |  | 5.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Beckman Coulter | 10.388
Beckman Synchron CX System | 27 | 6.45 | -0.15 | 5.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Beckman Coulter | 9.846
Beckman Synchron CX System | 27 | 9.21 | -0.29 | 5.6 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Beckman Coulter | 10.686
Beckman Synchron CX System | 27 | 7.18 | -0.22 | 5.5 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Beckman Coulter | 10.56
Beckman Synchron CX System | 24 | 5.79 | -0.11 | 5.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Beckman Coulter | 10.082
Beckman Synchron CX System | 24 | 9.51 | -0.29 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Beckman Coulter | 8.138
Beckman Synchron CX System | 24 | 7.11 | -0.29 | 6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Beckman Coulter | 11.47
Beckman Synchron CX System | 16 | 5.25 | 0.05 | 4 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 7.89
Beckman Synchron CX System | 16 | 8.15 | -0.55 | 4.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 7.486
Beckman Synchron CX System | 16 | 6.12 | -0.18 | 3.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 6.876
Beckman Synchron CX System | 12 | 8.43 | -0.07 | 3.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 7.574
Beckman Synchron CX System | 12 | 5.42 | 0.02 | 7.2 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 14.132
Beckman Synchron CX System | 12 | 6.48 | 0.08 | 7.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 14.192
Beckman Synchron LX System | 85 | 5 |  | 4.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Beckman Coulter | 9.604
Beckman Synchron LX System | 85 | 8.2 |  | 3.7 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Beckman Coulter | 7.252
Beckman Synchron LX System | 85 | 5.8 |  | 3.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Beckman Coulter | 7.448
Beckman Synchron LX System | 87 | 6.49 | -0.11 | 3.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Beckman Coulter | 7.338
Beckman Synchron LX System | 87 | 9.23 | -0.27 | 3.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Beckman Coulter | 6.59
Beckman Synchron LX System | 87 | 7.1 | -0.3 | 3.9 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Beckman Coulter | 7.344
Beckman Synchron LX System | 69 | 5.83 | -0.07 | 4.4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Beckman Coulter | 8.554
Beckman Synchron LX System | 69 | 9.71 | -0.09 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Beckman Coulter | 8.338
Beckman Synchron LX System | 69 | 7.16 | -0.24 | 3.6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Beckman Coulter | 6.816
Beckman Synchron LX System | 59 | 5.42 | 0.22 | 5.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 10.412
Beckman Synchron LX System | 59 | 8.61 | -0.09 | 5.3 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 10.298
Beckman Synchron LX System | 59 | 6.32 | 0.02 | 5 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 9.82
Beckman Synchron LX System | 39 | 8.55 | 0.05 | 4.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 8.478
Beckman Synchron LX System | 39 | 5.52 | 0.12 | 5.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 10.116
Beckman Synchron LX System | 39 | 6.41 | 0.01 | 4.5 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 8.83
Beckman Synchron LX System | 29 | 6.2 | -0.1 | 4.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 8.916
Beckman Synchron LX System | 29 | 7.69 | 0.09 | 3.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 7.342
Beckman Synchron LX System | 29 | 9.54 | 0.34 | 4.5 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 9.16
Beckman Synchron LX System | 21 | 5.77 | 0.17 | 4.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Beckman Coulter | 9.382
Beckman Synchron LX System | 21 | 9.45 | 0.05 | 2.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Beckman Coulter | 5.734
Beckman Synchron LX System | 21 | 7.15 | -0.05 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Beckman Coulter | 5.438
Beckman Synchron LX System | 17 | 5.49 | 0.09 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Beckman Coulter | 7.93
Beckman Synchron LX System | 17 | 8.22 | -0.08 | 3.2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Beckman Coulter | 6.192
Beckman Synchron LX System | 17 | 5.62 | -0.03 | 6.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Beckman Coulter | 12.71
Beckman Synchron LX System | 10 | 6.91 | -0.2 | 3.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 6.66
Beckman Synchron LX System | 10 | 9.41 | 0.09 | 2.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 4.402
Beckman Synchron LX System | 10 | 6.29 | 0.22 | 8.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 16.292
Beckman Synchron System | 291 | 10.6 |  | 4.8 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Beckman Coulter | 9.408
Beckman Synchron System | 291 | 5.2 |  | 5.4 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Beckman Coulter | 10.584
Beckman Synchron System | 291 | 7.9 |  | 4.8 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Beckman Coulter | 9.408
Beckman Synchron System | 310 | 6.6 |  | 4.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Beckman Coulter | 8.624
Beckman Synchron System | 310 | 6.6 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Beckman Coulter | 8.428
Beckman Synchron System | 310 | 11.4 |  | 4.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Beckman Coulter | 9.212
Beckman Synchron System | 310 | 6.6 |  | 4.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Beckman Coulter | 8.624
Beckman Synchron System | 310 | 6.6 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Beckman Coulter | 8.428
Beckman Synchron System | 310 | 11.4 |  | 4.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Beckman Coulter | 9.212
Beckman Synchron System | 385 | 5.3 |  | 4.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Beckman Coulter | 9.016
Beckman Synchron System | 385 | 11.7 |  | 5.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Beckman Coulter | 10.584
Beckman Synchron System | 385 | 7.3 |  | 4.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Beckman Coulter | 9.016
Beckman Synchron System | 353 | 9 |  | 4.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Beckman Coulter | 9.212
Beckman Synchron System | 353 | 11.2 |  | 4.4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Beckman Coulter | 8.624
Beckman Synchron System | 353 | 6 |  | 4.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Beckman Coulter | 8.232
Beckman Synchron System | 369 | 13.2 |  | 4.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Beckman Coulter | 9.016
Beckman Synchron System | 369 | 5.3 |  | 4.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Beckman Coulter | 9.408
Beckman Synchron System | 369 | 8.7 |  | 4.5 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Beckman Coulter | 8.82
Beckman Synchron System | 389 | 11.8 |  | 4.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Beckman Coulter | 8.82
Beckman Synchron System | 389 | 6.5 |  | 4 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Beckman Coulter | 7.84
Beckman Synchron System | 389 | 5 |  | 4.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Beckman Coulter | 9.212
Beckman Coulter UniCel DxC Synchron Systems | 172 | 6.57 | -0.01 | 2.7 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Beckman Coulter | 5.282
Beckman Coulter UniCel DxC Synchron Systems | 172 | 8.49 | 0.1 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Beckman Coulter | 5.588
Beckman Coulter UniCel DxC Synchron Systems | 172 | 5.61 | -0.04 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Beckman Coulter | 5.056
Beckman Coulter UniCel DxC Synchron Systems | 191 | 6.5 | 0.01 | 3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Beckman Coulter | 5.89
Beckman Coulter UniCel DxC Synchron Systems | 191 | 6.95 | -0.02 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Beckman Coulter | 5.076
Beckman Coulter UniCel DxC Synchron Systems | 191 | 9.66 | 0.01 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Beckman Coulter | 5.498
Beckman Coulter UniCel DxC Synchron Systems | 149 | 6.78 | -0.01 | 2.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Beckman Coulter | 5.086
Beckman Coulter UniCel DxC Synchron Systems | 149 | 10.37 | 0.09 | 3.1 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Beckman Coulter | 6.166
Beckman Coulter UniCel DxC Synchron Systems | 149 | 6.77 | -0.05 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Beckman Coulter | 4.85
Beckman Coulter UniCel DxC Synchron Systems | 114 | 8.7 | 0.07 | 2.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Beckman Coulter | 5.166
Beckman Coulter UniCel DxC Synchron Systems | 114 | 5.31 | -0.01 | 2.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Beckman Coulter | 4.89
Beckman Coulter UniCel DxC Synchron Systems | 117 | 5.78 | -0.06 | 2.4 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Beckman Coulter | 4.644
Beckman Coulter UniCel DxC Synchron Systems | 117 | 11.87 | 0.16 | 3.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Beckman Coulter | 6.628
Beckman Coulter UniCel DxC Synchron Systems | 117 | 9.62 | 0.09 | 2.8 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Beckman Coulter | 5.578
Beckman Coulter UniCel DxC Synchron Systems | 117 | 5.05 | 0.01 | 2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Beckman Coulter | 3.93
Beckman Coulter UniCel DxC Synchron Systems | 117 | 7.32 | -0.06 | 2.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Beckman Coulter | 4.84
Beckman Coulter UniCel DxC Synchron Systems | 119-157 | 11.87 | 0.18 | 3.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Beckman Coulter | 7.04
Beckman Coulter UniCel DxC Synchron Systems | 119-157 | 5.89 | -0.04 | 2.4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Beckman Coulter | 4.664
Beckman Coulter UniCel DxC Synchron Systems | 119-157 | 5.19 | 0.02 | 2.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Beckman Coulter | 4.528
Beckman Coulter UniCel DxC Synchron Systems | 119-157 | 8.17 | 0.03 | 2.2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Beckman Coulter | 4.342
Beckman Coulter UniCel DxC Synchron Systems | 119-157 | 9.41 | 0.11 | 3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Beckman Coulter | 5.99
Beckman Coulter UniCel DxC Synchron Systems | 111-142 | 5.31 | -0.01 | 2.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Beckman Coulter | 5.086
Beckman Coulter UniCel DxC Synchron Systems | 111-142 | 9.35 | 0.19 | 2.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Beckman Coulter | 4.894
Beckman Coulter UniCel DxC Synchron Systems | 111-142 | 5.28 | -0.03 | 2.4 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Beckman Coulter | 4.674
Beckman Coulter UniCel DxC Synchron Systems | 111-142 | 12.27 | 0.24 | 3.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Beckman Coulter | 6.316
Beckman Coulter UniCel DxC Synchron Systems | 111-142 | 5.86 | -0.08 | 2.8 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Beckman Coulter | 5.408
Beckman Coulter UniCel DxC Synchron Systems | 108 | 5.23 | -0.04 | 1.5 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Beckman Coulter | 2.9
Beckman Coulter UniCel DxC Synchron Systems | 108 | 10.68 | 0.09 | 2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Beckman Coulter | 4.01
Beckman Coulter UniCel DxC Synchron Systems | 108 | 6.07 | -0.13 | 1.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Beckman Coulter | 3.006
Beckman Coulter UniCel DxC Synchron Systems | 108 | 12.4 | 0.17 | 2.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Beckman Coulter | 4.482
Beckman Coulter UniCel DxC Synchron Systems | 108 | 7.5 | -0.01 | 1.8 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Beckman Coulter | 3.518
Beckman Coulter UniCel DxC Synchron Systems | 101-133 | 9.21 | 0.1 | 2.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Beckman Coulter | 4.804
Beckman Coulter UniCel DxC Synchron Systems | 101-133 | 5.93 | -0.08 | 2.2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Beckman Coulter | 4.232
Beckman Coulter UniCel DxC Synchron Systems | 101-133 | 11.83 | 0.12 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Beckman Coulter | 5.216
Beckman Coulter UniCel DxC Synchron Systems | 101-133 | 5.02 | 0 | 2.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Beckman Coulter | 4.704
Beckman Coulter UniCel DxC Synchron Systems | 101-133 | 7.53 | -0.05 | 1.9 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Beckman Coulter | 3.674
Beckman Coulter UniCel DxC Synchron Systems | 108-136 | 6.44 | 0.03 | 3.3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Beckman Coulter | 6.498
Beckman Coulter UniCel DxC Synchron Systems | 108-136 | 9.68 | 0.15 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Beckman Coulter | 5.638
Beckman Coulter UniCel DxC Synchron Systems | 108-136 | 5.33 | -0.01 | 3.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Beckman Coulter | 6.066
Beckman Coulter UniCel DxC Synchron Systems | 108-136 | 8.59 | 0.08 | 2.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Beckman Coulter | 5.372
Beckman Coulter UniCel DxC Synchron Systems | 108-136 | 7.27 | 0.02 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 4.528
Beckman Coulter UniCel DxC Synchron Systems | 111 | 7.24 | 0.09 | 3.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Beckman Coulter | 6.558
Beckman Coulter UniCel DxC Synchron Systems | 111 | 5.27 | 0.08 | 3.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Beckman Coulter | 7.332
Beckman Coulter UniCel DxC Synchron Systems | 111 | 8.5 | 0.08 | 2.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Beckman Coulter | 5.764
Beckman Coulter UniCel DxC Synchron Systems | 111 | 9.95 | 0.16 | 2.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Beckman Coulter | 4.668
Beckman Coulter UniCel DxC Synchron Systems | 111 | 6.12 | 0 | 2.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Beckman Coulter | 4.9
Beckman Coulter UniCel DxC Synchron Systems | 86 | 6.34 | 0.03 | 2.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Beckman Coulter | 4.93
Beckman Coulter UniCel DxC Synchron Systems | 86 | 9.23 | 0.12 | 2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Beckman Coulter | 4.04
Beckman Coulter UniCel DxC Synchron Systems | 86 | 5.27 | -0.04 | 2.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Beckman Coulter | 4.468
Beckman Coulter UniCel DxC Synchron Systems | 86 | 8.44 | 0.09 | 2.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Beckman Coulter | 4.206
Beckman Coulter UniCel DxC Synchron Systems | 86 | 7.67 | 0.05 | 2.6 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Beckman Coulter | 5.146
Beckman Coulter UniCel DxC Synchron Systems | 103 | 7.51 | 0.09 | 2.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Beckman Coulter | 5.774
Beckman Coulter UniCel DxC Synchron Systems | 103 | 5.21 | 0.01 | 3.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Beckman Coulter | 6.282
Beckman Coulter UniCel DxC Synchron Systems | 103 | 8.45 | 0.14 | 3 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Beckman Coulter | 6.02
Beckman Coulter UniCel DxC Synchron Systems | 103 | 10.63 | 0.3 | 3.2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Beckman Coulter | 6.572
Beckman Coulter UniCel DxC Synchron Systems | 103 | 5.88 | 0.01 | 2.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Beckman Coulter | 4.714
Beckman Coulter UniCel DxC Synchron Systems | 118 | 7.33 | 0.08 | 3.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 6.156
Beckman Coulter UniCel DxC Synchron Systems | 118 | 6.84 | 0.03 | 3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Beckman Coulter | 5.91
Beckman Coulter UniCel DxC Synchron Systems | 118 | 8.77 | 0.12 | 2.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Beckman Coulter | 5.216
Beckman Coulter UniCel DxC Synchron Systems | 118 | 9.78 | 0.28 | 2.7 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Beckman Coulter | 5.572
Beckman Coulter UniCel DxC Synchron Systems | 118 | 5.45 | 0 | 2.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Beckman Coulter | 4.508
Beckman Coulter UniCel DxC Synchron Systems | 233 | 7.01 | -0.1 | 3.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 6.76
Beckman Coulter UniCel DxC Synchron Systems | 233 | 9.45 | 0.13 | 3.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 7.186
Beckman Coulter UniCel DxC Synchron Systems | 233 | 6.06 | -0.01 | 4.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 8.222
Beckman Coulter UniCel DxC Synchron Systems | 187 | 7.97 | -0.08 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Beckman Coulter | 5.8
Beckman Coulter UniCel DxC Synchron Systems | 187 | 5.3 | 0 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Beckman Coulter | 6.076
Beckman Coulter UniCel DxC Synchron Systems | 187 | 6.09 | -0.05 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Beckman Coulter | 6.026
Beckman Coulter UniCel DxC Synchron Systems | 143 | 7.01 | -0.1 | 3.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 6.956
Beckman Coulter UniCel DxC Synchron Systems | 143 | 9.46 | 0.14 | 3.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 7.588
Beckman Coulter UniCel DxC Synchron Systems | 143 | 6.05 | -0.02 | 4.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 8.212
Beckman Coulter UniCel DxC Synchron Systems | 68 | 8.92 | 0.06 | 2.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 5.156
Beckman Coulter UniCel DxC Synchron Systems | 68 | 6.83 | -0.01 | 2.4 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 4.694
Beckman Coulter UniCel DxC Synchron Systems | 68 | 9.35 | -0.04 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 4.86
Beckman Coulter UniCel DxC Synchron Systems | 68 | 6.1 | -0.03 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 3.498
Beckman Coulter UniCel DxC Synchron Systems | 68 | 7.46 | -0.06 | 2.3 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 4.448
Beckman Coulter UniCel DxC Synchron Systems | 93 | 9.35 | -0.02 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Beckman Coulter | 4.096
Beckman Coulter UniCel DxC Synchron Systems | 93 | 5.02 | 0.1 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Beckman Coulter | 5.392
Beckman Coulter UniCel DxC Synchron Systems | 93 | 6 | -0.04 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Beckman Coulter | 4.468
Beckman Coulter UniCel DxC Synchron Systems | 93 | 5.27 | -0.01 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Beckman Coulter | 4.694
Beckman Coulter UniCel DxC Synchron Systems | 93 | 8.05 | -0.03 | 2.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Beckman Coulter | 4.478
Beckman Coulter UniCel DxC Synchron Systems | 84 | 5.6 | 0.14 | 2.7 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Beckman Coulter | 5.432
Beckman Coulter UniCel DxC Synchron Systems | 84 | 5.7 | 0.04 | 2.5 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Beckman Coulter | 4.94
Beckman Coulter UniCel DxC Synchron Systems | 84 | 9.39 | 0.08 | 2.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Beckman Coulter | 5.372
Beckman Coulter UniCel DxC Synchron Systems | 84 | 5.34 | 0.06 | 2.9 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Beckman Coulter | 5.744
Beckman Coulter UniCel DxC Synchron Systems | 84 | 7.32 | -0.09 | 2.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Beckman Coulter | 5.398
Beckman Coulter UniCel DxC Synchron Systems | 81 | 5.24 | -0.01 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 5.87
Beckman Coulter UniCel DxC Synchron Systems | 81 | 6.32 | -0.09 | 2.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 5.006
Beckman Coulter UniCel DxC Synchron Systems | 81 | 8.08 | -0.13 | 2.2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 4.182
Beckman Coulter UniCel DxC Synchron Systems | 81 | 5.39 | -0.08 | 2.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 5.212
Beckman Coulter UniCel DxC Synchron Systems | 81 | 9.39 | -0.09 | 3.1 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 5.986
Beckman Coulter UniCel DxC Synchron Systems | 73 | 7.97 | 0 | 3 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 5.88
Beckman Coulter UniCel DxC Synchron Systems | 73 | 5.81 | -0.08 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | 3.84
Beckman Coulter UniCel DxC Synchron Systems | 73 | 5.1 | -0.03 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 4.478
Beckman Coulter UniCel DxC Synchron Systems | 73 | 7.29 | -0.11 | 2.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 4.79
Beckman Coulter UniCel DxC Synchron Systems | 73 | 9.11 | -0.06 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 5.232
Beckman Coulter UniCel DxC Synchron Systems | 60 | 7.78 | -0.03 | 2.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 4.674
Beckman Coulter UniCel DxC Synchron Systems | 60 | 5.82 | -0.09 | 2.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 5.202
Beckman Coulter UniCel DxC Synchron Systems | 60 | 8.7 | 0.01 | 3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 5.89
Beckman Coulter UniCel DxC Synchron Systems | 60 | 5.86 | 0 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 4.9
Beckman Coulter UniCel DxC Synchron Systems | 60 | 9.56 | 0 | 2.8 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | 5.488
Beckman Coulter UniCel DxC Synchron Systems | 68 | 5.5 | 0.02 | 3.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 6.88
Beckman Coulter UniCel DxC Synchron Systems | 68 | 8.24 | -0.02 | 3.1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 6.056
Beckman Coulter UniCel DxC Synchron Systems | 68 | 5.14 | -0.01 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 5.674
Beckman Coulter UniCel DxC Synchron Systems | 68 | 10.12 | 0.07 | 2.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 5.362
Beckman Coulter UniCel DxC Synchron Systems | 68 | 5.93 | -0.06 | 3.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 6.016
Beckman Coulter UniCel DxC Synchron Systems | 59 | 5.88 | 0.08 | 3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 5.96
Beckman Coulter UniCel DxC Synchron Systems | 59 | 9.08 | -0.04 | 2.4 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 4.664
Beckman Coulter UniCel DxC Synchron Systems | 59 | 5.47 | 0 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 4.508
Beckman Coulter UniCel DxC Synchron Systems | 59 | 7.24 | -0.01 | 2.2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 4.302
Beckman Coulter UniCel DxC Synchron Systems | 59 | 9.25 | -0.01 | 2.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 4.89
Beckman Coulter UniCel DxC Synchron Systems | 65 | 5.28 | 0.07 | 3.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 6.146
Beckman Coulter UniCel DxC Synchron Systems | 65 | 6.66 | -0.02 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 4.684
Beckman Coulter UniCel DxC Synchron Systems | 65 | 8.32 |  | 2.7 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 5.292
Beckman Coulter UniCel DxC Synchron Systems | 65 | 5.98 | 0 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 4.116
Beckman Coulter UniCel DxC Synchron Systems | 65 | 10.3 | 0.18 | 2.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 5.276
Beckman Coulter UniCel DxC Synchron Systems | 69 | 6.32 | -0.09 | 3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 5.79
Beckman Coulter UniCel DxC Synchron Systems | 69 | 8.58 | -0.02 | 2.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Beckman Coulter | 4.88
Beckman Coulter UniCel DxC Synchron Systems | 69 | 5.39 | -0.03 | 3.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Beckman Coulter | 7.222
Beckman Coulter UniCel DxC Synchron Systems | 69 | 7.29 | -0.09 | 3.1 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Beckman Coulter | 5.986
Beckman Coulter UniCel DxC Synchron Systems | 69 | 9.81 | 0.06 | 3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Beckman Coulter | 5.94
JEOL BioMajesty JCA-BM series | 7 | 7.5393 | 0.8 | 10.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | BioMajesty | 21.968
JEOL BioMajesty JCA-BM series | 7 | 5.8741 | -1.5 | 13.7 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | BioMajesty | 25.352
JEOL BioMajesty JCA-BM series | 7 | 8.0152 | -6 | 6.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | BioMajesty | 7.524
JEOL BioMajesty JCA-BM series | 6 | 5.3617 | -3.1 | 7.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | BioMajesty | 11.796
Bio-Rad D-10 | 106 | 10.9 |  | 2.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad D-10 | 106 | 5.3 |  | 2.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad D-10 | 106 | 8.5 |  | 2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 3.92
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad D-10 | 130 | 12.3 |  | 2.4 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad D-10 | 130 | 12.3 |  | 2.4 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad D-10 | 169 | 5.5 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad D-10 | 169 | 12.6 |  | 2.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad D-10 | 169 | 8.1 |  | 2.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad D-10 | 167 | 9.6 |  | 2.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad D-10 | 167 | 11.5 |  | 2.6 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad D-10 | 167 | 6.3 |  | 2.5 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad D-10 | 177 | 13.2 |  | 2.4 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad D-10 | 177 | 5.3 |  | 3.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 6.076
Bio-Rad D-10 | 177 | 9 |  | 2.3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad D-10 | 200 | 12.1 |  | 2.8 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad D-10 | 200 | 6.9 |  | 2.9 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad D-10 | 200 | 5 |  | 3.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 7.644
Bio-Rad D-10 | 191 | 5.1 |  | 3.2 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad D-10 | 191 | 8.5 |  | 2.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad D-10 | 191 | 6.1 |  | 3 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad D-10 | 209 | 6.74 | 0.14 | 2.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 5.824
Bio-Rad D-10 | 209 | 9.54 | 0.04 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 5.528
Bio-Rad D-10 | 209 | 7.6 | 0.2 | 3.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 6.472
Bio-Rad D-10 | 198 | 6.09 | 0.19 | 2.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 5.286
Bio-Rad D-10 | 198 | 10 | 0.2 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 5.1
Bio-Rad D-10 | 198 | 7.56 | 0.16 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 4.864
Bio-Rad D-10 | 207 | 5.29 | 0.09 | 2.8 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 5.578
Bio-Rad D-10 | 207 | 8.94 | 0.24 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 5.14
Bio-Rad D-10 | 207 | 6.45 | 0.15 | 2.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 5.442
Bio-Rad D-10 | 199 | 8.72 | 0.22 | 2.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 4.924
Bio-Rad D-10 | 199 | 5.41 | 0.01 | 2.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 5.694
Bio-Rad D-10 | 199 | 6.44 | 0.04 | 2.8 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 5.528
Bio-Rad D-10 | 215 | 6.55 | 0.25 | 2.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 5.738
Bio-Rad D-10 | 215 | 7.92 | 0.32 | 2.3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 4.828
Bio-Rad D-10 | 215 | 9.46 | 0.26 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 5.552
Bio-Rad D-10 | 244 | 5.68 | 0.08 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 5.372
Bio-Rad D-10 | 244 | 9.64 | 0.24 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 5.336
Bio-Rad D-10 | 244 | 7.55 | 0.35 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 5.642
Bio-Rad D-10 | 230 | 5.55 | 0.15 | 2.9 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 5.834
Bio-Rad D-10 | 230 | 8.5 | 0.2 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 5.1
Bio-Rad D-10 | 230 | 5.8 | 0.15 | 3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 6.03
Bio-Rad D-10 | 210 | 7.16 | 0.05 | 2.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 5.342
Bio-Rad D-10 | 210 | 9.41 | 0.09 | 2.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 5.186
Bio-Rad D-10 | 210 | 6.14 | 0.07 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 5.166
Bio-Rad D-10 | 205 | 8.18 | 0.13 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 5.422
Bio-Rad D-10 | 205 | 5.32 | 0.02 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 5.508
Bio-Rad D-10 | 205 | 6.2 | 0.06 | 2.8 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 5.548
Bio-Rad D-10 | 217 | 6.6 | 0.02 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 5.508
Bio-Rad D-10 | 217 | 8.45 | 0.06 | 2.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 4.372
Bio-Rad D-10 | 217 | 5.69 | 0.04 | 2.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 5.332
Bio-Rad D-10 | 209 | 6.58 | 0.09 | 2.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 4.598
Bio-Rad D-10 | 209 | 7.13 | 0.16 | 2.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 4.276
Bio-Rad D-10 | 209 | 9.77 | 0.12 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 4.824
Bio-Rad D-10 | 198 | 6.92 | 0.13 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 5.422
Bio-Rad D-10 | 198 | 10.36 | 0.08 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 4.588
Bio-Rad D-10 | 198 | 6.92 | 0.1 | 2.4 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 4.804
Bio-Rad D-10 | 136 | 8.72 | 0.09 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 4.598
Bio-Rad D-10 | 136 | 5.36 | 0.04 | 2.6 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 5.136
Bio-Rad D-10 | 133 | 6 | 0.16 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 5.452
Bio-Rad D-10 | 133 | 11.83 | 0.12 | 2.6 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 5.216
Bio-Rad D-10 | 133 | 9.67 | 0.14 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 5.432
Bio-Rad D-10 | 133 | 5.09 | 0.05 | 3.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 6.518
Bio-Rad D-10 | 133 | 7.43 | 0.05 | 2.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 5.146
Bio-Rad D-10 | 137-198 | 11.94 | 0.25 | 2.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 5.542
Bio-Rad D-10 | 137-198 | 5.99 | 0.06 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 4.96
Bio-Rad D-10 | 137-198 | 5.14 | -0.03 | 2.9 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 5.654
Bio-Rad D-10 | 137-198 | 8.06 | -0.08 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 5.016
Bio-Rad D-10 | 137-198 | 9.4 | 0.1 | 2.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 5.392
Bio-Rad D-10 | 127-176 | 5.42 | 0.1 | 2.8 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 5.588
Bio-Rad D-10 | 127-176 | 9.42 | 0.26 | 2.5 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 5.16
Bio-Rad D-10 | 127-176 | 5.41 | 0.1 | 2.9 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 5.784
Bio-Rad D-10 | 127-176 | 12.46 | 0.43 | 2.4 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 5.134
Bio-Rad D-10 | 127-176 | 6.11 | 0.17 | 2.4 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 4.874
Bio-Rad D-10 | 128 | 5.32 | 0.05 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 5.734
Bio-Rad D-10 | 128 | 10.83 | 0.24 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 4.944
Bio-Rad D-10 | 128 | 6.33 | 0.13 | 2.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 5.03
Bio-Rad D-10 | 128 | 12.39 | 0.16 | 2.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 4.864
Bio-Rad D-10 | 128 | 7.68 | 0.17 | 2.4 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 4.874
Bio-Rad D-10 | 125-172 | 9.16 | 0.05 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 4.95
Bio-Rad D-10 | 125-172 | 6.05 | 0.04 | 2.4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 4.744
Bio-Rad D-10 | 125-172 | 11.85 | 0.14 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 5.236
Bio-Rad D-10 | 125-172 | 5.11 | 0.09 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 4.598
Bio-Rad D-10 | 125-172 | 7.65 | 0.07 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 4.382
Bio-Rad D-10 | 120-164 | 6.53 | 0.12 | 2.4 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 4.824
Bio-Rad D-10 | 120-164 | 9.81 | 0.28 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 5.572
Bio-Rad D-10 | 120-164 | 5.36 | 0.02 | 2.6 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 5.116
Bio-Rad D-10 | 120-164 | 8.75 | 0.24 | 2.2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 4.552
Bio-Rad D-10 | 120-164 | 7.36 | 0.11 | 2.4 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 4.814
Bio-Rad D-10 | 121 | 7.58 | 0.16 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 5.06
Bio-Rad D-10 | 121 | 5.45 | 0.25 | 2.3 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 4.758
Bio-Rad D-10 | 121 | 8.38 | 0.07 | 1.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 3.794
Bio-Rad D-10 | 121 | 10.39 | 0.06 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 3.784
Bio-Rad D-10 | 121 | 6.01 | 0.14 | 2.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 4.648
Bio-Rad D-10 | 156 | 7.36 | 0.11 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 4.618
Bio-Rad D-10 | 156 | 6.9 | 0.09 | 2.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 4.794
Bio-Rad D-10 | 156 | 8.71 | 0.06 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 4.176
Bio-Rad D-10 | 156 | 9.65 | 0.15 | 2.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 4.854
Bio-Rad D-10 | 156 | 5.48 | 0.03 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 5.322
Bio-Rad D-10 | 156 | 7.25 | 0.1 | 2.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 4.804
Bio-Rad D-10 | 156 | 5.2 | 0.01 | 2.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 5.498
Bio-Rad D-10 | 156 | 8.6 | 0.18 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 4.884
Bio-Rad D-10 | 156 | 9.94 | 0.15 | 2.4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 4.854
Bio-Rad D-10 | 156 | 6.25 | 0.13 | 2.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 5.226
Bio-Rad D-10 | 125 | 6.45 | 0.14 | 2.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 5.432
Bio-Rad D-10 | 125 | 9.15 | 0.04 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 5.136
Bio-Rad D-10 | 125 | 5.29 | -0.02 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 5.468
Bio-Rad D-10 | 125 | 8.5 | 0.15 | 2.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 4.854
Bio-Rad D-10 | 125 | 7.73 | 0.11 | 2.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 5.01
Bio-Rad D-10 | 153 | 9.45 | 0.08 | 3 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 5.96
Bio-Rad D-10 | 153 | 4.85 | -0.07 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 6.398
Bio-Rad D-10 | 153 | 6.13 | 0.09 | 3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 5.97
Bio-Rad D-10 | 153 | 5.39 | 0.11 | 2.9 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 5.794
Bio-Rad D-10 | 153 | 8.28 | 0.2 | 2.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 5.296
Bio-Rad D-10 | 155 | 5.41 | -0.05 | 3.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 6.81
Bio-Rad D-10 | 155 | 5.51 | -0.15 | 4.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 8.474
Bio-Rad D-10 | 155 | 9.49 | 0.18 | 2.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 4.688
Bio-Rad D-10 | 155 | 5.23 | -0.05 | 3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 5.83
Bio-Rad D-10 | 155 | 7.54 | 0.13 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 4.442
Bio-Rad D-10 | 123 | 6.55 | 0.14 | 2.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 4.844
Bio-Rad D-10 | 123 | 8.8 | 0.2 | 2.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 4.708
Bio-Rad D-10 | 123 | 5.45 | 0.03 | 2.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 5.322
Bio-Rad D-10 | 123 | 7.47 | 0.09 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 4.794
Bio-Rad D-10 | 123 | 9.89 | 0.14 | 2.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 4.844
Bio-Rad D-10 | 135 | 5.17 | -0.08 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 4.624
Bio-Rad D-10 | 135 | 6.51 | 0.1 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 4.216
Bio-Rad D-10 | 135 | 8.16 | -0.05 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 4.066
Bio-Rad D-10 | 135 | 5.44 | -0.03 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 4.282
Bio-Rad D-10 | 135 | 9.58 | 0.1 | 1.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 3.824
Bio-Rad D-10 | 117 | 7.89 | -0.08 | 2.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 5.408
Bio-Rad D-10 | 117 | 5.9 | 0.01 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 4.518
Bio-Rad D-10 | 117 | 5.07 | -0.06 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 4.448
Bio-Rad D-10 | 117 | 7.33 | -0.07 | 2.8 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 5.418
Bio-Rad D-10 | 117 | 9.23 | 0.06 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 4.568
Bio-Rad D-10 | 103 | 7.8 | -0.01 | 1.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 3.518
Bio-Rad D-10 | 103 | 5.92 | 0.01 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 3.538
Bio-Rad D-10 | 103 | 8.78 | 0.09 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 3.422
Bio-Rad D-10 | 103 | 6.03 | 0.17 | 1.9 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 3.894
Bio-Rad D-10 | 103 | 9.59 | 0.03 | 1.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 3.754
Bio-Rad D-10 | 124 | 5.58 | 0.1 | 2.6 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 5.196
Bio-Rad D-10 | 124 | 8.29 | 0.03 | 2.5 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 4.93
Bio-Rad D-10 | 124 | 5.09 | -0.06 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 5.036
Bio-Rad D-10 | 124 | 10.14 | 0.09 | 2.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 4.99
Bio-Rad D-10 | 124 | 6.01 | 0.02 | 2.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 4.332
Bio-Rad D-10 | 116 | 9.12 | 0.26 | 2.3 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 4.768
Bio-Rad D-10 | 116 | 6.98 | 0.14 | 2.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 5.04
Bio-Rad D-10 | 116 | 9.53 | 0.14 | 1.9 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 3.864
Bio-Rad D-10 | 116 | 6.19 | 0.06 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 4.372
Bio-Rad D-10 | 116 | 7.55 | 0.03 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 3.558
Bio-Rad D-10 | 100 | 5.99 | 0.19 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 4.502
Bio-Rad D-10 | 100 | 9.25 | 0.13 | 2.1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 4.246
Bio-Rad D-10 | 100 | 5.58 | 0.11 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 4.226
Bio-Rad D-10 | 100 | 7.24 | -0.01 | 2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 3.91
Bio-Rad D-10 | 100 | 9.49 | 0.23 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 4.15
Bio-Rad D-10 | 112 | 5.3 | 0.09 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 4.402
Bio-Rad D-10 | 112 | 6.82 | 0.14 | 1.9 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 3.864
Bio-Rad D-10 | 112 | 8.49 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad D-10 | 112 | 6.13 | 0.15 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 3.874
Bio-Rad D-10 | 112 | 10.25 | 0.13 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 3.854
Bio-Rad D-10 | 37 | 5.9381 | -0.9 | 5.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Bio-Rad | 9.88
Bio-Rad D-10 | 36 | 7.3106 | -1.5 | 4.9 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Bio-Rad | 8.104
Bio-Rad D-10 | 53 | 6.0845 | 0.7 | 4.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Bio-Rad | 9.716
Bio-Rad D-10 | 53 | 7.5302 | 0.9 | 5.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Bio-Rad | 10.896
Bio-Rad D-10 | 39 | 6.1394 | -0.6 | 6.5 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Bio-Rad | 12.14
Bio-Rad D-10 | 39 | 7.2557 | -2.2 | 4.5 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Bio-Rad | 6.62
Bio-Rad D-10 | 59 | 6.3224 | 1.4 | 4.7 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Bio-Rad | 10.612
Bio-Rad D-10 | 60 | 7.5942 | 1.5 | 4.8 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Bio-Rad | 10.908
Bio-Rad D-10 | 36 | 7.3655 | -1.1 | 4.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Bio-Rad | 8.504
Bio-Rad D-10 | 37 | 5.993 | -0.2 | 5.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Bio-Rad | 11.168
Bio-Rad D-10 | 72 | 7.6857 | 2.4 | 5.3 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Bio-Rad | 12.788
Bio-Rad D-10 | 72 | 6.1669 | 1.7 | 5.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Bio-Rad | 12.284
Bio-Rad D-10 | 106 | 8.8661 | 3.3 | 5.2 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Bio-Rad | 13.492
Bio-Rad D-10 | 107 | 5.6636 | 0.2 | 6.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Bio-Rad | 13.136
Bio-Rad D-10 | 91 | 8.8935 | 3.6 | 3.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Bio-Rad | 10.264
Bio-Rad D-10 | 92 | 5.7368 | 1 | 4.3 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Bio-Rad | 9.428
Bio-Rad D-10 | 98 | 5.6087 | 1.3 | 7.2 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 15.412
Bio-Rad D-10 | 94 | 5.5904 | 1.1 | 6.4 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 13.644
Bio-Rad D-10 | 72 | 5.508 | 0.2 | 4.7 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 9.412
Bio-Rad D-10 | 72 | 5.508 | 0.2 | 4.4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 8.824
Bio-Rad D-100 | 17-24 | 6.38 | -0.03 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 3.106
Bio-Rad D-100 | 17-24 | 9.41 | -0.12 | 2.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 3.996
Bio-Rad D-100 | 17-24 | 5.25 | -0.09 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 4.026
Bio-Rad D-100 | 17-24 | 8.47 | -0.04 | 2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 3.88
Bio-Rad D-100 | 17-24 | 7.18 | -0.07 | 1.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 3.262
Bio-Rad D-100 | 23 | 7.33 | -0.09 | 2.4 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 4.614
Bio-Rad D-100 | 23 | 5.2 | 0 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 3.92
Bio-Rad D-100 | 23 | 8.13 | -0.18 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 3.152
Bio-Rad D-100 | 23 | 10.06 | -0.27 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 3.454
Bio-Rad D-100 | 23 | 5.83 | -0.04 | 1.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 3.292
Bio-Rad D-100 | 41 | 7.19 | -0.06 | 1.7 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 3.272
Bio-Rad D-100 | 41 | 6.76 | -0.05 | 1.7 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 3.282
Bio-Rad D-100 | 41 | 8.49 | -0.16 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 2.388
Bio-Rad D-100 | 41 | 9.32 | -0.18 | 2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 3.74
Bio-Rad D-100 | 41 | 5.42 | -0.03 | 1.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 3.302
Bio-Rad D-100 | 60 | 7.02 | -0.13 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 2.614
Bio-Rad D-100 | 60 | 5.15 | -0.04 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 3.88
Bio-Rad D-100 | 60 | 8.18 | -0.24 | 1.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 3.288
Bio-Rad D-100 | 60 | 9.52 | -0.27 | 1.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 2.866
Bio-Rad D-100 | 60 | 6.1 | -0.02 | 1.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 2.92
Bio-Rad D-100 | 59 | 6.23 | -0.08 | 1.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 3.644
Bio-Rad D-100 | 59 | 8.85 | -0.26 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 2.68
Bio-Rad D-100 | 59 | 5.18 | -0.13 | 2.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 3.986
Bio-Rad D-100 | 59 | 8.13 | -0.22 | 1.6 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 2.916
Bio-Rad D-100 | 59 | 7.48 | -0.14 | 1.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 3.192
Bio-Rad D-100 | 85 | 9.14 | -0.23 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 3.298
Bio-Rad D-100 | 85 | 4.86 | -0.06 | 2.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 5.036
Bio-Rad D-100 | 85 | 6 | -0.04 | 2.2 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 4.272
Bio-Rad D-100 | 85 | 5.28 | 0 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad D-100 | 85 | 8 | -0.08 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 3.84
Bio-Rad D-100 | 102 | 5.33 | -0.13 | 2.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 4.182
Bio-Rad D-100 | 102 | 5.58 | -0.08 | 2.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 4.624
Bio-Rad D-100 | 102 | 9.11 | -0.2 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 3.132
Bio-Rad D-100 | 102 | 5.11 | -0.17 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 4.534
Bio-Rad D-100 | 102 | 7.23 | -0.18 | 2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 3.74
Bio-Rad D-100 | 94 | 6.3 | -0.11 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 3.614
Bio-Rad D-100 | 94 | 8.45 | -0.15 | 1.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 2.79
Bio-Rad D-100 | 94 | 5.25 | -0.17 | 2.2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 4.142
Bio-Rad D-100 | 94 | 7.21 | -0.17 | 1.6 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 2.966
Bio-Rad D-100 | 94 | 9.55 | -0.2 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 2.152
Bio-Rad D-100 | 119 | 4.99 | -0.26 | 2.3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 4.248
Bio-Rad D-100 | 119 | 6.27 | -0.14 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 3.976
Bio-Rad D-100 | 119 | 7.95 | -0.26 | 2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 3.66
Bio-Rad D-100 | 119 | 5.26 | -0.21 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 4.102
Bio-Rad D-100 | 119 | 9.27 | -0.21 | 2.1 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 3.906
Bio-Rad D-100 | 116 | 7.86 | -0.11 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 4.202
Bio-Rad D-100 | 116 | 5.75 | -0.14 | 2.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 4.172
Bio-Rad D-100 | 116 | 4.92 | -0.21 | 2.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 4.69
Bio-Rad D-100 | 116 | 7.22 | -0.18 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 4.132
Bio-Rad D-100 | 116 | 9.02 | -0.15 | 2.1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 3.966
Bio-Rad D-100 | 109 | 7.72 | -0.09 | 1.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 3.438
Bio-Rad D-100 | 109 | 5.83 | -0.08 | 1.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 3.644
Bio-Rad D-100 | 109 | 8.64 | -0.05 | 1.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 3.674
Bio-Rad D-100 | 109 | 5.91 | 0.05 | 2.1 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 4.166
Bio-Rad D-100 | 109 | 9.45 | -0.11 | 1.7 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 3.222
Bio-Rad D-100 | 131 | 5.46 | -0.02 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 3.508
Bio-Rad D-100 | 131 | 8.22 | -0.04 | 1.7 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 3.292
Bio-Rad D-100 | 131 | 5.03 | -0.12 | 2.1 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 3.996
Bio-Rad D-100 | 131 | 10.02 | -0.03 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 3.302
Bio-Rad D-100 | 131 | 5.92 | -0.07 | 1.9 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 3.654
Bio-Rad D-100 | 131 | 8.93 | 0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 3.402
Bio-Rad D-100 | 131 | 6.86 | 0.02 | 1.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 3.156
Bio-Rad D-100 | 131 | 9.35 | -0.04 | 1.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 2.704
Bio-Rad D-100 | 131 | 6.02 | -0.11 | 1.6 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 3.026
Bio-Rad D-100 | 131 | 7.42 | -0.1 | 1.7 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 3.232
Bio-Rad D-100 | 112 | 5.82 | 0.02 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 3.352
Bio-Rad D-100 | 112 | 9.06 | -0.06 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 2.88
Bio-Rad D-100 | 112 | 5.42 | -0.05 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 4.066
Bio-Rad D-100 | 112 | 7.13 | -0.12 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 2.82
Bio-Rad D-100 | 112 | 9.27 | 0.01 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 2.754
Bio-Rad D-100 | 133 | 5.12 | -0.09 | 2.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 4.026
Bio-Rad D-100 | 133 | 6.66 | -0.02 | 1.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 3.312
Bio-Rad D-100 | 133 | 8.33 |  | 2.1 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad D-100 | 133 | 5.93 | -0.05 | 1.5 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 2.89
Bio-Rad D-100 | 133 | 9.99 | -0.13 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 3.202
Bio-Rad D-100 | 16 | 6.0296 | 0.1 | 2.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Bio-Rad | 4.216
Bio-Rad D-100 | 19 | 7.3746 | -0.8 | 1.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Bio-Rad | 2.532
Bio-Rad D-100 | 11 | 5.9381 | -0.9 | 1.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Bio-Rad | 2.236
Bio-Rad D-100 | 11 | 7.3746 | -0.8 | 2.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Bio-Rad | 3.316
Bio-Rad D-100 | 18 | 6.2309 | 0.4 | 2.6 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Bio-Rad | 5.496
Bio-Rad D-100 | 18 | 7.4021 | -0.6 | 1.5 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Bio-Rad | 2.34
Bio-Rad D-100 | 19 | 6.1852 | -0.1 | 3.2 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Bio-Rad | 6.172
Bio-Rad D-100 | 19 | 7.3838 | -0.8 | 2.5 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Bio-Rad | 4.1
Bio-Rad D-100 | 28 | 7.3655 | -1.1 | 2.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Bio-Rad | 3.016
Bio-Rad D-100 | 28 | 6.0296 | 0.2 | 2.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Bio-Rad | 5.688
Bio-Rad D-100 | 59 | 7.4021 | -0.7 | 3.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Bio-Rad | 6.16
Bio-Rad D-100 | 59 | 6.0022 | -0.1 | 3.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Bio-Rad | 7.544
Bio-Rad D-100 | 49 | 8.6831 | 1.3 | 2.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Bio-Rad | 5.808
Bio-Rad D-100 | 49 | 5.6087 | -0.4 | 3.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Bio-Rad | 6.656
Bio-Rad D-100 | 124 | 8.4635 | -1.1 | 2.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Bio-Rad | 4.388
Bio-Rad D-100 | 123 | 5.4531 | -2.1 | 3.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Bio-Rad | 4.956
Bio-Rad D-100 | 78 | 5.6087 | 1.3 | 2.9 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 6.984
Bio-Rad D-100 | 74 | 5.5904 | 1.1 | 3.1 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 7.176
Bio-Rad D-100 | 78 | 5.4897 | 0 | 3.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 6.076
Bio-Rad D-100 | 77 | 5.4897 | 0 | 2.7 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 5.292
Bio-Rad Diastat | 15 | 10 |  | 6.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 11.956
Bio-Rad Diastat | 15 | 5 |  | 5.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 11.564
Bio-Rad Diastat | 15 | 8 |  | 5.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 10.78
Bio-Rad Diastat | 13 | 6.9 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 9.996
Bio-Rad Diastat | 13 | 6.8 |  | 5.5 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 10.78
Bio-Rad Diastat | 13 | 11.8 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 7.056
Bio-Rad Diastat | 13 | 6.9 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 9.996
Bio-Rad Diastat | 13 | 6.8 |  | 5.5 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 10.78
Bio-Rad Diastat | 13 | 11.8 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 7.056
Bio-Rad Diastat | 11 | 5.2 |  | 3.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 6.86
Bio-Rad Diastat | 11 | 11.8 |  | 4.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 8.624
Bio-Rad Diastat | 11 | 7.3 |  | 3.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 7.252
Bio-Rad Diastat | 7 | 8.9 |  |  | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 0
Bio-Rad Diastat | 7 | 10.9 |  |  | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 0
Bio-Rad Diastat | 7 | 5.7 |  |  | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 0
Bio-Rad in2it | 10 | 6.1 | 0.2 | 3.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 7.256
Bio-Rad in2it | 10 | 10.28 | 0.48 | 5.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 11.456
Bio-Rad in2it | 10 | 7.41 | 0.01 | 5.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 11.182
Bio-Rad in2it | 10 | 5.41 | -0.19 | 7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 13.53
Bio-Rad in2it | 10 | 9.01 | -0.39 | 3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 5.49
Bio-Rad in2it | 10 | 7.24 | 0.04 | 5.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 10.624
Bio-Rad in2it | 11 | 5.3 | -0.1 | 5.2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 10.092
Bio-Rad in2it | 11 | 8.13 | -0.17 | 3.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 7.474
Bio-Rad in2it | 11 | 5.55 | -0.1 | 7.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 13.816
Bio-Rad not specified/other | 115 | 5.4714 | -0.2 | 4.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 8.62
Bio-Rad not specified/other | 114 | 5.4714 | -0.2 | 4.4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 8.424
Bio-Rad not specified/other | 53 | 6.2309 | 0.4 | 4.4 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Bio-Rad | 9.024
Bio-Rad not specified/other | 53 | 7.4204 | -0.4 | 4.1 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Bio-Rad | 7.636
Bio-Rad not specified/other | 47 | 7.4478 | -0.2 | 5.3 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Bio-Rad | 10.188
Bio-Rad not specified/other | 47 | 6.0845 | 0.8 | 5.5 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Bio-Rad | 11.58
Bio-Rad not specified/other | 10 | 7.7223 | 2.8 | 2.3 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Bio-Rad | 7.308
Bio-Rad not specified/other | 10 | 6.2126 | 2.2 | 4.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Bio-Rad | 11.804
Bio-Rad Variant | 10 | 13.6 |  | 2.7 | GH2-04 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad Variant | 10 | 5.6 |  | 8 | GH2-05 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 15.68
Bio-Rad Variant | 10 | 10.2 |  | 5.9 | GH2-06 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 11.564
Bio-Rad Variant | 9 | 7.6 |  |  | GH2-04 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 0
Bio-Rad Variant | 9 | 7.6 |  |  | GH2-05 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 0
Bio-Rad Variant | 9 | 14.8 |  |  | GH2-06 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 0
Bio-Rad Variant | 10 | 5.4 |  | 6.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 12.74
Bio-Rad Variant | 10 | 14.1 |  | 4.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 8.232
Bio-Rad Variant | 10 | 8.8 |  | 3.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 7.644
Bio-Rad Variant A1c | 20 | 10.5 |  | 1.7 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 3.332
Bio-Rad Variant A1c | 20 | 5.2 |  | 1.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 2.548
Bio-Rad Variant A1c | 20 | 8.4 |  | 1.9 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 3.724
Bio-Rad Variant A1c | 23 | 6.9 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad Variant A1c | 23 | 6.9 |  | 3.1 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 6.076
Bio-Rad Variant A1c | 23 | 11.6 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 7.056
Bio-Rad Variant A1c | 23 | 6.9 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad Variant A1c | 23 | 6.9 |  | 3.1 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 6.076
Bio-Rad Variant A1c | 23 | 11.6 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 7.056
Bio-Rad Variant A1c | 17 | 5.2 |  | 2.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant A1c | 17 | 11.5 |  | 2.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant A1c | 17 | 7.5 |  | 3.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 7.252
Bio-Rad Variant A1c | 15 | 9.1 |  | 3.2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant A1c | 15 | 11 |  | 3 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant A1c | 15 | 6.2 |  | 2.3 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant A1c | 10 | 12.7 |  | 2.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant A1c | 10 | 5.1 |  | 4.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 8.036
Bio-Rad Variant A1c | 10 | 8.7 |  | 1.3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 2.548
Bio-Rad Variant A1c | 13 | 12 |  | 3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant A1c | 13 | 6.8 |  | 1.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 3.136
Bio-Rad Variant A1c | 13 | 4.9 |  | 2.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad Variant II | 139 | 6.81 | 0.21 | 2.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 5.698
Bio-Rad Variant II | 139 | 9.75 | 0.25 | 2.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 5.542
Bio-Rad Variant II | 139 | 7.67 | 0.27 | 3.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 6.934
Bio-Rad Variant II | 118 | 6.14 | 0.24 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 6.316
Bio-Rad Variant II | 118 | 10.07 | 0.27 | 2.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 4.778
Bio-Rad Variant II | 118 | 7.64 | 0.24 | 2.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 5.532
Bio-Rad Variant II | 110 | 5.32 | 0.12 | 2.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 5.02
Bio-Rad Variant II | 110 | 8.87 | 0.17 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 5.07
Bio-Rad Variant II | 110 | 6.4 | 0.1 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 4.412
Bio-Rad Variant II | 107 | 8.65 | 0.15 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 4.462
Bio-Rad Variant II | 107 | 5.47 | 0.07 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 5.558
Bio-Rad Variant II | 107 | 6.48 | 0.08 | 2.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 4.392
Bio-Rad Variant II | 11 | 6.18 | -0.12 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 7.328
Bio-Rad Variant II | 11 | 7.55 | -0.05 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 5.046
Bio-Rad Variant II | 11 | 9.11 | -0.09 | 4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 7.75
Bio-Rad Variant II | 111 | 5.6 | 0 | 2.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II | 111 | 9.61 | 0.21 | 2.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 5.11
Bio-Rad Variant II | 111 | 7.46 | 0.26 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 5.16
Bio-Rad Variant II | 103 | 5.32 | -0.08 | 2.7 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 5.212
Bio-Rad Variant II | 103 | 8.36 | 0.06 | 2.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 4.568
Bio-Rad Variant II | 103 | 5.68 | 0.03 | 3.7 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 7.282
Bio-Rad Variant II | 97 | 7.04 | -0.07 | 2.1 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 4.046
Bio-Rad Variant II | 97 | 9.37 | 0.05 | 2.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 4.362
Bio-Rad Variant II | 97 | 5.97 | -0.1 | 2.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 4.212
Bio-Rad Variant II | 88 | 8.06 | 0.01 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 4.714
Bio-Rad Variant II | 88 | 5.24 | -0.06 | 2.4 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 4.644
Bio-Rad Variant II | 88 | 6.06 | -0.08 | 2.4 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 4.624
Bio-Rad Variant II | 27 | 9.41 | 0.04 | 2.3 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 4.548
Bio-Rad Variant II | 27 | 4.83 | -0.09 | 3.2 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 6.182
Bio-Rad Variant II | 27 | 6.1 | 0.06 | 2.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 4.96
Bio-Rad Variant II | 27 | 5.33 | 0.05 | 2.5 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 4.95
Bio-Rad Variant II | 27 | 8.18 | 0.1 | 2.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 4.216
Bio-Rad Variant II | 22 | 5.35 | -0.11 | 2.6 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 4.986
Bio-Rad Variant II | 22 | 5.69 | 0.03 | 3.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 6.694
Bio-Rad Variant II | 22 | 9.47 | 0.16 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 4.08
Bio-Rad Variant II | 22 | 5.15 | -0.13 | 2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 3.79
Bio-Rad Variant II | 22 | 7.45 | 0.04 | 1.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 3.764
Bio-Rad Variant II | 18 | 6.47 | 0.06 | 1.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.412
Bio-Rad Variant II | 18 | 8.79 | 0.19 | 1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 2.15
Bio-Rad Variant II | 18 | 5.31 | -0.11 | 2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 3.81
Bio-Rad Variant II | 18 | 7.44 | 0.06 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 2.412
Bio-Rad Variant II | 18 | 9.91 | 0.16 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 2.512
Bio-Rad Variant II | 21 | 5.16 | -0.09 | 3.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 6.574
Bio-Rad Variant II | 21 | 6.49 | 0.08 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 5.372
Bio-Rad Variant II | 21 | 8.34 | 0.13 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 3.854
Bio-Rad Variant II | 21 | 5.4 | -0.07 | 2.9 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 5.614
Bio-Rad Variant II | 21 | 9.64 | 0.16 | 2.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 5.452
Bio-Rad Variant II | 16 | 7.93 | -0.04 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 5.252
Bio-Rad Variant II | 16 | 5.84 | -0.05 | 2.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 5.634
Bio-Rad Variant II | 16 | 4.98 | -0.15 | 2.6 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 4.946
Bio-Rad Variant II | 16 | 7.35 | -0.05 | 4.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 8.182
Bio-Rad Variant II | 16 | 9.29 | 0.12 | 2.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 4.824
Bio-Rad Variant II | 17 | 7.91 | 0.1 | 2.9 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 5.784
Bio-Rad Variant II | 17 | 5.93 | 0.02 | 2.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 5.116
Bio-Rad Variant II | 17 | 8.89 | 0.2 | 2.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 5.688
Bio-Rad Variant II | 17 | 6.1 | 0.24 | 3.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 6.904
Bio-Rad Variant II | 17 | 9.67 | 0.11 | 2.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 5.206
Bio-Rad Variant II | 18 | 5.57 | 0.09 | 1.9 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 3.814
Bio-Rad Variant II | 18 | 8.37 | 0.11 | 2.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 4.618
Bio-Rad Variant II | 18 | 5.15 | 0 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 6.468
Bio-Rad Variant II | 18 | 10.31 | 0.26 | 2.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 4.572
Bio-Rad Variant II | 18 | 6.13 | 0.14 | 2.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 4.256
Bio-Rad Variant II | 17 | 9.11 | 0.25 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 3.386
Bio-Rad Variant II | 17 | 6.96 | 0.12 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 4.628
Bio-Rad Variant II | 17 | 9.55 | 0.16 | 2.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 4.276
Bio-Rad Variant II | 17 | 6.18 | 0.05 | 1.3 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 2.598
Bio-Rad Variant II | 17 | 7.59 | 0.07 | 2.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 4.186
Bio-Rad Variant II | 18 | 5.91 | 0.11 | 2.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 5.206
Bio-Rad Variant II | 18 | 9.28 | 0.16 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad Variant II | 18 | 5.48 | 0.01 | 3.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 6.086
Bio-Rad Variant II | 18 | 7.23 | -0.02 | 2.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 5.076
Bio-Rad Variant II | 18 | 9.46 | 0.2 | 1.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 3.924
Bio-Rad Variant II | 15 | 5.31 | 0.1 | 5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 9.9
Bio-Rad Variant II | 15 | 6.81 | 0.13 | 2.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 5.422
Bio-Rad Variant II | 15 | 8.51 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant II | 15 | 6.11 | 0.13 | 3.5 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 6.99
Bio-Rad Variant II | 15 | 10.24 | 0.12 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 4.432
Bio-Rad Variant II | 254 | 11.2 |  | 2.8 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II | 254 | 5.3 |  | 2.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II | 254 | 8.7 |  | 2.9 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad Variant II | 239 | 7.1 |  | 3.2 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant II | 239 | 7.1 |  | 3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 239 | 12.2 |  | 3 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 239 | 7.1 |  | 3.2 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant II | 239 | 7.1 |  | 3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 239 | 12.2 |  | 3 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 253 | 5.4 |  | 3.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 6.664
Bio-Rad Variant II | 253 | 12.3 |  | 2.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant II | 253 | 7.9 |  | 2.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II | 200 | 9.6 |  | 2.3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II | 200 | 11.5 |  | 2.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad Variant II | 200 | 6.3 |  | 2.8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II | 175 | 13.4 |  | 2.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant II | 175 | 5.3 |  | 3.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 6.86
Bio-Rad Variant II | 175 | 9.2 |  | 2.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad Variant II | 167 | 12.4 |  | 2.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II | 167 | 7 |  | 3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 167 | 5 |  | 4 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 7.84
Bio-Rad Variant II | 146 | 5.1 |  | 3.6 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 7.056
Bio-Rad Variant II | 146 | 8.7 |  | 2.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II | 146 | 6.2 |  | 2.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II Turbo | 147 | 6.76 | 0.16 | 2.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 5.844
Bio-Rad Variant II Turbo | 147 | 9.45 | -0.05 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 4.654
Bio-Rad Variant II Turbo | 147 | 7.58 | 0.18 | 2.6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 5.276
Bio-Rad Variant II Turbo | 143 | 6.04 | 0.14 | 3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 6.02
Bio-Rad Variant II Turbo | 143 | 9.66 | -0.14 | 2.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 4.368
Bio-Rad Variant II Turbo | 143 | 7.5 | 0.1 | 2.3 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 4.608
Bio-Rad Variant II Turbo | 146 | 5.28 | 0.08 | 2.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 5.372
Bio-Rad Variant II Turbo | 146 | 8.74 | 0.04 | 2.2 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 4.352
Bio-Rad Variant II Turbo | 146 | 6.38 | 0.08 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 4.392
Bio-Rad Variant II Turbo | 148 | 8.59 | 0.09 | 2.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 4.794
Bio-Rad Variant II Turbo | 148 | 5.42 | 0.02 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 5.312
Bio-Rad Variant II Turbo | 148 | 6.45 | 0.05 | 2.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 4.754
Bio-Rad Variant II Turbo | 96 | 6.45 | 0.15 | 3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 6.03
Bio-Rad Variant II Turbo | 96 | 7.82 | 0.22 | 2.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 5.512
Bio-Rad Variant II Turbo | 96 | 9.45 | 0.25 | 2.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 4.758
Bio-Rad Variant II Turbo | 183 | 5.53 | -0.07 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 5.026
Bio-Rad Variant II Turbo | 183 | 9.52 | 0.12 | 2.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 4.432
Bio-Rad Variant II Turbo | 183 | 7.35 | 0.15 | 2.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 4.266
Bio-Rad Variant II Turbo | 174 | 5.4 | 0 | 2.9 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 5.684
Bio-Rad Variant II Turbo | 174 | 8.42 | 0.12 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 5.02
Bio-Rad Variant II Turbo | 174 | 5.66 | 0.01 | 4.2 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 8.242
Bio-Rad Variant II Turbo | 152 | 7.16 | 0.05 | 2.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 5.146
Bio-Rad Variant II Turbo | 152 | 9.43 | 0.11 | 2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 4.03
Bio-Rad Variant II Turbo | 152 | 6.05 | -0.02 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 5.076
Bio-Rad Variant II Turbo | 133 | 8.15 | 0.1 | 2.3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 4.608
Bio-Rad Variant II Turbo | 133 | 5.31 | 0.01 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 5.498
Bio-Rad Variant II Turbo | 133 | 6.2 | 0.06 | 2.3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 4.568
Bio-Rad Variant II Turbo | 55 | 9.44 | 0.07 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 4.186
Bio-Rad Variant II Turbo | 55 | 4.89 | -0.03 | 2.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 4.478
Bio-Rad Variant II Turbo | 55 | 6.09 | 0.05 | 1.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 3.578
Bio-Rad Variant II Turbo | 55 | 5.32 | 0.04 | 2.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 4.156
Bio-Rad Variant II Turbo | 55 | 8.24 | 0.16 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 4.08
Bio-Rad Variant II Turbo | 42 | 5.4 | -0.06 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 4.84
Bio-Rad Variant II Turbo | 42 | 5.84 | 0.18 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 5.472
Bio-Rad Variant II Turbo | 42 | 9.45 | 0.14 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 3.472
Bio-Rad Variant II Turbo | 42 | 5.2 | -0.08 | 2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 3.84
Bio-Rad Variant II Turbo | 42 | 7.5 | 0.09 | 1.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 3.618
Bio-Rad Variant II Turbo | 39 | 6.46 | 0.05 | 2.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 5.734
Bio-Rad Variant II Turbo | 39 | 8.85 | 0.25 | 2.4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 4.954
Bio-Rad Variant II Turbo | 39 | 5.37 | -0.05 | 3.3 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 6.418
Bio-Rad Variant II Turbo | 39 | 7.44 | 0.06 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 4.764
Bio-Rad Variant II Turbo | 39 | 9.92 | 0.17 | 2.5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 5.07
Bio-Rad Variant II Turbo | 29 | 5.06 | -0.19 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 5.102
Bio-Rad Variant II Turbo | 29 | 6.42 | 0.01 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 3.734
Bio-Rad Variant II Turbo | 29 | 8.19 | -0.02 | 1.7 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 3.312
Bio-Rad Variant II Turbo | 29 | 5.31 | -0.16 | 2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 3.76
Bio-Rad Variant II Turbo | 29 | 9.53 | 0.05 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 3.382
Bio-Rad Variant II Turbo 2.0 | 19 | 6.27 | 0.37 | 2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 4.29
Bio-Rad Variant II Turbo 2.0 | 19 | 9.93 | 0.13 | 2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 4.05
Bio-Rad Variant II Turbo 2.0 | 19 | 7.67 | 0.27 | 1.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 3.602
Bio-Rad Variant II Turbo 2.0 | 22 | 5.39 | 0.19 | 1.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 3.13
Bio-Rad Variant II Turbo 2.0 | 22 | 8.91 | 0.21 | 1.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 2.954
Bio-Rad Variant II Turbo 2.0 | 22 | 6.51 | 0.21 | 1.5 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 3.15
Bio-Rad Variant II Turbo 2.0 | 30 | 8.7 | 0.2 | 2.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 4.316
Bio-Rad Variant II Turbo 2.0 | 30 | 5.57 | 0.17 | 2.4 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 4.874
Bio-Rad Variant II Turbo 2.0 | 30 | 6.59 | 0.19 | 2.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 4.502
Bio-Rad Variant II Turbo 2.0 | 166 | 6.43 | 0.13 | 2.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 4.834
Bio-Rad Variant II Turbo 2.0 | 166 | 7.82 | 0.22 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 5.12
Bio-Rad Variant II Turbo 2.0 | 166 | 9.36 | 0.16 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 5.648
Bio-Rad Variant II Turbo 2.0 | 47 | 5.69 | 0.09 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 5.97
Bio-Rad Variant II Turbo 2.0 | 47 | 9.64 | 0.24 | 1.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 3.572
Bio-Rad Variant II Turbo 2.0 | 47 | 7.54 | 0.34 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 5.24
Bio-Rad Variant II Turbo 2.0 | 48 | 5.5 | 0.1 | 2.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 4.216
Bio-Rad Variant II Turbo 2.0 | 48 | 8.52 | 0.22 | 1.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 3.944
Bio-Rad Variant II Turbo 2.0 | 48 | 5.96 | 0.31 | 3.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 7.758
Bio-Rad Variant II Turbo 2.0 | 51 | 7.12 | 0.01 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 4.322
Bio-Rad Variant II Turbo 2.0 | 51 | 9.35 | 0.03 | 2.3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 4.538
Bio-Rad Variant II Turbo 2.0 | 51 | 6.11 | 0.04 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 5.136
Bio-Rad Variant II Turbo 2.0 | 60 | 8.13 | 0.08 | 1.9 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 3.804
Bio-Rad Variant II Turbo 2.0 | 60 | 5.27 | -0.03 | 2.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 4.87
Bio-Rad Variant II Turbo 2.0 | 60 | 6.15 | 0.01 | 1.9 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 3.734
Bio-Rad Variant II Turbo 2.0 | 125 | 9.5 | 0.13 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 4.246
Bio-Rad Variant II Turbo 2.0 | 125 | 4.89 | -0.03 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 5.262
Bio-Rad Variant II Turbo 2.0 | 125 | 6.12 | 0.08 | 1.9 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 3.804
Bio-Rad Variant II Turbo 2.0 | 125 | 5.38 | 0.1 | 2.3 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 4.608
Bio-Rad Variant II Turbo 2.0 | 125 | 8.24 | 0.16 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 4.08
Bio-Rad Variant II Turbo 2.0 | 120 | 5.47 | 0.01 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 4.91
Bio-Rad Variant II Turbo 2.0 | 120 | 5.93 | 0.27 | 2.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 4.778
Bio-Rad Variant II Turbo 2.0 | 120 | 9.48 | 0.17 | 1.6 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 3.306
Bio-Rad Variant II Turbo 2.0 | 120 | 5.23 | -0.05 | 2.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 5.046
Bio-Rad Variant II Turbo 2.0 | 120 | 7.51 | 0.1 | 1.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 3.824
Bio-Rad Variant II Turbo 2.0 | 106 | 6.48 | 0.07 | 2.1 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 4.186
Bio-Rad Variant II Turbo 2.0 | 106 | 8.85 | 0.25 | 2 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 4.17
Bio-Rad Variant II Turbo 2.0 | 106 | 5.39 | -0.03 | 2.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 5.262
Bio-Rad Variant II Turbo 2.0 | 106 | 7.41 | 0.03 | 1.9 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 3.754
Bio-Rad Variant II Turbo 2.0 | 106 | 9.88 | 0.13 | 1.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 3.854
Bio-Rad Variant II Turbo 2.0 | 131 | 5.05 | -0.2 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 5.092
Bio-Rad Variant II Turbo 2.0 | 131 | 6.41 | 0 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 3.724
Bio-Rad Variant II Turbo 2.0 | 131 | 8.19 | -0.02 | 1.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 3.508
Bio-Rad Variant II Turbo 2.0 | 131 | 5.33 | -0.14 | 2.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 4.564
Bio-Rad Variant II Turbo 2.0 | 131 | 9.52 | 0.04 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 3.372
Bio-Rad Variant II Turbo 2.0 | 138 | 7.96 | -0.01 | 2.6 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 5.086
Bio-Rad Variant II Turbo 2.0 | 138 | 5.79 | -0.1 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 3.82
Bio-Rad Variant II Turbo 2.0 | 138 | 5 | -0.13 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 4.378
Bio-Rad Variant II Turbo 2.0 | 138 | 7.28 | -0.12 | 2.4 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 4.584
Bio-Rad Variant II Turbo 2.0 | 138 | 9.18 | 0.01 | 2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 3.93
Bio-Rad Variant II Turbo 2.0 | 131 | 7.8 | -0.01 | 2.2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 4.302
Bio-Rad Variant II Turbo 2.0 | 131 | 5.85 | -0.06 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 4.252
Bio-Rad Variant II Turbo 2.0 | 131 | 8.84 | 0.15 | 2.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 4.266
Bio-Rad Variant II Turbo 2.0 | 131 | 5.97 | 0.11 | 2.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 4.618
Bio-Rad Variant II Turbo 2.0 | 131 | 9.59 | 0.03 | 2.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 4.734
Bio-Rad Variant II Turbo 2.0 | 139 | 5.5 | 0.02 | 2.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 4.528
Bio-Rad Variant II Turbo 2.0 | 139 | 8.32 | 0.06 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 3.784
Bio-Rad Variant II Turbo 2.0 | 139 | 5.04 | -0.11 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 4.986
Bio-Rad Variant II Turbo 2.0 | 139 | 10.22 | 0.17 | 1.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 3.894
Bio-Rad Variant II Turbo 2.0 | 139 | 5.94 | -0.05 | 2.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 4.458
Bio-Rad Variant II Turbo 2.0 | 134 | 9.03 | 0.17 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 4.482
Bio-Rad Variant II Turbo 2.0 | 134 | 6.9 | 0.06 | 2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 3.98
Bio-Rad Variant II Turbo 2.0 | 134 | 9.49 | 0.1 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 4.02
Bio-Rad Variant II Turbo 2.0 | 134 | 6.05 | -0.08 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 4.232
Bio-Rad Variant II Turbo 2.0 | 134 | 7.52 | 0 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 3.528
Bio-Rad Variant II Turbo 2.0 | 113 | 5.85 | 0.05 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 4.362
Bio-Rad Variant II Turbo 2.0 | 113 | 9.21 | 0.09 | 1.7 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 3.422
Bio-Rad Variant II Turbo 2.0 | 113 | 5.43 | -0.04 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 4.076
Bio-Rad Variant II Turbo 2.0 | 113 | 7.16 | -0.09 | 1.8 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 3.438
Bio-Rad Variant II Turbo 2.0 | 113 | 9.37 | 0.11 | 1.7 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 3.442
Bio-Rad Variant II Turbo 2.0 | 121 | 5.16 | -0.05 | 2.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 4.654
Bio-Rad Variant II Turbo 2.0 | 121 | 6.74 | 0.06 | 1.8 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 3.588
Bio-Rad Variant II Turbo 2.0 | 121 | 8.43 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant II Turbo 2.0 | 121 | 5.97 | -0.01 | 2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 3.91
Bio-Rad Variant II Turbo 2.0 | 121 | 10.09 | -0.03 | 1.5 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 2.91
Bio-Rad Variant II Turbo | 40 | 10.9 |  | 2.4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad Variant II Turbo | 40 | 5.3 |  | 2.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II Turbo | 40 | 8.4 |  | 2.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant II Turbo | 54 | 6.9 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 54 | 6.9 |  | 2.2 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant II Turbo | 54 | 12 |  | 2.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 54 | 6.9 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 54 | 6.9 |  | 2.2 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant II Turbo | 54 | 12 |  | 2.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 95 | 5.4 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad Variant II Turbo | 95 | 12 |  | 2.6 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 95 | 7.8 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II Turbo | 103 | 9.4 |  | 2.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad Variant II Turbo | 103 | 11.4 |  | 2.6 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II Turbo | 103 | 6.3 |  | 2.4 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 4.704
Bio-Rad Variant II Turbo | 128 | 13.5 |  | 3.3 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 6.468
Bio-Rad Variant II Turbo | 128 | 5.5 |  | 2.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad Variant II Turbo | 128 | 9.3 |  | 2.7 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad Variant II Turbo | 132 | 12 |  | 2.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant II Turbo | 132 | 7 |  | 2.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 5.292
Bio-Rad Variant II Turbo | 132 | 5.2 |  | 3.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant II Turbo | 130 | 5.1 |  | 2.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 5.488
Bio-Rad Variant II Turbo | 130 | 8.5 |  | 2.5 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 4.9
Bio-Rad Variant II Turbo | 130 | 6 |  | 2.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant series | 38 | 6.1577 | 1.5 | 5.8 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Bio-Rad | 12.868
Bio-Rad Variant series | 40 | 7.5576 | 1.2 | 3.9 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Bio-Rad | 8.844
Bio-Rad Variant series | 86 | 6.0754 | 0.6 | 4.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Bio-Rad | 8.636
Bio-Rad Variant series | 86 | 7.5576 | 1.2 | 3.8 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Bio-Rad | 8.648
Bio-Rad Variant series | 37 | 6.4139 | 2.4 | 6.3 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Bio-Rad | 14.748
Bio-Rad Variant series | 37 | 7.6125 | 1.7 | 3.9 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Bio-Rad | 9.344
Bio-Rad Variant series | 112 | 6.3041 | 1.2 | 3.9 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Bio-Rad | 8.844
Bio-Rad Variant series | 112 | 7.5668 | 1.2 | 3 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Bio-Rad | 7.08
Bio-Rad Variant series | 47 | 7.8322 | 4 | 4.7 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Bio-Rad | 13.212
Bio-Rad Variant series | 47 | 6.3133 | 3.3 | 5.3 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Bio-Rad | 13.688
Bio-Rad Variant series | 147 | 7.6217 | 1.7 | 4.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Bio-Rad | 9.736
Bio-Rad Variant series | 148 | 6.0937 | 0.9 | 4.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Bio-Rad | 10.504
Bio-Rad Variant series | 66 | 8.8386 | 3 | 5.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Bio-Rad | 13.388
Bio-Rad Variant series | 67 | 5.6819 | 0.4 | 8.9 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Bio-Rad | 17.844
Bio-Rad Variant series | 162 | 8.8295 | 2.9 | 2.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Bio-Rad | 8.388
Bio-Rad Variant series | 163 | 5.6636 | 0.2 | 5.2 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Bio-Rad | 10.392
Bio-Rad Variant series | 89 | 5.5355 | 0.5 | 8.1 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 16.376
Bio-Rad Variant series | 85 | 5.5263 | 0.4 | 8.3 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Bio-Rad | 16.668
Bio-Rad Variant series | 121 | 5.4897 | 0 | 6.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 12.74
Bio-Rad Variant series | 120 | 5.4714 | -0.2 | 6.4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Bio-Rad | 12.344
Bio-Rad Variant II | 83 | 6.58 | 0 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 5.096
Bio-Rad Variant II | 83 | 8.51 | 0.12 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 5.608
Bio-Rad Variant II | 83 | 5.57 | -0.08 | 2.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 5.212
Bio-Rad Variant II | 84 | 6.44 | -0.05 | 2.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 4.458
Bio-Rad Variant II | 84 | 6.98 | 0.01 | 2.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 4.126
Bio-Rad Variant II | 84 | 9.8 | 0.15 | 2.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 5.246
Bio-Rad Variant II | 70 | 6.9 | 0.11 | 2.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 5.206
Bio-Rad Variant II | 70 | 10.43 | 0.15 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 5.246
Bio-Rad Variant II | 70 | 6.92 | 0.1 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 5.784
Bio-Rad Variant II | 60 | 8.77 | 0.14 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 5.432
Bio-Rad Variant II | 60 | 5.29 | -0.03 | 2.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 5.654
Bio-Rad Variant II | 60 | 5.94 | 0.1 | 3.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 6.372
Bio-Rad Variant II | 60 | 12.14 | 0.43 | 3.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 6.506
Bio-Rad Variant II | 60 | 9.92 | 0.39 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 6.466
Bio-Rad Variant II | 60 | 5.04 | 0 | 3.2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant II | 60 | 7.51 | 0.13 | 3.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 6.206
Bio-Rad Variant II | 52-65 | 12.12 | 0.43 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 6.506
Bio-Rad Variant II | 52-65 | 6.08 | 0.15 | 2.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 5.638
Bio-Rad Variant II | 52-65 | 5.21 | 0.04 | 3.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 6.508
Bio-Rad Variant II | 52-65 | 8.3 | 0.16 | 3.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 6.824
Bio-Rad Variant II | 52-65 | 9.53 | 0.23 | 2.8 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 5.718
Bio-Rad Variant II | 50-59 | 5.45 | 0.13 | 3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 6.01
Bio-Rad Variant II | 50-59 | 9.53 | 0.37 | 2.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 5.074
Bio-Rad Variant II | 50-59 | 5.45 | 0.14 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 5.04
Bio-Rad Variant II | 50-59 | 12.55 | 0.52 | 2.4 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 5.224
Bio-Rad Variant II | 50-59 | 6.1 | 0.16 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 4.668
Bio-Rad Variant II | 44 | 5.39 | 0.12 | 3.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 6.196
Bio-Rad Variant II | 44 | 10.85 | 0.26 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 4.964
Bio-Rad Variant II | 44 | 6.37 | 0.17 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 4.482
Bio-Rad Variant II | 44 | 12.54 | 0.31 | 3 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 6.19
Bio-Rad Variant II | 44 | 7.74 | 0.23 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 4.542
Bio-Rad Variant II | 44-53 | 9.23 | 0.12 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 4.432
Bio-Rad Variant II | 44-53 | 6.12 | 0.11 | 2.7 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 5.402
Bio-Rad Variant II | 44-53 | 11.91 | 0.2 | 2.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 4.708
Bio-Rad Variant II | 44-53 | 5.11 | 0.09 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 5.578
Bio-Rad Variant II | 44-53 | 7.67 | 0.09 | 1.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 3.422
Bio-Rad Variant II | 43-50 | 6.51 | 0.1 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 5.196
Bio-Rad Variant II | 43-50 | 9.75 | 0.22 | 2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 4.14
Bio-Rad Variant II | 43-50 | 5.34 | 0 | 2.2 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 4.312
Bio-Rad Variant II | 43-50 | 8.73 | 0.22 | 1.9 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 3.944
Bio-Rad Variant II | 43-50 | 7.29 | 0.04 | 2.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 4.156
Bio-Rad Variant II | 34 | 7.18 | 0.03 | 1.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 3.754
Bio-Rad Variant II | 34 | 5.19 | 0 | 1.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 3.528
Bio-Rad Variant II | 34 | 8.47 | 0.05 | 1.7 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 3.382
Bio-Rad Variant II | 34 | 9.81 | 0.02 | 1.9 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 3.744
Bio-Rad Variant II | 34 | 6.23 | 0.11 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 4.814
Bio-Rad Variant II | 26 | 6.39 | 0.08 | 1.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 3.412
Bio-Rad Variant II | 26 | 9.17 | 0.06 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II | 26 | 5.31 | 0 | 3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II | 26 | 8.4 | 0.05 | 1.6 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 3.186
Bio-Rad Variant II | 26 | 7.7 | 0.08 | 2.1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 4.196
Bio-Rad Variant II | 38 | 7.59 | 0.17 | 2.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 4.482
Bio-Rad Variant II | 38 | 5.33 | 0.13 | 1.5 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 3.07
Bio-Rad Variant II | 38 | 8.42 | 0.11 | 2.5 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 5.01
Bio-Rad Variant II | 38 | 10.56 | 0.23 | 2.7 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 5.522
Bio-Rad Variant II | 38 | 6.01 | 0.14 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 3.864
Bio-Rad Variant II | 39 | 7.32 | 0.07 | 2.4 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 4.774
Bio-Rad Variant II | 39 | 6.92 | 0.11 | 2.5 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 5.01
Bio-Rad Variant II | 39 | 8.66 | 0.01 | 2.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 4.714
Bio-Rad Variant II | 39 | 9.56 | 0.06 | 2.6 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 5.156
Bio-Rad Variant II | 39 | 5.47 | 0.02 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 5.312
Bio-Rad Variant II Turbo | 127 | 6.7 | 0.12 | 2.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 4.432
Bio-Rad Variant II Turbo | 127 | 8.56 | 0.17 | 2.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 4.286
Bio-Rad Variant II Turbo | 127 | 5.7 | 0.05 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 5.146
Bio-Rad Variant II Turbo | 151 | 6.57 | 0.08 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 4.98
Bio-Rad Variant II Turbo | 151 | 7.13 | 0.16 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 4.472
Bio-Rad Variant II Turbo | 151 | 9.84 | 0.19 | 2.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 4.502
Bio-Rad Variant II Turbo | 104 | 6.79 | 0 | 3.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 6.272
Bio-Rad Variant II Turbo | 104 | 10.33 | 0.05 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 4.95
Bio-Rad Variant II Turbo | 104 | 6.82 | 0 | 3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 5.88
Bio-Rad Variant II Turbo | 85 | 8.69 | 0.06 | 2.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 4.96
Bio-Rad Variant II Turbo | 85 | 5.3 | -0.02 | 2.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 5.664
Bio-Rad Variant II Turbo | 82 | 5.8 | -0.04 | 1.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 3.292
Bio-Rad Variant II Turbo | 82 | 11.73 | 0.02 | 2.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 4.92
Bio-Rad Variant II Turbo | 82 | 9.6 | 0.07 | 2.4 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 4.774
Bio-Rad Variant II Turbo | 82 | 5.01 | -0.03 | 2.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 4.478
Bio-Rad Variant II Turbo | 82 | 7.38 | 0 | 2.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad Variant II Turbo | 83-99 | 11.73 | 0.04 | 2.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 4.744
Bio-Rad Variant II Turbo | 83-99 | 5.92 | -0.01 | 1.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 3.714
Bio-Rad Variant II Turbo | 83-99 | 5.15 | -0.02 | 2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 3.9
Bio-Rad Variant II Turbo | 83-99 | 8.15 | 0.01 | 2.1 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 4.126
Bio-Rad Variant II Turbo | 83-99 | 9.32 | 0.02 | 2.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 4.528
Bio-Rad Variant II Turbo | 82-89 | 5.34 | 0.02 | 2.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 4.92
Bio-Rad Variant II Turbo | 82-89 | 9.34 | 0.18 | 1.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 3.904
Bio-Rad Variant II Turbo | 82-89 | 5.33 | 0.02 | 2.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 4.136
Bio-Rad Variant II Turbo | 82-89 | 12.3 | 0.27 | 2.2 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 4.582
Bio-Rad Variant II Turbo | 82-89 | 5.98 | 0.04 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 3.176
Bio-Rad Variant II Turbo | 79 | 5.27 | 0 | 2.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II Turbo | 79 | 10.73 | 0.14 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 4.844
Bio-Rad Variant II Turbo | 79 | 6.24 | 0.04 | 2.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 5.136
Bio-Rad Variant II Turbo | 79 | 12.36 | 0.13 | 2.8 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 5.618
Bio-Rad Variant II Turbo | 79 | 7.66 | 0.15 | 2.4 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 4.854
Bio-Rad Variant II Turbo | 77-84 | 9.16 | 0.05 | 2.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 4.754
Bio-Rad Variant II Turbo | 77-84 | 5.99 | -0.02 | 2.1 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 4.096
Bio-Rad Variant II Turbo | 77-84 | 11.73 | 0.02 | 2.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 5.508
Bio-Rad Variant II Turbo | 77-84 | 5 | -0.02 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 4.488
Bio-Rad Variant II Turbo | 77-84 | 7.65 | 0.07 | 2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 3.99
Bio-Rad Variant II Turbo | 73-78 | 6.48 | 0.07 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 5.166
Bio-Rad Variant II Turbo | 73-78 | 9.84 | 0.31 | 2.3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 4.818
Bio-Rad Variant II Turbo | 73-78 | 5.36 | 0.02 | 2.8 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 5.508
Bio-Rad Variant II Turbo | 73-78 | 8.81 | 0.3 | 2.6 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 5.396
Bio-Rad Variant II Turbo | 73-78 | 7.4 | 0.15 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 5.05
Bio-Rad Variant II Turbo | 62 | 7.12 | -0.03 | 2.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 4.282
Bio-Rad Variant II Turbo | 62 | 5.13 | -0.06 | 2.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 4.84
Bio-Rad Variant II Turbo | 62 | 8.48 | 0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 4.764
Bio-Rad Variant II Turbo | 62 | 9.83 | 0.04 | 2.7 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 5.332
Bio-Rad Variant II Turbo | 62 | 6.13 | 0.01 | 2.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 5.694
Bio-Rad Variant II Turbo | 56 | 6.29 | -0.02 | 3.3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 6.448
Bio-Rad Variant II Turbo | 56 | 9.15 | 0.04 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 5.136
Bio-Rad Variant II Turbo | 56 | 5.21 | -0.1 | 3.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 6.172
Bio-Rad Variant II Turbo | 56 | 8.41 | 0.06 | 2.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 4.764
Bio-Rad Variant II Turbo | 56 | 7.68 | 0.06 | 2.4 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 4.764
Bio-Rad Variant II Turbo | 62 | 7.51 | 0.09 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 3.422
Bio-Rad Variant II Turbo | 62 | 5.26 | 0.06 | 2.1 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 4.176
Bio-Rad Variant II Turbo | 62 | 8.42 | 0.11 | 2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 4.03
Bio-Rad Variant II Turbo | 62 | 10.5 | 0.17 | 2.2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 4.482
Bio-Rad Variant II Turbo | 62 | 5.9 | 0.03 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 3.754
Bio-Rad Variant II Turbo | 67 | 7.31 | 0.06 | 2.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 4.96
Bio-Rad Variant II Turbo | 67 | 6.85 | 0.04 | 2.3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 4.548
Bio-Rad Variant II Turbo | 67 | 8.72 | 0.07 | 2.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 5.558
Bio-Rad Variant II Turbo | 67 | 9.66 | 0.16 | 2.9 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 5.844
Bio-Rad Variant II Turbo | 67 | 5.41 | -0.04 | 2.5 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 4.86
Bio-Rad Variant II Turbo 2.0 | 124 | 6.64 | 0.06 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 5.156
Bio-Rad Variant II Turbo 2.0 | 124 | 8.54 | 0.15 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 4.658
Bio-Rad Variant II Turbo 2.0 | 124 | 5.65 | 0 | 3.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 6.076
Bio-Rad Variant II Turbo 2.0 | 103 | 6.45 | -0.04 | 2 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 3.88
Bio-Rad Variant II Turbo 2.0 | 103 | 7.03 | 0.06 | 2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 3.98
Bio-Rad Variant II Turbo 2.0 | 103 | 9.71 | 0.06 | 1.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 3.392
Bio-Rad Variant II Turbo 2.0 | 144 | 6.89 | 0.1 | 2.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 4.608
Bio-Rad Variant II Turbo 2.0 | 144 | 10.26 | -0.02 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 4.88
Bio-Rad Variant II Turbo 2.0 | 144 | 6.88 | 0.06 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 5.156
Bio-Rad Variant II Turbo 2.0 | 121 | 8.67 | 0.04 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 5.332
Bio-Rad Variant II Turbo 2.0 | 121 | 5.25 | -0.07 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 5.222
Bio-Rad Variant II Turbo 2.0 | 132 | 5.84 | 0 | 2.3 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II Turbo 2.0 | 132 | 11.57 | -0.14 | 2.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 4.368
Bio-Rad Variant II Turbo 2.0 | 132 | 9.55 | 0.02 | 2.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 4.92
Bio-Rad Variant II Turbo 2.0 | 132 | 5.01 | -0.03 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 5.262
Bio-Rad Variant II Turbo 2.0 | 132 | 7.35 | -0.03 | 2.3 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 4.478
Bio-Rad Variant II Turbo 2.0 | 141-171 | 11.59 | -0.1 | 2.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 4.604
Bio-Rad Variant II Turbo 2.0 | 141-171 | 5.92 | -0.01 | 2 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 3.91
Bio-Rad Variant II Turbo 2.0 | 141-171 | 5.14 | -0.03 | 2.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 4.282
Bio-Rad Variant II Turbo 2.0 | 141-171 | 8.09 | -0.05 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 3.086
Bio-Rad Variant II Turbo 2.0 | 141-171 | 9.18 | -0.12 | 1.6 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 3.016
Bio-Rad Variant II Turbo 2.0 | 146-173 | 5.37 | 0.05 | 2.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 4.166
Bio-Rad Variant II Turbo 2.0 | 146-173 | 9.24 | 0.08 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 3.412
Bio-Rad Variant II Turbo 2.0 | 146-173 | 5.35 | 0.04 | 2.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 4.548
Bio-Rad Variant II Turbo 2.0 | 146-173 | 12 | -0.03 | 1.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 3.302
Bio-Rad Variant II Turbo 2.0 | 146-173 | 5.99 | 0.05 | 1.9 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 3.774
Bio-Rad Variant II Turbo 2.0 | 145 | 5.23 | -0.04 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 5.056
Bio-Rad Variant II Turbo 2.0 | 145 | 10.62 | 0.03 | 1.9 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 3.754
Bio-Rad Variant II Turbo 2.0 | 145 | 6.23 | 0.03 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 4.342
Bio-Rad Variant II Turbo 2.0 | 145 | 12.09 | -0.14 | 2.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 4.172
Bio-Rad Variant II Turbo 2.0 | 145 | 7.53 | 0.02 | 2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 3.94
Bio-Rad Variant II Turbo 2.0 | 147-173 | 9.13 | 0.02 | 2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 3.94
Bio-Rad Variant II Turbo 2.0 | 147-173 | 5.99 | -0.02 | 1.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 3.508
Bio-Rad Variant II Turbo 2.0 | 147-173 | 11.49 | -0.22 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 3.7
Bio-Rad Variant II Turbo 2.0 | 147-173 | 4.98 | -0.04 | 2 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 3.88
Bio-Rad Variant II Turbo 2.0 | 147-173 | 7.54 | -0.04 | 1.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 3.292
Bio-Rad Variant II Turbo 2.0 | 138-157 | 6.41 | 0 | 1.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 3.332
Bio-Rad Variant II Turbo 2.0 | 138-157 | 9.63 | 0.1 | 1.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 3.432
Bio-Rad Variant II Turbo 2.0 | 138-157 | 5.31 | -0.03 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 4.086
Bio-Rad Variant II Turbo 2.0 | 138-157 | 8.57 | 0.06 | 1.6 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 3.196
Bio-Rad Variant II Turbo 2.0 | 138-157 | 7.24 | -0.01 | 1.6 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 3.126
Bio-Rad Variant II Turbo 2.0 | 142 | 7.5 | 0.08 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 4.98
Bio-Rad Variant II Turbo 2.0 | 142 | 5.3 | 0.1 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 5.196
Bio-Rad Variant II Turbo 2.0 | 142 | 8.31 | 0 | 2.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 4.116
Bio-Rad Variant II Turbo 2.0 | 142 | 10.4 | 0.07 | 2.1 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 4.186
Bio-Rad Variant II Turbo 2.0 | 142 | 5.92 | 0.05 | 2.2 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 4.362
Bio-Rad Variant II Turbo 2.0 | 152 | 7.29 | 0.04 | 2 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 3.96
Bio-Rad Variant II Turbo 2.0 | 152 | 6.86 | 0.05 | 2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 3.97
Bio-Rad Variant II Turbo 2.0 | 152 | 8.58 | -0.07 | 1.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 3.458
Bio-Rad Variant II Turbo 2.0 | 152 | 9.54 | 0.04 | 1.9 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 3.764
Bio-Rad Variant II Turbo 2.0 | 152 | 5.49 | 0.04 | 2.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 4.352
Bio-Rad Variant II Turbo 2.0 | 142 | 7.16 | 0.01 | 1.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 3.538
Bio-Rad Variant II Turbo 2.0 | 142 | 5.11 | -0.08 | 2.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 5.212
Bio-Rad Variant II Turbo 2.0 | 142 | 8.46 | 0.04 | 2.2 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 4.352
Bio-Rad Variant II Turbo 2.0 | 142 | 9.75 | -0.04 | 2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 3.88
Bio-Rad Variant II Turbo 2.0 | 142 | 6.13 | 0.01 | 1.8 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 3.538
Bio-Rad Variant II Turbo 2.0 | 115 | 6.31 | 0 | 2.3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 4.508
Bio-Rad Variant II Turbo 2.0 | 115 | 9.11 | 0 | 2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 3.92
Bio-Rad Variant II Turbo 2.0 | 115 | 5.2 | -0.11 | 2.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 5.574
Bio-Rad Variant II Turbo 2.0 | 115 | 8.36 | 0.01 | 2.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 4.126
Bio-Rad Variant II Turbo 2.0 | 115 | 7.58 | -0.04 | 1.9 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 3.684
Dade Behring Dimension | 490 | 10.5 |  | 3.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Dade Behring | 6.076
Dade Behring Dimension | 490 | 5.6 |  | 3.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Dade Behring | 7.056
Dade Behring Dimension | 490 | 8.1 |  | 3.2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Dade Behring | 6.272
Dade Behring Dimension | 494 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Dade Behring | 7.252
Dade Behring Dimension | 494 | 7.1 |  | 3.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Dade Behring | 7.056
Dade Behring Dimension | 494 | 11.5 |  | 3.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Dade Behring | 7.252
Dade Behring Dimension | 494 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Dade Behring | 7.252
Dade Behring Dimension | 494 | 7.1 |  | 3.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Dade Behring | 7.056
Dade Behring Dimension | 494 | 11.5 |  | 3.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Dade Behring | 7.252
Dade Behring Dimension | 593 | 5.8 |  | 3.1 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Dade Behring | 6.076
Dade Behring Dimension | 593 | 11.3 |  | 3.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Dade Behring | 6.86
Dade Behring Dimension | 593 | 7.6 |  | 2.8 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Dade Behring | 5.488
Dade Behring Dimension | 527 | 8.9 |  | 3.1 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Dade Behring | 6.076
Dade Behring Dimension | 527 | 10.9 |  | 3.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Dade Behring | 6.076
Dade Behring Dimension | 527 | 6.3 |  | 3.3 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Dade Behring | 6.468
Dade Behring Dimension | 574 | 12.9 |  | 4.4 | GH2-01 | 13 | 2008 | CAP | Fresh | Siemens | 8.624
Dade Behring Dimension | 574 | 5.6 |  | 3.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Siemens | 6.076
Dade Behring Dimension | 574 | 8.6 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Siemens | 5.88
Dade Behring Dimension | 578 | 12.3 |  | 4.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 9.016
Dade Behring Dimension | 578 | 6.9 |  | 3.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 7.252
Dade Behring Dimension | 578 | 5.4 |  | 3.5 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 6.86
Siemens Dimension Vista | 13 | 11.4 |  | 2.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 4.704
Siemens Dimension Vista | 13 | 7.2 |  | 4.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 9.016
Siemens Dimension Vista | 13 | 5 |  | 4.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 9.212
Diasys InnovaStar | 7 | 7.2191 | -2.7 | 6.7 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Diasys | 10.432
Diasys InnovaStar | 7 | 5.9656 | -0.5 | 3.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Diasys | 6.948
EKF Diagnostics | 11 | 8.6831 | 1.3 | 3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | EKF Diagnostics | 7.18
EKF Diagnostics | 11 | 5.7551 | 1.2 | 5.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | EKF Diagnostics | 11.784
EKF Diagnostics | 18 | 5.5995 | 1.2 | 8.2 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | EKF Diagnostics | 17.272
EKF Diagnostics | 18 | 5.627 | 1.5 | 6.2 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | EKF Diagnostics | 13.652
HemoCue HbA1c 501 | 38 | 7.3106 | -1.7 | 7 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | HemoCue | 12.02
HemoCue HbA1c 501 | 38 | 5.9198 | -1 | 8 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | HemoCue | 14.68
HemoCue HbA1c 501 | 14 | 8.4086 | -1.7 | 4.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | HemoCue | 6.728
HemoCue HbA1c 501 | 14 | 5.3891 | -2.8 | 8.9 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | HemoCue | 14.644
HemoCue HbA1c 501 | 13 | 5.4165 | -0.8 | 9 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | HemoCue | 16.84
HemoCue HbA1c 501 | 13 | 5.4897 | 0 | 9.4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | HemoCue | 18.424
JEOL BioMajesty JCA-BM series | 8-10 | 11.48 | -0.21 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | JEOL | 5.474
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-12 | 5.93 | 2015 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-13 | 5.17 | 2015 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-14 | 8.14 | 2015 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-15 | 9.3 | 2015 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 8-10 | 5.28 | -0.04 | 2.8 | GH5-01 | 5.32 | 2016 | CAP | Fresh | JEOL | 5.448
JEOL BioMajesty JCA-BM series | 8-10 | 9.1 | -0.06 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | JEOL | 3.272
JEOL BioMajesty JCA-BM series | 8-10 | 5.3 | -0.01 | 2.8 | GH5-03 | 5.31 | 2016 | CAP | Fresh | JEOL | 5.478
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 8-10 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 7-10 | 8.97 | -0.14 | 1.8 | GH5-11 | 9.11 | 2016 | CAP | Fresh | JEOL | 3.388
JEOL BioMajesty JCA-BM series | 7-10 |  |  |  | GH5-12 | 6.01 | 2016 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 7-10 | 12.02 | 0.31 | 1.7 | GH5-13 | 11.71 | 2016 | CAP | Fresh | JEOL | 3.642
JEOL BioMajesty JCA-BM series | 7-10 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | JEOL | 0
JEOL BioMajesty JCA-BM series | 7-10 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | JEOL | 0
Lifotronic | 21 | 5.5538 | 0.7 | 8.8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Lifotronic | 17.948
Lifotronic | 21 | 5.508 | 0.2 | 9.9 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Lifotronic | 19.604
Lifotronic | 13 | 5.6361 | 1.6 | 8.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Lifotronic | 18.26
Lifotronic | 12 | 5.7368 | 2.7 | 3.6 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Lifotronic | 9.756
Medinor | 6 | 5.6453 | -4.1 | 12.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Medinor | 19.616
Medinor | 6 | 6.9629 | -5.3 | 17.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Medinor | 28.216
Medinor | 10 | 6.2401 | 0.5 | 8.3 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Medinor | 16.768
Medinor | 10.0 | 7.457 | 0 | 6.9 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Medinor | 13.524
Medinor NycoCard | 9 | 7.5851 | 1.3 | 8.6 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Medinor | 18.156
Medinor NycoCard | 9 | 6.3316 | 3.5 | 7.6 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Medinor | 18.396
Arkray Adams HA-8160 series | 75 | 6.2401 | 0.5 | 4.6 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Menarini | 9.516
Arkray Adams HA-8160 series | 74 | 7.4204 | -0.4 | 3.9 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Menarini | 7.244
Arkray Adams HA-8160 series | 76 | 6.295 | 1.1 | 3.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Menarini | 7.176
Arkray Adams HA-8160 series | 76 | 7.5393 | 0.9 | 2.6 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Menarini | 5.996
Arkray Adams HA-8160 series | 66 | 7.4204 | -0.5 | 4.2 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Menarini | 7.732
Arkray Adams HA-8160 series | 66 | 6.0022 | -0.1 | 4.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Menarini | 8.524
Arkray Adams HA-8160 series | 82 | 7.4844 | 0.2 | 3.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Menarini | 7.648
Arkray Adams HA-8160 series | 81 | 6.0113 | 0 | 3.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Menarini | 7.448
Arkray Adams HA-8160 series | 59 | 8.6282 | 0.7 | 3.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Menarini | 7.756
Arkray Adams HA-8160 series | 62 | 5.6087 | -0.4 | 5.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Menarini | 9.596
Arkray Adams HA-8160 series | 74 | 8.619 | 0.6 | 3.2 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Menarini | 6.872
Arkray Adams HA-8160 series | 73 | 5.5904 | -0.6 | 4.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Menarini | 8.024
Arkray Adams HA-8160 series | 49 | 5.5446 | 0.6 | 4.9 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 10.204
Arkray Adams HA-8160 series | 50 | 5.5172 | 0.3 | 5.1 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 10.296
Arkray Adams HA-8160 series | 51 | 5.5172 | 0.3 | 5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 10.1
Arkray Adams HA-8160 series | 50 | 5.508 | 0.2 | 4.9 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 9.804
Arkray Adams HA-8160 series | 87 | 5.993 | -0.3 | 2.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Menarini | 5.384
Arkray Adams HA-8160 series | 92 | 7.3655 | -0.9 | 3 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Menarini | 4.98
Arkray Adams HA-8160 series | 91 | 6.0388 | 0.2 | 3.7 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Menarini | 7.452
Arkray Adams HA-8160 series | 91 | 7.5119 | 0.7 | 3.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Menarini | 6.776
Arkray Adams HA-8180 series | 78 | 6.2858 | 1 | 3.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Menarini | 7.076
Arkray Adams HA-8180 series | 78 | 7.4387 | -0.2 | 3.2 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Menarini | 6.072
Arkray Adams HA-8180 series | 100 | 6.2767 | 0.9 | 3.2 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Menarini | 7.172
Arkray Adams HA-8180 series | 100 | 7.521 | 0.7 | 2.7 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Menarini | 5.992
Arkray Adams HA-8180 series | 121 | 7.3746 | -1 | 4.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Menarini | 7.82
Arkray Adams HA-8180 series | 116 | 5.9747 | -0.4 | 5.2 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Menarini | 9.792
Arkray Adams HA-8180 series | 144 | 7.5302 | 0.7 | 2.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Menarini | 6.188
Arkray Adams HA-8180 series | 144 | 6.0845 | 0.8 | 3 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Menarini | 6.68
Arkray Adams HA-8180 series | 99 | 8.372 | -2.1 | 5.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Menarini | 9.072
Arkray Adams HA-8180 series | 102 | 5.5538 | -1 | 5.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Menarini | 9.584
Arkray Adams HA-8180 series | 214 | 8.6465 | 0.9 | 2.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Menarini | 6.388
Arkray Adams HA-8180 series | 213 | 5.6361 | -0.1 | 3.8 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Menarini | 7.348
Arkray Adams HA-8180 series | 99 | 5.5355 | 0.5 | 4.2 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 8.732
Arkray Adams HA-8180 series | 104 | 5.5355 | 0.5 | 4.2 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 8.732
Arkray Adams HA-8180 series | 184 | 5.5995 | 1.2 | 3.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 7.276
Arkray Adams HA-8180 series | 185 | 5.6087 | 1.3 | 3.2 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 7.572
Arkray Adams HA-8180 series | 72 | 5.9656 | -0.6 | 3.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Menarini | 7.044
Arkray Adams HA-8180 series | 72 | 7.3655 | -0.9 | 3.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Menarini | 5.176
Arkray Adams HA-8180 series | 82 | 6.0205 | 0 | 3 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Menarini | 5.88
Arkray Adams HA-8180 series | 82 | 7.5119 | 0.7 | 2.9 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Menarini | 6.384
Menarini (ARKRAY) not specified/other | 24 | 8.5275 | -0.4 | 4.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Menarini | 8.616
Menarini (ARKRAY) not specified/other | 24 | 5.5446 | -1.1 | 3.9 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Menarini | 6.544
Menarini (ARKRAY) not specified/other | 15 | 5.4806 | -0.1 | 4.8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 9.308
Menarini (ARKRAY) not specified/other | 15 | 5.4714 | -0.2 | 5 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Menarini | 9.6
Menarini (ARKRAY) not specified/other | 12 | 5.5446 | 0.6 | 3.8 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 8.048
Menarini (ARKRAY) not specified/other | 12 | 5.5812 | 1 | 3.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Menarini | 7.076
Menarini (ARKRAY) not specified/other | 48 | 7.3472 | -1.3 | 3 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Menarini | 4.58
Menarini (ARKRAY) not specified/other | 47 | 5.9473 | -0.7 | 3.2 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Menarini | 5.572
Menarini (ARKRAY) not specified/other | 36 | 6.2584 | 0.7 | 3.5 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Menarini | 7.56
Menarini (ARKRAY) not specified/other | 36 | 7.3929 | -0.7 | 3.9 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Menarini | 6.944
Menarini (ARKRAY) not specified/other | 6 | 7.4295 | -0.4 | 3.6 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Menarini | 6.656
Menarini (ARKRAY) not specified/other | 6 | 6.0113 | 0 | 5.1 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Menarini | 9.996
A1cNOW | 15 | 10.4 |  | 8.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Metrika | 15.876
A1cNOW | 15 | 5.3 |  | 6.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Metrika | 12.348
A1cNOW | 15 | 8.2 |  | 7.6 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Metrika | 14.896
A1cNOW | 18 | 6.7 |  | 7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Metrika | 13.72
A1cNOW | 18 | 6.7 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Metrika | 15.092
A1cNOW | 18 | 11.4 |  | 7.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Metrika | 14.7
A1cNOW | 18 | 6.7 |  | 7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Metrika | 13.72
A1cNOW | 18 | 6.7 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Metrika | 15.092
A1cNOW | 18 | 11.4 |  | 7.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Metrika | 14.7
A1cNOW | 23 | 5.3 |  | 6.1 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Metrika | 11.956
A1cNOW | 23 | 11.7 |  | 6.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Metrika | 12.348
A1cNOW | 23 | 7.4 |  | 7.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Metrika | 15.092
A1cNOW | 19 | 8.2 |  | 7.4 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Metrika | 14.504
A1cNOW | 19 | 10.2 |  | 6.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Metrika | 12.74
A1cNOW | 19 | 5.7 |  | 5.7 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Metrika | 11.172
A1cNOW | 19 | 11.9 |  | 5.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Metrika | 10.976
A1cNOW | 19 | 4.9 |  | 3.7 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Metrika | 7.252
A1cNOW | 19 | 8.2 |  | 6.6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Metrika | 12.936
A1cNOW | 20 | 10.4 |  | 6.9 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Metrika | 13.524
A1cNOW | 20 | 6.1 |  | 7.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Metrika | 15.092
A1cNOW | 20 | 4.8 |  | 7 | GH2-06 | 5 | 2008 | CAP | Fresh | Metrika | 13.72
A1cNOW | 17 | 4.6 |  | 5 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Metrika | 9.8
A1cNOW | 17 | 7.2 |  | 6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Metrika | 11.76
A1cNOW | 17 | 5.3 |  | 5.2 | GH2-03 | 6 | 2009 | CAP | Fresh | Metrika | 10.192
A1cNOW | 19 | 5.98 | -0.62 | 7.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Metrika | 14.864
A1cNOW | 19 | 8.51 | -0.99 | 6.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Metrika | 11.75
A1cNOW | 19 | 6.72 | -0.68 | 6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Metrika | 11.08
A1cNOW | 21 | 5.63 | -0.27 | 5.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Metrika | 10.706
A1cNOW | 21 | 9.09 | -0.71 | 4.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Metrika | 8.306
A1cNOW | 21 | 6.84 | -0.56 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Metrika | 6.3
Mindray bs series | 8 | 5.6087 | 1.3 | 12.4 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Mindray | 25.604
Mindray bs series | 8 | 5.6544 | 1.8 | 11.8 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Mindray | 24.928
Not specified/Other | 14 | 5.9564 | -0.7 | 8.7 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Not specified | 16.352
Not specified/Other | 18 | 7.3655 | -0.9 | 7.4 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Not specified | 13.604
Not specified/Other | 123 | 6.0113 | -0.1 | 5.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Not specified | 10.68
Not specified/Other | 123 | 7.4661 | 0.2 | 5.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Not specified | 10.196
Not specified/Other | 24 | 6.5145 | 3.5 | 10.2 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Not specified | 23.492
Not specified/Other | 17 | 7.5851 | 1.4 | 9.5 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Not specified | 20.02
Not specified/Other | 113 | 6.2584 | 0.7 | 5.6 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Not specified | 11.676
Not specified/Other | 111 | 7.457 | 0 | 6.6 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Not specified | 12.936
Not specified/Other | 127 | 7.4478 | -0.2 | 6.4 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Not specified | 12.344
Not specified/Other | 124 | 6.0571 | 0.5 | 6 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Not specified | 12.26
Not specified/Other | 11 | 8.6922 | 1.4 | 4.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Not specified | 10.416
Not specified/Other | 13 | 5.5812 | -0.7 | 8.7 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Not specified | 16.352
Not specified/Other | 96 | 8.4909 | -0.8 | 6.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Not specified | 12.528
Not specified/Other | 94 | 5.5538 | -1 | 7 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Not specified | 12.72
Not specified/Other | 21 | 5.5263 | 0.4 | 12 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Not specified | 23.92
Not specified/Other | 21 | 5.5995 | 1.2 | 8.9 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Not specified | 18.644
Not specified/Other | 213 | 5.5263 | 0.4 | 5.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Not specified | 11.18
Not specified/Other | 212 | 5.5355 | 0.5 | 4.9 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Not specified | 10.104
Not specified/Other | 14 | 7.3289 | -1.5 | 4.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Not specified | 8.104
Not specified/Other | 14 | 5.9381 | -0.8 | 5.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Not specified | 10.568
Olympus AU 400/400e | 15 | 4.8 |  | 5.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Olympus | 10.388
Olympus AU 400/400e | 15 | 8.1 |  | 7.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Olympus | 14.308
Olympus AU 400/400e | 15 | 5.6 |  | 5.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Olympus | 11.172
Olympus AU system | 22 | 11.2 |  | 4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Olympus | 7.84
Olympus AU system | 22 | 5.5 |  | 3.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Olympus | 7.644
Olympus AU system | 22 | 8.9 |  | 5.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Olympus | 10.78
Olympus AU system | 26 | 7.3 |  | 4.3 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Olympus | 8.428
Olympus AU system | 26 | 7.2 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Olympus | 8.428
Olympus AU system | 26 | 12.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Olympus | 8.82
Olympus AU system | 26 | 7.3 |  | 4.3 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Olympus | 8.428
Olympus AU system | 26 | 7.2 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Olympus | 8.428
Olympus AU system | 26 | 12.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Olympus | 8.82
Olympus AU system | 30 | 5.6 |  | 6.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Olympus | 12.152
Olympus AU system | 30 | 12.2 |  | 6.1 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Olympus | 11.956
Olympus AU system | 30 | 8.2 |  | 5.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Olympus | 10.78
Olympus AU system | 25 | 9.4 |  | 3.6 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Olympus | 7.056
Olympus AU system | 25 | 11 |  | 4.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Olympus | 8.036
Olympus AU system | 25 | 6.1 |  | 3.6 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Olympus | 7.056
Olympus AU system | 30 | 12.8 |  | 6 | GH2-01 | 13 | 2008 | CAP | Fresh | Olympus | 11.76
Olympus AU system | 30 | 5.2 |  | 3.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Olympus | 6.86
Olympus AU system | 30 | 8.9 |  | 5.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Olympus | 10.584
Olympus AU system | 25 | 11.9 |  | 6.3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Olympus | 12.348
Olympus AU system | 25 | 6.7 |  | 4.9 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Olympus | 9.604
Olympus AU system | 25 | 4.9 |  | 4 | GH2-06 | 5 | 2008 | CAP | Fresh | Olympus | 7.84
Olympus AU system | 25 | 6.46 | -0.14 | 4.4 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Olympus | 8.484
Olympus AU system | 25 | 9.24 | -0.26 | 4.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Olympus | 9.148
Olympus AU system | 25 | 7.32 | -0.08 | 4.8 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Olympus | 9.328
Olympus AU system | 19 | 5.74 | -0.16 | 3.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Olympus | 7.484
Olympus AU system | 19 | 9.67 | -0.13 | 5.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Olympus | 9.866
Olympus AU system | 19 | 7.39 | -0.01 | 4.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Olympus | 8.81
Vitros 5.1 FS. 4600. 5600 | 7 | 8.9759 | 4.5 | 4.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Vitros | 13.908
Vitros 5.1 FS. 4600. 5600 | 7 | 5.8649 | 2.4 | 5 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Vitros | 12.2
Thermo Fisher Scientific/Konelab | 5 | 5 |  |  | GH2-01 | 5.4 | 2007 | CAP | Fresh | Fisher | 0
Thermo Fisher Scientific/Konelab | 5 | 11 |  |  | GH2-02 | 11.6 | 2007 | CAP | Fresh | Fisher | 0
Thermo Fisher Scientific/Konelab | 5 | 7.5 |  |  | GH2-03 | 7.6 | 2007 | CAP | Fresh | Fisher | 0
Primus | 9 | 15.2 |  |  | GH2-04 | 10.7 | 2006 | CAP | Fresh | Primus | 0
Primus | 9 | 6 |  |  | GH2-05 | 5.30 | 2006 | CAP | Fresh | Primus | 0
Primus | 9 | 11.2 |  |  | GH2-06 | 8.40 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 8.3 |  |  | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 8.3 |  |  | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 16.9 |  |  | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 0
Primus | 15 | 6.2 |  | 4.3 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Primus | 8.428
Primus | 15 | 17.1 |  | 2.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Primus | 4.508
Primus | 15 | 10 |  | 2.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Primus | 5.684
Primus | 16 | 12.1 |  | 5.2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Primus | 10.192
Primus | 16 | 15.4 |  | 5.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Primus | 9.996
Primus | 16 | 7.3 |  | 5.1 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Primus | 9.996
Primus | 16 | 19.2 |  | 5.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Primus | 9.996
Primus | 16 | 6.1 |  | 5.7 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Primus | 11.172
Primus | 16 | 11.8 |  | 6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Primus | 11.76
Primus | 20 | 17.5 |  | 2.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Primus | 4.116
Primus | 20 | 8.7 |  | 1.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Primus | 3.332
Primus | 20 | 5.7 |  | 4.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Primus | 8.232
Primus HPLC (affinity) | 25 | 10.9 |  | 3.2 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Primus | 6.272
Primus HPLC (affinity) | 25 | 5.1 |  | 3.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Primus | 6.86
Primus HPLC (affinity) | 25 | 8.4 |  | 3.3 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Primus | 6.468
Primus HPLC (affinity) | 27 | 6.5 |  | 4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 7.84
Primus HPLC (affinity) | 27 | 6.5 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 7.644
Primus HPLC (affinity) | 27 | 11.5 |  | 3.2 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 6.272
Primus HPLC (affinity) | 27 | 6.5 |  | 4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 7.84
Primus HPLC (affinity) | 27 | 6.5 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 7.644
Primus HPLC (affinity) | 27 | 11.5 |  | 3.2 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 6.272
Primus HPLC (affinity) | 36 | 5.3 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Primus | 5.292
Primus HPLC (affinity) | 36 | 11.7 |  | 2.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Primus | 4.9
Primus HPLC (affinity) | 36 | 7.6 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Primus | 4.508
Primus HPLC (affinity) | 25 | 9 |  | 4.8 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Primus | 9.408
Primus HPLC (affinity) | 25 | 11 |  | 4.7 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Primus | 9.212
Primus HPLC (affinity) | 25 | 6 |  | 4.6 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Primus | 9.016
Primus HPLC (affinity) | 24 | 12.8 |  | 3.2 | GH2-01 | 13 | 2008 | CAP | Fresh | Primus | 6.272
Primus HPLC (affinity) | 24 | 5.3 |  | 2.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Primus | 5.488
Primus HPLC (affinity) | 24 | 8.6 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Primus | 5.88
Primus HPLC (affinity) | 32 | 12 |  | 2.2 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Primus | 4.312
Primus HPLC (affinity) | 32 | 6.8 |  | 3.2 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Primus | 6.272
Primus HPLC (affinity) | 32 | 5.1 |  | 2.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Primus | 5.292
Primus HPLC (affinity) | 12 | 5.1 |  | 2.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Primus | 5.488
Primus HPLC (affinity) | 12 | 8.3 |  | 2.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Primus | 5.488
Primus HPLC (affinity) | 12 | 5.9 |  | 2.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Primus | 4.508
Roche Diagnostics Cobas | 33 | 6.58 | 0.28 | 2.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche Diagnostics | 5.964
Roche Diagnostics Cobas | 33 | 7.95 | 0.35 | 2.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche Diagnostics | 6.034
Roche Diagnostics Cobas | 33 | 9.41 | 0.21 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche Diagnostics | 4.914
Roche Diagnostics Cobas | 12 | 5.55 | -0.05 | 5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche Diagnostics | 9.75
Roche Diagnostics Cobas | 12 | 9.53 | 0.13 | 5.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche Diagnostics | 11.302
Roche Diagnostics Cobas | 12 | 7.23 | 0.03 | 4.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche Diagnostics | 8.066
Roche Diagnostics Cobas | 15 | 5.5 | 0.1 | 5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche Diagnostics | 9.9
Roche Diagnostics Cobas | 15 | 8.33 | 0.03 | 5.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche Diagnostics | 11.398
Roche Diagnostics Cobas | 15 | 5.7 | 0.05 | 12.2 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche Diagnostics | 23.962
Roche Diagnostics Cobas | 15 | 7.09 | -0.02 | 3.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche Diagnostics | 6.448
Roche Diagnostics Cobas | 15 | 9.44 | 0.12 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche Diagnostics | 5.804
Roche Diagnostics Cobas | 15 | 6.03 | -0.04 | 4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche Diagnostics | 7.8
Roche Diagnostics Cobas | 23 | 8.07 | 0.02 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche Diagnostics | 5.9
Roche Diagnostics Cobas | 23 | 5.38 | 0.08 | 4.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche Diagnostics | 8.9
Roche Diagnostics Cobas | 23 | 6.21 | 0.07 | 4.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche Diagnostics | 8.106
Roche Diagnostics Cobas | 25 | 6.68 | 0.1 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche Diagnostics | 5.196
Roche Diagnostics Cobas | 25 | 8.66 | 0.27 | 2.4 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche Diagnostics | 4.974
Roche Diagnostics Cobas | 25 | 5.7 | 0.05 | 4.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche Diagnostics | 9.066
Roche Diagnostics Cobas | 29 | 6.6 | 0.11 | 3.1 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche Diagnostics | 6.186
Roche Diagnostics Cobas | 29 | 7.11 | 0.14 | 2.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche Diagnostics | 4.844
Roche Diagnostics Cobas | 29 | 9.94 | 0.29 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche Diagnostics | 5.778
Roche Diagnostics Cobas | 30 | 6.8 | 0.01 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche Diagnostics | 5.302
Roche Diagnostics Cobas | 30 | 10.39 | 0.11 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche Diagnostics | 5.206
Roche Diagnostics Cobas | 30 | 6.78 | -0.04 | 2.7 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche Diagnostics | 5.252
Roche Diagnostics Cobas | 19 | 8.73 | 0.1 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche Diagnostics | 5.588
Roche Diagnostics Cobas | 19 | 5.35 | 0.03 | 4.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche Diagnostics | 8.066
Roche Diagnostics Cobas | 18 | 5.83 | -0.01 | 3.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche Diagnostics | 7.438
Roche Diagnostics Cobas | 18 | 11.82 | 0.11 | 3.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche Diagnostics | 6.186
Roche Diagnostics Cobas | 18 | 9.72 | 0.19 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche Diagnostics | 5.482
Roche Diagnostics Cobas | 18 | 5.05 | 0.01 | 3.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche Diagnostics | 7.458
Roche Diagnostics Cobas | 18 | 7.38 | 0 | 3.4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas | 18-26 | 11.62 | -0.07 | 3.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche Diagnostics | 6.986
Roche Diagnostics Cobas | 18-26 | 5.81 | -0.12 | 4.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche Diagnostics | 9.288
Roche Diagnostics Cobas | 18-26 | 5.07 | -0.1 | 5.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche Diagnostics | 11.268
Roche Diagnostics Cobas | 18-26 | 8.19 | 0.05 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche Diagnostics | 4.95
Roche Diagnostics Cobas | 18-26 | 9.38 | 0.08 | 3.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche Diagnostics | 6.548
Roche Diagnostics Cobas | 16-25 | 5.15 | -0.17 | 4.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche Diagnostics | 8.846
Roche Diagnostics Cobas | 16-25 | 9.28 | 0.12 | 3.6 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche Diagnostics | 7.176
Roche Diagnostics Cobas | 16-25 | 5.17 | -0.14 | 4.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche Diagnostics | 8.092
Roche Diagnostics Cobas | 16-25 | 12.11 | 0.08 | 2.6 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche Diagnostics | 5.176
Roche Diagnostics Cobas | 16-25 | 5.85 | -0.09 | 2.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche Diagnostics | 5.202
Roche Diagnostics Cobas | 18 | 5.06 | -0.21 | 3.5 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche Diagnostics | 6.65
Roche Diagnostics Cobas | 18 | 10.42 | -0.17 | 3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche Diagnostics | 5.71
Roche Diagnostics Cobas | 18 | 5.94 | -0.26 | 2.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche Diagnostics | 4.64
Roche Diagnostics Cobas | 18 | 11.95 | -0.28 | 3.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche Diagnostics | 5.796
Roche Diagnostics Cobas | 18 | 7.38 | -0.13 | 2.7 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche Diagnostics | 5.162
Roche Diagnostics Cobas | 16-28 | 9.15 | 0.04 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche Diagnostics | 4.548
Roche Diagnostics Cobas | 16-28 | 5.83 | -0.18 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche Diagnostics | 5.308
Roche Diagnostics Cobas | 16-28 | 11.52 | -0.19 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche Diagnostics | 3.73
Roche Diagnostics Cobas | 16-28 | 4.75 | -0.27 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche Diagnostics | 5.806
Roche Diagnostics Cobas | 16-28 | 7.4 | -0.18 | 1.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche Diagnostics | 3.348
Roche Diagnostics Cobas | 20-28 | 6.16 | -0.25 | 3.8 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche Diagnostics | 7.198
Roche Diagnostics Cobas | 20-28 | 9.44 | -0.09 | 3.3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche Diagnostics | 6.378
Roche Diagnostics Cobas | 20-28 | 5.07 | -0.27 | 5.2 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche Diagnostics | 9.922
Roche Diagnostics Cobas | 20-28 | 8.39 | -0.12 | 2.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche Diagnostics | 5.172
Roche Diagnostics Cobas | 20-28 | 7.15 | -0.1 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 4.8
Roche Diagnostics Cobas | 20 | 7.21 | -0.21 | 2.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche Diagnostics | 4.102
Roche Diagnostics Cobas | 20 | 4.87 | -0.33 | 3.4 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche Diagnostics | 6.334
Roche Diagnostics Cobas | 20 | 8.19 | -0.12 | 3.3 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche Diagnostics | 6.348
Roche Diagnostics Cobas | 20 | 10.32 | -0.01 | 3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche Diagnostics | 5.87
Roche Diagnostics Cobas | 20 | 5.56 | -0.31 | 3.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche Diagnostics | 6.158
Roche Diagnostics Cobas | 26 | 7.1 | -0.15 | 1.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 3.378
Roche Diagnostics Cobas | 26 | 6.56 | -0.25 | 2.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche Diagnostics | 4.062
Roche Diagnostics Cobas | 26 | 8.64 | -0.01 | 1.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche Diagnostics | 3.126
Roche Diagnostics Cobas | 26 | 9.55 | 0.05 | 1.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche Diagnostics | 2.794
Roche Diagnostics Cobas | 26 | 5.16 | -0.29 | 2.6 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche Diagnostics | 4.806
Roche Diagnostics Cobas | 21 | 6.93 | -0.22 | 2.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche Diagnostics | 5.464
Roche Diagnostics Cobas | 21 | 4.89 | -0.3 | 4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche Diagnostics | 7.54
Roche Diagnostics Cobas | 21 | 8.35 | -0.07 | 2.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche Diagnostics | 4.046
Roche Diagnostics Cobas | 21 | 9.88 | 0.09 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche Diagnostics | 3.618
Roche Diagnostics Cobas | 21 | 5.88 | -0.24 | 3.1 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche Diagnostics | 5.836
Roche Diagnostics Cobas | 12 | 6.24 | -0.07 | 1.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche Diagnostics | 2.674
Roche Diagnostics Cobas | 12 | 9.26 | 0.15 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche Diagnostics | 3.09
Roche Diagnostics Cobas | 12 | 5.18 | -0.13 | 2.5 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche Diagnostics | 4.77
Roche Diagnostics Cobas | 12 | 8.33 | -0.02 | 1.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche Diagnostics | 2.724
Roche Diagnostics Cobas | 12 | 7.64 | 0.02 | 1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche Diagnostics | 1.98
Roche Diagnostics Cobas | 18 | 9.36 | -0.01 | 1.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche Diagnostics | 3.322
Roche Diagnostics Cobas | 18 | 4.77 | -0.15 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche Diagnostics | 5.926
Roche Diagnostics Cobas | 18 | 5.91 | -0.13 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche Diagnostics | 4.378
Roche Diagnostics Cobas | 18 | 5.18 | -0.1 | 1.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche Diagnostics | 2.056
Roche Diagnostics Cobas | 18 | 7.96 | -0.12 | 1.4 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche Diagnostics | 2.624
Roche Diagnostics Cobas | 19 | 5.18 | -0.28 | 2.7 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche Diagnostics | 5.012
Roche Diagnostics Cobas | 19 | 5.43 | -0.23 | 1.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche Diagnostics | 3.494
Roche Diagnostics Cobas | 19 | 9.33 | 0.02 | 1.2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche Diagnostics | 2.372
Roche Diagnostics Cobas | 19 | 4.99 | -0.29 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche Diagnostics | 6.178
Roche Diagnostics Cobas | 19 | 7.31 | -0.1 | 1.1 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche Diagnostics | 2.056
Roche Diagnostics Cobas | 14 | 6.36 | -0.05 | 1.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 3.478
Roche Diagnostics Cobas | 14 | 8.71 | 0.11 | 1.8 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche Diagnostics | 3.638
Roche Diagnostics Cobas | 14 | 5.29 | -0.13 | 2.3 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche Diagnostics | 4.378
Roche Diagnostics Cobas | 14 | 7.39 | 0.01 | 2.1 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche Diagnostics | 4.126
Roche Diagnostics Cobas | 14 | 9.96 | 0.21 | 2.6 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche Diagnostics | 5.306
Roche Diagnostics Cobas | 21 | 5.14 | -0.11 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche Diagnostics | 4.986
Roche Diagnostics Cobas | 21 | 6.35 | -0.06 | 2.2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 4.252
Roche Diagnostics Cobas | 21 | 8.31 | 0.1 | 2.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche Diagnostics | 4.608
Roche Diagnostics Cobas | 21 | 5.34 | -0.13 | 2.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche Diagnostics | 5.162
Roche Diagnostics Cobas | 21 | 9.75 | 0.27 | 1.4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche Diagnostics | 3.014
Roche Diagnostics Cobas | 18 | 8.14 | 0.17 | 2.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche Diagnostics | 4.286
Roche Diagnostics Cobas | 18 | 5.85 | -0.04 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche Diagnostics | 4.664
Roche Diagnostics Cobas | 18 | 5.03 | -0.1 | 3.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche Diagnostics | 6.368
Roche Diagnostics Cobas | 18 | 7.53 | 0.13 | 2.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche Diagnostics | 5.814
Roche Diagnostics Cobas | 18 | 9.31 | 0.14 | 1.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche Diagnostics | 3.276
Roche Diagnostics Cobas | 17 | 7.9 | 0.09 | 2.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche Diagnostics | 4.206
Roche Diagnostics Cobas | 17 | 5.85 | -0.06 | 2.5 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche Diagnostics | 4.84
Roche Diagnostics Cobas | 17 | 8.82 | 0.13 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche Diagnostics | 4.442
Roche Diagnostics Cobas | 17 | 5.76 | -0.1 | 3.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche Diagnostics | 6.368
Roche Diagnostics Cobas | 17 | 9.72 | 0.16 | 1.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche Diagnostics | 3.884
Roche Diagnostics Cobas | 21 | 5.44 | -0.04 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche Diagnostics | 5.448
Roche Diagnostics Cobas | 21 | 8.3 | 0.04 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche Diagnostics | 4.352
Roche Diagnostics Cobas | 21 | 5.06 | -0.09 | 3.4 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche Diagnostics | 6.574
Roche Diagnostics Cobas | 21 | 10.12 | 0.07 | 2.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche Diagnostics | 4.774
Roche Diagnostics Cobas | 21 | 5.92 | -0.07 | 2.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche Diagnostics | 4.438
Roche Diagnostics Cobas | 23 | 8.9 | 0.04 | 2.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche Diagnostics | 5.332
Roche Diagnostics Cobas | 23 | 6.75 | -0.09 | 2.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche Diagnostics | 4.81
Roche Diagnostics Cobas | 23 | 9.43 | 0.04 | 3.3 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche Diagnostics | 6.508
Roche Diagnostics Cobas | 23 | 6.06 | -0.07 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche Diagnostics | 4.83
Roche Diagnostics Cobas | 23 | 7.48 | -0.04 | 2.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche Diagnostics | 5.448
Roche Diagnostics Cobas | 12 | 5.63 | -0.17 | 1.9 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche Diagnostics | 3.554
Roche Diagnostics Cobas | 12 | 9.24 | 0.12 | 1.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche Diagnostics | 3.844
Roche Diagnostics Cobas | 12 | 5.32 | -0.15 | 2.7 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche Diagnostics | 5.142
Roche Diagnostics Cobas | 12 | 7.28 | 0.03 | 1.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche Diagnostics | 3.166
Roche Diagnostics Cobas | 12 | 9.35 | 0.09 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche Diagnostics | 2.834
Roche Diagnostics Cobas | 21 | 4.95 | -0.26 | 3.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche Diagnostics | 6.404
Roche Diagnostics Cobas | 21 | 6.53 | -0.15 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche Diagnostics | 4.946
Roche Diagnostics Cobas | 21 | 8.25 |  | 2.6 | GH5-13 |  | 2021 | CAP | Fresh | Roche Diagnostics | 5.096
Roche Diagnostics Cobas | 21 | 5.83 | -0.15 | 2.7 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche Diagnostics | 5.142
Roche Diagnostics Cobas | 21 | 10.21 | 0.09 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche Diagnostics | 4.402
Roche Diagnostics Cobas c500 series | 237 | 6.95 | -0.16 | 2.4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche Diagnostics | 4.544
Roche Diagnostics Cobas c500 series | 237 | 9.14 | -0.18 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche Diagnostics | 5.504
Roche Diagnostics Cobas c500 series | 237 | 6.06 | -0.01 | 3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche Diagnostics | 5.87
Roche Diagnostics Cobas c500 series | 274 | 7.82 | -0.23 | 2.8 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche Diagnostics | 5.258
Roche Diagnostics Cobas c500 series | 274 | 5.36 | 0.06 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche Diagnostics | 6.136
Roche Diagnostics Cobas c500 series | 274 | 6.12 | -0.02 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche Diagnostics | 6.056
Roche Diagnostics Cobas c500 series | 350 | 6.63 | 0.05 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche Diagnostics | 5.538
Roche Diagnostics Cobas c500 series | 350 | 8.39 | 0 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics Cobas c500 series | 350 | 5.73 | 0.08 | 2.9 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche Diagnostics | 5.764
Roche Diagnostics Cobas c500 series | 316 | 6.52 | 0.03 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche Diagnostics | 4.93
Roche Diagnostics Cobas c500 series | 316 | 6.95 | -0.02 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche Diagnostics | 5.076
Roche Diagnostics Cobas c500 series | 316 | 9.51 | -0.14 | 2.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche Diagnostics | 5.152
Roche Diagnostics Cobas c500 series | 358 | 6.77 | -0.02 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche Diagnostics | 5.272
Roche Diagnostics Cobas c500 series | 358 | 10.01 | -0.27 | 2.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche Diagnostics | 5.218
Roche Diagnostics Cobas c500 series | 358 | 6.76 | -0.06 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche Diagnostics | 4.84
Roche Diagnostics Cobas c500 series | 293 | 8.51 | -0.12 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche Diagnostics | 5.564
Roche Diagnostics Cobas c500 series | 293 | 5.39 | 0.07 | 2.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche Diagnostics | 5.558
Roche Diagnostics Cobas c500 series | 294 | 5.84 | 0 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche Diagnostics | 5.292
Roche Diagnostics Cobas c500 series | 294 | 11.22 | -0.49 | 3.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche Diagnostics | 5.782
Roche Diagnostics Cobas c500 series | 294 | 9.32 | -0.21 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche Diagnostics | 5.866
Roche Diagnostics Cobas c500 series | 294 | 5.11 | 0.07 | 3.5 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche Diagnostics | 6.93
Roche Diagnostics Cobas c500 series | 294 | 7.27 | -0.11 | 2.4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche Diagnostics | 4.594
Roche Diagnostics Cobas c500 series | 307-385 | 11.37 | -0.32 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche Diagnostics | 5.364
Roche Diagnostics Cobas c500 series | 307-385 | 5.89 | -0.04 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche Diagnostics | 5.252
Roche Diagnostics Cobas c500 series | 307-385 | 5.19 | 0.02 | 3.5 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche Diagnostics | 6.88
Roche Diagnostics Cobas c500 series | 307-385 | 8.04 | -0.1 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche Diagnostics | 4.8
Roche Diagnostics Cobas c500 series | 307-385 | 9.17 | -0.13 | 2.9 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche Diagnostics | 5.554
Roche Diagnostics Cobas c500 series | 319-381 | 5.23 | -0.09 | 3.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche Diagnostics | 6.182
Roche Diagnostics Cobas c500 series | 319-381 | 9.11 | -0.05 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche Diagnostics | 5.438
Roche Diagnostics Cobas c500 series | 319-381 | 5.23 | -0.08 | 3.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche Diagnostics | 6.388
Roche Diagnostics Cobas c500 series | 319-381 | 11.82 | -0.21 | 3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche Diagnostics | 5.67
Roche Diagnostics Cobas c500 series | 319-381 | 5.9 | -0.04 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche Diagnostics | 4.468
Roche Diagnostics Cobas c500 series | 321 | 5.19 | -0.08 | 2.8 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche Diagnostics | 5.408
Roche Diagnostics Cobas c500 series | 321 | 10.32 | -0.27 | 2.7 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche Diagnostics | 5.022
Roche Diagnostics Cobas c500 series | 321 | 6.07 | -0.13 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche Diagnostics | 4.182
Roche Diagnostics Cobas c500 series | 321 | 11.86 | -0.37 | 2.7 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche Diagnostics | 4.922
Roche Diagnostics Cobas c500 series | 321 | 7.49 | -0.02 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche Diagnostics | 4.488
Roche Diagnostics Cobas c500 series | 331-399 | 9.03 | -0.08 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche Diagnostics | 4.232
Roche Diagnostics Cobas c500 series | 331-399 | 5.9 | -0.11 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche Diagnostics | 4.398
Roche Diagnostics Cobas c500 series | 331-399 | 11.43 | -0.28 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche Diagnostics | 4.816
Roche Diagnostics Cobas c500 series | 331-399 | 4.87 | -0.15 | 2.9 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche Diagnostics | 5.534
Roche Diagnostics Cobas c500 series | 331-399 | 7.52 | -0.06 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche Diagnostics | 4.252
Roche Diagnostics Cobas c500 series | 340-396 | 6.28 | -0.13 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche Diagnostics | 4.966
Roche Diagnostics Cobas c500 series | 340-396 | 9.45 | -0.08 | 2.5 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche Diagnostics | 4.82
Roche Diagnostics Cobas c500 series | 340-396 | 5.21 | -0.13 | 2.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche Diagnostics | 5.554
Roche Diagnostics Cobas c500 series | 340-396 | 8.4 | -0.11 | 2.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche Diagnostics | 4.594
Roche Diagnostics Cobas c500 series | 340-396 | 7.2 | -0.05 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 4.458
Roche Diagnostics Cobas c500 series | 346 | 7.31 | -0.11 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche Diagnostics | 4.79
Roche Diagnostics Cobas c500 series | 346 | 5 | -0.2 | 2.9 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche Diagnostics | 5.484
Roche Diagnostics Cobas c500 series | 346 | 8.23 | -0.08 | 2.6 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche Diagnostics | 5.016
Roche Diagnostics Cobas c500 series | 346 | 10.28 | -0.05 | 2.7 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche Diagnostics | 5.242
Roche Diagnostics Cobas c500 series | 346 | 5.7 | -0.17 | 2.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche Diagnostics | 5.514
Roche Diagnostics Cobas c500 series | 406 | 7.16 | -0.09 | 2.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 4.026
Roche Diagnostics Cobas c500 series | 406 | 6.67 | -0.14 | 2.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche Diagnostics | 4.172
Roche Diagnostics Cobas c500 series | 406 | 8.58 | -0.07 | 2.2 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche Diagnostics | 4.242
Roche Diagnostics Cobas c500 series | 406 | 9.48 | -0.02 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche Diagnostics | 4.488
Roche Diagnostics Cobas c500 series | 406 | 5.31 | -0.14 | 2.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche Diagnostics | 4.172
Roche Diagnostics Cobas c500 series | 427 | 7.05 | -0.1 | 2.5 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche Diagnostics | 4.8
Roche Diagnostics Cobas c500 series | 427 | 5.06 | -0.13 | 3.3 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche Diagnostics | 6.338
Roche Diagnostics Cobas c500 series | 427 | 8.36 | -0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche Diagnostics | 4.644
Roche Diagnostics Cobas c500 series | 427 | 9.77 | -0.02 | 2.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche Diagnostics | 5.076
Roche Diagnostics Cobas c500 series | 427 | 6.03 | -0.09 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche Diagnostics | 4.614
Roche Diagnostics Cobas c500 series | 367 | 6.24 | -0.07 | 2.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche Diagnostics | 4.634
Roche Diagnostics Cobas c500 series | 367 | 9.13 | 0.02 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche Diagnostics | 5.116
Roche Diagnostics Cobas c500 series | 367 | 5.25 | -0.06 | 2.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche Diagnostics | 4.448
Roche Diagnostics Cobas c500 series | 367 | 8.34 | -0.01 | 2.5 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche Diagnostics | 4.89
Roche Diagnostics Cobas c500 series | 367 | 7.63 | 0.01 | 2.2 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche Diagnostics | 4.322
Roche Diagnostics Cobas c500 series | 413 | 9.28 | -0.09 | 2.4 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche Diagnostics | 4.614
Roche Diagnostics Cobas c500 series | 413 | 4.86 | -0.06 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche Diagnostics | 5.232
Roche Diagnostics Cobas c500 series | 413 | 6 | -0.04 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche Diagnostics | 4.468
Roche Diagnostics Cobas c500 series | 413 | 5.23 | -0.05 | 2.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche Diagnostics | 5.046
Roche Diagnostics Cobas c500 series | 413 | 8.02 | -0.06 | 2.5 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche Diagnostics | 4.84
Roche Diagnostics Cobas c500 series | 416 | 5.26 | -0.2 | 3.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche Diagnostics | 6.072
Roche Diagnostics Cobas c500 series | 416 | 5.52 | -0.14 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche Diagnostics | 5.348
Roche Diagnostics Cobas c500 series | 416 | 9.2 | -0.11 | 2.4 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche Diagnostics | 4.594
Roche Diagnostics Cobas c500 series | 416 | 5.08 | -0.2 | 2.8 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche Diagnostics | 5.288
Roche Diagnostics Cobas c500 series | 416 | 7.34 | -0.07 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche Diagnostics | 4.242
Roche Diagnostics Cobas c500 series | 364 | 6.38 | -0.03 | 2.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 4.478
Roche Diagnostics Cobas c500 series | 364 | 8.63 | 0.03 | 2.1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche Diagnostics | 4.146
Roche Diagnostics Cobas c500 series | 364 | 5.36 | -0.06 | 2.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche Diagnostics | 4.644
Roche Diagnostics Cobas c500 series | 364 | 7.41 | 0.03 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche Diagnostics | 4.734
Roche Diagnostics Cobas c500 series | 364 | 9.78 | 0.03 | 2.3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche Diagnostics | 4.538
Roche Diagnostics Cobas c500 series | 406 | 5.18 | -0.07 | 2.3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche Diagnostics | 4.438
Roche Diagnostics Cobas c500 series | 406 | 6.42 | 0.01 | 2.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 4.714
Roche Diagnostics Cobas c500 series | 406 | 8.32 | 0.11 | 2.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche Diagnostics | 4.618
Roche Diagnostics Cobas c500 series | 406 | 5.43 | -0.04 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche Diagnostics | 4.272
Roche Diagnostics Cobas c500 series | 406 | 9.57 | 0.09 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche Diagnostics | 4.598
Roche Diagnostics Cobas c500 series | 357 | 8.11 | 0.14 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche Diagnostics | 4.06
Roche Diagnostics Cobas c500 series | 357 | 5.91 | 0.02 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche Diagnostics | 4.528
Roche Diagnostics Cobas c500 series | 357 | 5.06 | -0.07 | 2.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche Diagnostics | 4.83
Roche Diagnostics Cobas c500 series | 357 | 7.49 | 0.09 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche Diagnostics | 4.402
Roche Diagnostics Cobas c500 series | 357 | 9.22 | 0.05 | 2.5 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche Diagnostics | 4.95
Roche Diagnostics Cobas c500 series | 347 | 7.85 | 0.04 | 2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche Diagnostics | 3.96
Roche Diagnostics Cobas c500 series | 347 | 5.9 | -0.01 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche Diagnostics | 4.302
Roche Diagnostics Cobas c500 series | 347 | 8.74 | 0.05 | 2.3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche Diagnostics | 4.558
Roche Diagnostics Cobas c500 series | 347 | 5.85 | -0.01 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche Diagnostics | 4.89
Roche Diagnostics Cobas c500 series | 347 | 9.58 | 0.02 | 2.5 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche Diagnostics | 4.92
Roche Diagnostics Cobas c500 series | 400 | 5.47 | -0.01 | 2.4 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche Diagnostics | 4.694
Roche Diagnostics Cobas c500 series | 400 | 8.34 | 0.08 | 2.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche Diagnostics | 4.588
Roche Diagnostics Cobas c500 series | 400 | 5.13 | -0.02 | 2.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche Diagnostics | 5.272
Roche Diagnostics Cobas c500 series | 400 | 10.09 | 0.04 | 2.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche Diagnostics | 4.548
Roche Diagnostics Cobas c500 series | 400 | 5.99 | 0 | 2.4 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche Diagnostics | 4.704
Roche Diagnostics Cobas c500 series | 394 | 8.89 | 0.03 | 2.5 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche Diagnostics | 4.93
Roche Diagnostics Cobas c500 series | 394 | 6.78 | -0.06 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche Diagnostics | 4.448
Roche Diagnostics Cobas c500 series | 394 | 9.41 | 0.02 | 2.6 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche Diagnostics | 5.116
Roche Diagnostics Cobas c500 series | 394 | 6.09 | -0.04 | 2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche Diagnostics | 3.88
Roche Diagnostics Cobas c500 series | 394 | 7.48 | -0.04 | 2.3 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche Diagnostics | 4.468
Roche Diagnostics Cobas c500 series | 350 | 5.67 | -0.13 | 2.4 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche Diagnostics | 4.574
Roche Diagnostics Cobas c500 series | 350 | 9.1 | -0.02 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche Diagnostics | 4.488
Roche Diagnostics Cobas c500 series | 350 | 5.39 | -0.08 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche Diagnostics | 4.232
Roche Diagnostics Cobas c500 series | 350 | 7.23 | -0.02 | 2.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche Diagnostics | 4.096
Roche Diagnostics Cobas c500 series | 350 | 9.26 | 0 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche Diagnostics | 4.704
Roche Diagnostics Cobas c500 series | 409 | 5.1 | -0.11 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche Diagnostics | 5.574
Roche Diagnostics Cobas c500 series | 409 | 6.63 | 0.05 | 2.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche Diagnostics | 4.362
Roche Diagnostics Cobas c500 series | 409 | 8.25 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Roche Diagnostics | 4.704
Roche Diagnostics Cobas c500 series | 409 | 5.93 | -0.05 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche Diagnostics | 4.654
Roche Diagnostics Cobas c500 series | 409 | 10.1 | -0.02 | 2.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche Diagnostics | 5.076
Roche Diagnostics Cobas c500/700 | 171 | 8.47 | -0.03 | 3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche Diagnostics | 5.85
Roche Diagnostics Cobas c500/700 | 171 | 5.58 | 0.18 | 3.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche Diagnostics | 6.648
Roche Diagnostics Cobas c500/700 | 171 | 6.52 | 0.12 | 3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche Diagnostics | 6
Roche Diagnostics Cobas c500/700 | 12 | 6.29 | -0.01 | 3.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche Diagnostics | 6.654
Roche Diagnostics Cobas c500/700 | 12 | 7.68 | 0.08 | 2.2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche Diagnostics | 4.392
Roche Diagnostics Cobas c500/700 | 12 | 9.35 | 0.15 | 4.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche Diagnostics | 8.186
Roche Diagnostics Cobas c500/700 | 248 | 5.62 | 0.02 | 2.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche Diagnostics | 4.92
Roche Diagnostics Cobas c500/700 | 248 | 9.26 | -0.14 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche Diagnostics | 4.956
Roche Diagnostics Cobas c500/700 | 248 | 7.15 | -0.05 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche Diagnostics | 4.654
Roche Diagnostics Cobas c500/700 | 262 | 5.5 | 0.1 | 2.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche Diagnostics | 5
Roche Diagnostics Cobas c500/700 | 262 | 8.12 | -0.18 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche Diagnostics | 4.72
Roche Diagnostics Cobas c500/700 | 262 | 5.61 | -0.04 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche Diagnostics | 5.056
Roche Diagnostics Cobas c501 | 6 | 5.8 |  |  | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche Diagnostics | 0
Roche Diagnostics Cobas c501 | 6 | 11.4 |  |  | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche Diagnostics | 0
Roche Diagnostics Cobas c501 | 6 | 7.5 |  |  | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche Diagnostics | 0
Roche Diagnostics Cobas c501 | 20 | 9 |  | 3.1 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche Diagnostics | 6.076
Roche Diagnostics Cobas c501 | 20 | 11 |  | 3.9 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas c501 | 20 | 6.3 |  | 3.4 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas c501 | 45 | 12.3 |  | 3.7 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas c501 | 45 | 5.4 |  | 3.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche Diagnostics | 6.272
Roche Diagnostics Cobas c501 | 45 | 8.6 |  | 3.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas c501 | 73 | 11.4 |  | 3.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas c501 | 73 | 6.7 |  | 3.2 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche Diagnostics | 6.272
Roche Diagnostics Cobas c501 | 73 | 5.3 |  | 2.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche Diagnostics | 5.096
Roche Diagnostics Cobas c501 | 108 | 5.3 |  | 3.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas c501 | 108 | 8.2 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas c501 | 108 | 6 |  | 3.4 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas c501 | 133 | 6.7 | 0.1 | 3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche Diagnostics | 5.98
Roche Diagnostics Cobas c501 | 133 | 9.29 | -0.21 | 3.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche Diagnostics | 7.042
Roche Diagnostics Cobas c501 | 133 | 7.35 | -0.05 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche Diagnostics | 7.202
Roche Diagnostics Cobas c501 | 133 | 6.07 | 0.17 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche Diagnostics | 5.658
Roche Diagnostics Cobas c501 | 133 | 9.51 | -0.29 | 3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche Diagnostics | 5.59
Roche Diagnostics Cobas c501 | 133 | 7.4 | 0 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics Cobas c501 | 161 | 5.4 | 0.2 | 3.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche Diagnostics | 6.668
Roche Diagnostics Cobas c501 | 161 | 8.55 | -0.15 | 3.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche Diagnostics | 5.926
Roche Diagnostics Cobas c501 | 161 | 6.37 | 0.07 | 2.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche Diagnostics | 5.754
Roche Diagnostics Cobas c513 | 13 | 7.43 | 0.01 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche Diagnostics | 3.342
Roche Diagnostics Cobas c513 | 13 | 5.12 | -0.08 | 2.3 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche Diagnostics | 4.428
Roche Diagnostics Cobas c513 | 13 | 8.36 | 0.05 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche Diagnostics | 3.382
Roche Diagnostics Cobas c513 | 13 | 10.6 | 0.27 | 1.8 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche Diagnostics | 3.798
Roche Diagnostics Cobas c513 | 13 | 5.78 | -0.09 | 2.2 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche Diagnostics | 4.222
Roche Diagnostics Cobas c513 | 34 | 7.24 | -0.01 | 1.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 2.93
Roche Diagnostics Cobas c513 | 34 | 6.78 | -0.03 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche Diagnostics | 2.714
Roche Diagnostics Cobas c513 | 34 | 8.69 | 0.04 | 1.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche Diagnostics | 3.176
Roche Diagnostics Cobas c513 | 34 | 9.67 | 0.17 | 1.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche Diagnostics | 3.11
Roche Diagnostics Cobas c513 | 34 | 5.42 | -0.03 | 1.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche Diagnostics | 2.518
Roche Diagnostics Cobas c513 | 44 | 7.15 | 0 | 2.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche Diagnostics | 4.508
Roche Diagnostics Cobas c513 | 44 | 5.19 | 0 | 2.2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche Diagnostics | 4.312
Roche Diagnostics Cobas c513 | 44 | 8.46 | 0.04 | 1.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche Diagnostics | 3.568
Roche Diagnostics Cobas c513 | 44 | 10.03 | 0.24 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche Diagnostics | 3.768
Roche Diagnostics Cobas c513 | 44 | 6.1 | -0.02 | 1.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche Diagnostics | 3.116
Roche Diagnostics Cobas c513 | 47 | 6.29 | -0.02 | 1.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche Diagnostics | 2.92
Roche Diagnostics Cobas c513 | 47 | 9.32 | 0.21 | 1.7 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche Diagnostics | 3.542
Roche Diagnostics Cobas c513 | 47 | 5.28 | -0.03 | 1.4 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche Diagnostics | 2.714
Roche Diagnostics Cobas c513 | 47 | 8.45 | 0.1 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche Diagnostics | 3.628
Roche Diagnostics Cobas c513 | 47 | 7.7 | 0.08 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche Diagnostics | 3.02
Roche Diagnostics Cobas c513 | 55 | 9.39 | 0.02 | 2 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche Diagnostics | 3.94
Roche Diagnostics Cobas c513 | 55 | 4.93 | 0.01 | 2.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche Diagnostics | 4.126
Roche Diagnostics Cobas c513 | 55 | 6.04 | 0 | 1.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche Diagnostics | 3.136
Roche Diagnostics Cobas c513 | 55 | 5.28 | 0 | 2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche Diagnostics | 3.92
Roche Diagnostics Cobas c513 | 55 | 8.04 | -0.04 | 1.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche Diagnostics | 3.292
Roche Diagnostics Cobas c513 | 63 | 5.42 | -0.04 | 1.9 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche Diagnostics | 3.684
Roche Diagnostics Cobas c513 | 63 | 5.64 | -0.02 | 1.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche Diagnostics | 3.508
Roche Diagnostics Cobas c513 | 63 | 9.53 | 0.22 | 1.8 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche Diagnostics | 3.748
Roche Diagnostics Cobas c513 | 63 | 5.23 | -0.05 | 1.8 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche Diagnostics | 3.478
Roche Diagnostics Cobas c513 | 63 | 7.45 | 0.04 | 1.6 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche Diagnostics | 3.176
Roche Diagnostics Cobas c513 | 54 | 6.43 | 0.02 | 1.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 2.568
Roche Diagnostics Cobas c513 | 54 | 8.66 | 0.06 | 1.6 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche Diagnostics | 3.196
Roche Diagnostics Cobas c513 | 54 | 5.42 | 0 | 1.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche Diagnostics | 2.744
Roche Diagnostics Cobas c513 | 54 | 7.4 | 0.02 | 1.7 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche Diagnostics | 3.352
Roche Diagnostics Cobas c513 | 54 | 9.9 | 0.15 | 1.8 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche Diagnostics | 3.678
Roche Diagnostics Cobas c513 | 59 | 5.21 | -0.04 | 1.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche Diagnostics | 3.096
Roche Diagnostics Cobas c513 | 59 | 6.41 | 0 | 1.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 3.332
Roche Diagnostics Cobas c513 | 59 | 8.24 | 0.03 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche Diagnostics | 3.754
Roche Diagnostics Cobas c513 | 59 | 5.43 | -0.04 | 1.5 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche Diagnostics | 2.9
Roche Diagnostics Cobas c513 | 59 | 9.57 | 0.09 | 1.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche Diagnostics | 3.814
Roche Diagnostics Cobas c513 | 56 | 8.02 | 0.05 | 1.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche Diagnostics | 3.382
Roche Diagnostics Cobas c513 | 56 | 5.88 | -0.01 | 1.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche Diagnostics | 2.538
Roche Diagnostics Cobas c513 | 56 | 5.1 | -0.03 | 1.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche Diagnostics | 2.91
Roche Diagnostics Cobas c513 | 56 | 7.39 | -0.01 | 1.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche Diagnostics | 2.93
Roche Diagnostics Cobas c513 | 56 | 9.21 | 0.04 | 1.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche Diagnostics | 3.176
Roche Diagnostics Cobas c513 | 58 | 7.79 | -0.02 | 1.9 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche Diagnostics | 3.704
Roche Diagnostics Cobas c513 | 58 | 5.86 | -0.05 | 1.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche Diagnostics | 3.282
Roche Diagnostics Cobas c513 | 58 | 8.7 | 0.01 | 1.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche Diagnostics | 3.538
Roche Diagnostics Cobas c513 | 58 | 5.82 | -0.04 | 1.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche Diagnostics | 3.096
Roche Diagnostics Cobas c513 | 58 | 9.58 | 0.02 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche Diagnostics | 3.94
Roche Diagnostics Cobas c513 | 60 | 5.45 | -0.03 | 1.9 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche Diagnostics | 3.694
Roche Diagnostics Cobas c513 | 60 | 8.29 | 0.03 | 1.8 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche Diagnostics | 3.558
Roche Diagnostics Cobas c513 | 60 | 5.13 | -0.02 | 1.8 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche Diagnostics | 3.508
Roche Diagnostics Cobas c513 | 60 | 10.18 | 0.13 | 2.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche Diagnostics | 4.442
Roche Diagnostics Cobas c513 | 60 | 5.95 | -0.04 | 1.9 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche Diagnostics | 3.684
Roche Diagnostics Cobas c513 | 63 | 8.95 | 0.09 | 1.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche Diagnostics | 3.618
Roche Diagnostics Cobas c513 | 63 | 6.81 | -0.03 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche Diagnostics | 2.91
Roche Diagnostics Cobas c513 | 63 | 9.54 | 0.15 | 1.7 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche Diagnostics | 3.482
Roche Diagnostics Cobas c513 | 63 | 6.13 | 0 | 1.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche Diagnostics | 2.744
Roche Diagnostics Cobas c513 | 63 | 7.51 | -0.01 | 1.5 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche Diagnostics | 2.93
Roche Diagnostics Cobas c513 | 61 | 5.82 | 0.02 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche Diagnostics | 3.352
Roche Diagnostics Cobas c513 | 61 | 9.26 | 0.14 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche Diagnostics | 3.276
Roche Diagnostics Cobas c513 | 61 | 5.5 | 0.03 | 1.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche Diagnostics | 2.578
Roche Diagnostics Cobas c513 | 61 | 7.32 | 0.07 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche Diagnostics | 2.618
Roche Diagnostics Cobas c513 | 61 | 9.38 | 0.12 | 1.6 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche Diagnostics | 3.256
Roche Diagnostics Cobas c513 | 64 | 5.27 | 0.06 | 1.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche Diagnostics | 3
Roche Diagnostics Cobas c513 | 64 | 6.73 | 0.05 | 1.5 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche Diagnostics | 2.99
Roche Diagnostics Cobas c513 | 64 | 8.37 |  | 2.1 | GH5-13 |  | 2021 | CAP | Fresh | Roche Diagnostics | 4.116
Roche Diagnostics Cobas c513 | 64 | 6.04 | 0.06 | 1.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche Diagnostics | 2.804
Roche Diagnostics Cobas c513 | 64 | 10.41 | 0.29 | 1.5 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche Diagnostics | 3.23
Roche Diagnostics Cobas Integra | 248 | 11.1 |  | 4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Roche Diagnostics | 7.84
Roche Diagnostics Cobas Integra | 248 | 5.5 |  | 3.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas Integra | 248 | 8.7 |  | 3.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 253 | 12.1 |  | 3.9 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 253 | 12.1 |  | 3.9 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas Integra | 239 | 5.5 |  | 3.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra | 239 | 11.8 |  | 3.7 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche Diagnostics | 7.252
Roche Diagnostics Cobas Integra | 239 | 7.7 |  | 3.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas Integra | 177 | 9.2 |  | 4.3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche Diagnostics | 8.428
Roche Diagnostics Cobas Integra | 177 | 11.1 |  | 4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche Diagnostics | 7.84
Roche Diagnostics Cobas Integra | 177 | 6.3 |  | 3.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche Diagnostics | 6.272
Roche Diagnostics Cobas Integra | 146 | 12.7 |  | 3.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas Integra | 146 | 5.4 |  | 3 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche Diagnostics | 5.88
Roche Diagnostics Cobas Integra | 146 | 8.9 |  | 3.1 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche Diagnostics | 6.076
Roche Diagnostics Cobas Integra | 121 | 11.7 |  | 3.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra | 121 | 6.8 |  | 3.3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra | 121 | 5.2 |  | 3 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche Diagnostics | 5.88
Roche Diagnostics Cobas Integra 400 | 38 | 6.87 | 0.27 | 3.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche Diagnostics | 7.326
Roche Diagnostics Cobas Integra 400 | 38 | 9.75 | 0.25 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche Diagnostics | 4.954
Roche Diagnostics Cobas Integra 400 | 38 | 7.71 | 0.31 | 3.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche Diagnostics | 6.778
Roche Diagnostics Cobas Integra 400 | 42 | 6.1 | 0.2 | 2.5 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche Diagnostics | 5.1
Roche Diagnostics Cobas Integra 400 | 42 | 9.89 | 0.09 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche Diagnostics | 4.99
Roche Diagnostics Cobas Integra 400 | 42 | 7.68 | 0.28 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche Diagnostics | 5.768
Roche Diagnostics Cobas Integra 400 | 47 | 5.36 | 0.16 | 2.9 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche Diagnostics | 5.844
Roche Diagnostics Cobas Integra 400 | 47 | 8.84 | 0.14 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche Diagnostics | 5.04
Roche Diagnostics Cobas Integra 400 | 47 | 6.5 | 0.2 | 2.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche Diagnostics | 5.884
Roche Diagnostics Cobas Integra 400 | 52 | 8.68 | 0.18 | 3.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche Diagnostics | 7.628
Roche Diagnostics Cobas Integra 400 | 52 | 5.58 | 0.18 | 3.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche Diagnostics | 7.04
Roche Diagnostics Cobas Integra 400 | 52 | 6.63 | 0.23 | 3.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche Diagnostics | 7.874
Roche Diagnostics Cobas Integra 400 | 210 | 6.23 | -0.07 | 2.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche Diagnostics | 5.614
Roche Diagnostics Cobas Integra 400 | 210 | 7.5 | -0.1 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche Diagnostics | 4.8
Roche Diagnostics Cobas Integra 400 | 210 | 9.07 | -0.13 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche Diagnostics | 5.358
Roche Diagnostics Cobas Integra 400 | 53 | 5.6 | 0 | 2.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics Cobas Integra 400 | 53 | 9.5 | 0.1 | 2.8 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche Diagnostics | 5.588
Roche Diagnostics Cobas Integra 400 | 53 | 7.33 | 0.13 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche Diagnostics | 4.834
Roche Diagnostics Cobas Integra 400 | 51 | 5.47 | 0.07 | 3.2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche Diagnostics | 6.342
Roche Diagnostics Cobas Integra 400 | 51 | 8.4 | 0.1 | 3.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche Diagnostics | 7.548
Roche Diagnostics Cobas Integra 400 | 51 | 5.65 | 0 | 3.3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra 400 | 43 | 7.18 | 0.07 | 3.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche Diagnostics | 7.322
Roche Diagnostics Cobas Integra 400 | 43 | 9.54 | 0.22 | 3.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche Diagnostics | 7.864
Roche Diagnostics Cobas Integra 400 | 43 | 6.15 | 0.08 | 3.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche Diagnostics | 7.724
Roche Diagnostics Cobas Integra 400 | 41 | 7.96 | -0.09 | 3.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche Diagnostics | 6.574
Roche Diagnostics Cobas Integra 400 | 41 | 5.29 | -0.01 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche Diagnostics | 6.066
Roche Diagnostics Cobas Integra 400 | 41 | 6.08 | -0.06 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche Diagnostics | 5.82
Roche Diagnostics Cobas Integra 400 | 60 | 6.63 | 0.05 | 4.3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche Diagnostics | 8.478
Roche Diagnostics Cobas Integra 400 | 60 | 8.55 | 0.16 | 4.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche Diagnostics | 8.392
Roche Diagnostics Cobas Integra 400 | 60 | 5.65 | 0 | 4.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche Diagnostics | 8.036
Roche Diagnostics Cobas Integra 400 | 50 | 6.46 | -0.03 | 2.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche Diagnostics | 5.654
Roche Diagnostics Cobas Integra 400 | 50 | 6.95 | -0.02 | 2.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche Diagnostics | 5.664
Roche Diagnostics Cobas Integra 400 | 50 | 9.73 | 0.08 | 3 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche Diagnostics | 5.96
Roche Diagnostics Cobas Integra 400 | 58 | 6.78 | -0.01 | 3.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche Diagnostics | 6.262
Roche Diagnostics Cobas Integra 400 | 58 | 10.09 | -0.19 | 3.7 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche Diagnostics | 7.062
Roche Diagnostics Cobas Integra 400 | 58 | 6.83 | 0.01 | 3.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche Diagnostics | 6.478
Roche Diagnostics Cobas Integra 400 | 24 | 8.71 | 0.08 | 3.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche Diagnostics | 6.548
Roche Diagnostics Cobas Integra 400 | 24 | 5.36 | 0.04 | 3.4 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche Diagnostics | 6.704
Roche Diagnostics Cobas Integra 400 | 26 | 5.77 | -0.07 | 3.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche Diagnostics | 6.202
Roche Diagnostics Cobas Integra 400 | 26 | 11.26 | -0.45 | 3.4 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche Diagnostics | 6.214
Roche Diagnostics Cobas Integra 400 | 26 | 9.52 | -0.01 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche Diagnostics | 6.066
Roche Diagnostics Cobas Integra 400 | 26 | 5.06 | 0.02 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche Diagnostics | 5.508
Roche Diagnostics Cobas Integra 400 | 26 | 7.4 | 0.02 | 3.9 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche Diagnostics | 7.664
Roche Diagnostics Cobas Integra 400 | 25-64 | 11.21 | -0.48 | 4.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche Diagnostics | 8.34
Roche Diagnostics Cobas Integra 400 | 25-64 | 5.91 | -0.02 | 3.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche Diagnostics | 6.448
Roche Diagnostics Cobas Integra 400 | 25-64 | 5.22 | 0.05 | 3.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche Diagnostics | 7.302
Roche Diagnostics Cobas Integra 400 | 25-64 | 8.28 | 0.14 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche Diagnostics | 7
Roche Diagnostics Cobas Integra 400 | 25-64 | 9.36 | 0.06 | 3.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche Diagnostics | 6.528
Roche Diagnostics Cobas Integra 400 | 31-57 | 5.23 | -0.09 | 3.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche Diagnostics | 7.554
Roche Diagnostics Cobas Integra 400 | 31-57 | 9.16 | 0 | 3.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra 400 | 31-57 | 5.25 | -0.06 | 4 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche Diagnostics | 7.78
Roche Diagnostics Cobas Integra 400 | 31-57 | 11.85 | -0.18 | 3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche Diagnostics | 5.7
Roche Diagnostics Cobas Integra 400 | 31-57 | 5.88 | -0.06 | 2.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche Diagnostics | 4.84
Roche Diagnostics Cobas Integra 400 | 34 | 5.23 | -0.04 | 3.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche Diagnostics | 6.428
Roche Diagnostics Cobas Integra 400 | 34 | 10.5 | -0.09 | 2.5 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche Diagnostics | 4.81
Roche Diagnostics Cobas Integra 400 | 34 | 6.06 | -0.14 | 3.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche Diagnostics | 6.916
Roche Diagnostics Cobas Integra 400 | 34 | 12.06 | -0.17 | 3 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche Diagnostics | 5.71
Roche Diagnostics Cobas Integra 400 | 34 | 7.45 | -0.06 | 2.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche Diagnostics | 5.036
Roche Diagnostics Cobas Integra 400 | 30-60 | 9.16 | 0.05 | 3.1 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche Diagnostics | 6.126
Roche Diagnostics Cobas Integra 400 | 30-60 | 5.99 | -0.02 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche Diagnostics | 5.468
Roche Diagnostics Cobas Integra 400 | 30-60 | 11.64 | -0.07 | 3.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche Diagnostics | 6.79
Roche Diagnostics Cobas Integra 400 | 30-60 | 5.16 | 0.14 | 3.6 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche Diagnostics | 7.196
Roche Diagnostics Cobas Integra 400 | 30-60 | 7.61 | 0.03 | 2.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche Diagnostics | 5.518
Roche Diagnostics Cobas Integra 400 | 28-54 | 6.42 | 0.01 | 3.2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche Diagnostics | 6.282
Roche Diagnostics Cobas Integra 400 | 28-54 | 9.62 | 0.09 | 2.9 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche Diagnostics | 5.774
Roche Diagnostics Cobas Integra 400 | 28-54 | 5.43 | 0.09 | 3.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche Diagnostics | 6.754
Roche Diagnostics Cobas Integra 400 | 28-54 | 8.66 | 0.15 | 3.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche Diagnostics | 7.598
Roche Diagnostics Cobas Integra 400 | 28-54 | 7.39 | 0.14 | 3.6 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 7.196
Roche Diagnostics Cobas Integra 400 | 25 | 7.46 | 0.04 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche Diagnostics | 3.372
Roche Diagnostics Cobas Integra 400 | 25 | 5.27 | 0.07 | 3.5 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche Diagnostics | 6.93
Roche Diagnostics Cobas Integra 400 | 25 | 8.42 | 0.11 | 2.2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche Diagnostics | 4.422
Roche Diagnostics Cobas Integra 400 | 25 | 10.55 | 0.22 | 2.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche Diagnostics | 4.728
Roche Diagnostics Cobas Integra 400 | 25 | 5.93 | 0.06 | 2.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche Diagnostics | 4.96
Roche Diagnostics Cobas Integra 400 | 50 | 7.22 | -0.03 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 5.458
Roche Diagnostics Cobas Integra 400 | 50 | 6.73 | -0.08 | 2.7 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche Diagnostics | 5.212
Roche Diagnostics Cobas Integra 400 | 50 | 8.77 | 0.12 | 2.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche Diagnostics | 5.608
Roche Diagnostics Cobas Integra 400 | 50 | 9.94 | 0.44 | 3.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche Diagnostics | 7.3
Roche Diagnostics Cobas Integra 400 | 50 | 5.51 | 0.06 | 3.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche Diagnostics | 6.724
Roche Diagnostics Cobas Integra 400 | 47 | 6.96 | -0.19 | 2.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche Diagnostics | 4.122
Roche Diagnostics Cobas Integra 400 | 47 | 5.19 | 0 | 2.9 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche Diagnostics | 5.684
Roche Diagnostics Cobas Integra 400 | 47 | 8.37 | -0.05 | 1.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche Diagnostics | 3.674
Roche Diagnostics Cobas Integra 400 | 47 | 9.94 | 0.15 | 2.2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche Diagnostics | 4.462
Roche Diagnostics Cobas Integra 400 | 47 | 6.07 | -0.05 | 2.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche Diagnostics | 5.634
Roche Diagnostics Cobas Integra 400 | 19 | 6.12 | -0.19 | 2.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche Diagnostics | 5.494
Roche Diagnostics Cobas Integra 400 | 19 | 9.13 | 0.02 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche Diagnostics | 5.116
Roche Diagnostics Cobas Integra 400 | 19 | 5.21 | -0.1 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche Diagnostics | 5.388
Roche Diagnostics Cobas Integra 400 | 19 | 8.31 | -0.04 | 2.3 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche Diagnostics | 4.468
Roche Diagnostics Cobas Integra 400 | 19 | 7.64 | 0.02 | 2.6 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche Diagnostics | 5.116
Roche Diagnostics Cobas Integra 400 | 42 | 9.32 | -0.05 | 2.2 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche Diagnostics | 4.262
Roche Diagnostics Cobas Integra 400 | 42 | 4.94 | 0.02 | 2.4 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche Diagnostics | 4.724
Roche Diagnostics Cobas Integra 400 | 42 | 6.01 | -0.03 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche Diagnostics | 4.674
Roche Diagnostics Cobas Integra 400 | 42 | 5.29 | 0.01 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche Diagnostics | 4.714
Roche Diagnostics Cobas Integra 400 | 42 | 7.93 | -0.15 | 1.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche Diagnostics | 2.986
Roche Diagnostics Cobas Integra 400 | 44 | 5.4 | -0.06 | 4.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche Diagnostics | 8.76
Roche Diagnostics Cobas Integra 400 | 44 | 5.64 | -0.02 | 3.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche Diagnostics | 6.644
Roche Diagnostics Cobas Integra 400 | 44 | 9.32 | 0.01 | 2.6 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche Diagnostics | 5.106
Roche Diagnostics Cobas Integra 400 | 44 | 5.28 | 0 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra 400 | 44 | 7.35 | -0.06 | 3.3 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche Diagnostics | 6.408
Roche Diagnostics Cobas Integra 400 | 23 | 6.37 | -0.04 | 2.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 4.272
Roche Diagnostics Cobas Integra 400 | 23 | 8.72 | 0.12 | 2.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche Diagnostics | 4.628
Roche Diagnostics Cobas Integra 400 | 23 | 5.35 | -0.07 | 3.2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche Diagnostics | 6.202
Roche Diagnostics Cobas Integra 400 | 23 | 7.47 | 0.09 | 2.5 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche Diagnostics | 4.99
Roche Diagnostics Cobas Integra 400 | 23 | 9.9 | 0.15 | 2.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche Diagnostics | 4.854
Roche Diagnostics Cobas Integra 400 | 41 | 5.17 | -0.08 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche Diagnostics | 4.624
Roche Diagnostics Cobas Integra 400 | 41 | 6.36 | -0.05 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche Diagnostics | 3.674
Roche Diagnostics Cobas Integra 400 | 41 | 8.33 | 0.12 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche Diagnostics | 4.236
Roche Diagnostics Cobas Integra 400 | 41 | 5.46 | -0.01 | 1.9 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche Diagnostics | 3.714
Roche Diagnostics Cobas Integra 400 | 41 | 9.54 | 0.06 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche Diagnostics | 3.196
Roche Diagnostics Cobas Integra 400 | 16 | 8.06 | 0.09 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche Diagnostics | 5.382
Roche Diagnostics Cobas Integra 400 | 16 | 5.83 | -0.06 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche Diagnostics | 4.644
Roche Diagnostics Cobas Integra 400 | 16 | 5.15 | 0.02 | 2.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche Diagnostics | 4.136
Roche Diagnostics Cobas Integra 400 | 16 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche Diagnostics | 0
Roche Diagnostics Cobas Integra 400 | 16 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche Diagnostics | 0
Roche Diagnostics Cobas Integra 400 | 19 | 7.85 | 0.04 | 2.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche Diagnostics | 4.744
Roche Diagnostics Cobas Integra 400 | 19 | 5.86 | -0.05 | 3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche Diagnostics | 5.83
Roche Diagnostics Cobas Integra 400 | 19 | 8.78 | 0.09 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche Diagnostics | 4.402
Roche Diagnostics Cobas Integra 400 | 19 | 5.85 | -0.01 | 3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche Diagnostics | 5.87
Roche Diagnostics Cobas Integra 400 | 19 | 9.66 | 0.1 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche Diagnostics | 4.02
Roche Diagnostics Cobas Integra 400 | 36 | 5.47 | -0.01 | 2.6 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche Diagnostics | 5.086
Roche Diagnostics Cobas Integra 400 | 36 | 8.41 | 0.15 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche Diagnostics | 4.462
Roche Diagnostics Cobas Integra 400 | 36 | 5.19 | 0.04 | 2.4 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche Diagnostics | 4.744
Roche Diagnostics Cobas Integra 400 | 36 | 10.3 | 0.25 | 2.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche Diagnostics | 4.758
Roche Diagnostics Cobas Integra 400 | 36 | 6.01 | 0.02 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche Diagnostics | 4.92
Roche Diagnostics Cobas Integra 400 | 36 | 8.85 | -0.01 | 3.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche Diagnostics | 6.654
Roche Diagnostics Cobas Integra 400 | 36 | 6.76 | -0.08 | 3.2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche Diagnostics | 6.192
Roche Diagnostics Cobas Integra 400 | 36 | 9.48 | 0.09 | 2.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche Diagnostics | 5.578
Roche Diagnostics Cobas Integra 400 | 36 | 6.07 | -0.06 | 3 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche Diagnostics | 5.82
Roche Diagnostics Cobas Integra 400 | 36 | 7.38 | -0.14 | 4.2 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche Diagnostics | 8.092
Roche Diagnostics Cobas Integra 400 | 18 | 5.75 | -0.05 | 3.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche Diagnostics | 6.222
Roche Diagnostics Cobas Integra 400 | 18 | 9.27 | 0.15 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche Diagnostics | 4.658
Roche Diagnostics Cobas Integra 400 | 18 | 5.45 | -0.02 | 2.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche Diagnostics | 4.88
Roche Diagnostics Cobas Integra 400 | 18 | 7.34 | 0.09 | 2.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche Diagnostics | 4.99
Roche Diagnostics Cobas Integra 400 | 18 | 9.43 | 0.17 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche Diagnostics | 4.874
Roche Diagnostics Cobas Integra 400 | 32 | 5.18 | -0.03 | 2.6 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche Diagnostics | 5.066
Roche Diagnostics Cobas Integra 400 | 32 | 6.65 | -0.03 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche Diagnostics | 5.066
Roche Diagnostics Cobas Integra 400 | 32 | 8.5 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Roche Diagnostics | 4.312
Roche Diagnostics Cobas Integra 400 | 32 | 5.98 | 0 | 2.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche Diagnostics | 4.312
Roche Diagnostics Cobas Integra 400 | 32 | 10.31 | 0.19 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche Diagnostics | 4.502
Roche Diagnostics Cobas Integra 400 (non TQ) | 20 | 5.3 |  | 2.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics | 5.684
Roche Diagnostics Cobas Integra 400 (non TQ) | 20 | 8.7 |  | 3.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra 400 (non TQ) | 20 | 6.2 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra 400 Gen.2 | 26 | 5.2 |  | 4.1 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics | 8.036
Roche Diagnostics Cobas Integra 400 Gen.2 | 26 | 8.5 |  | 2.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics | 4.704
Roche Diagnostics Cobas Integra 400 Gen.2 | 26 | 6.1 |  | 3.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas Integra 800 | 139 | 6.78 | 0.18 | 3.2 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche Diagnostics | 6.452
Roche Diagnostics Cobas Integra 800 | 139 | 9.59 | 0.09 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche Diagnostics | 5.578
Roche Diagnostics Cobas Integra 800 | 139 | 7.55 | 0.15 | 3.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche Diagnostics | 6.618
Roche Diagnostics Cobas Integra 800 | 115 | 6.07 | 0.17 | 3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche Diagnostics | 6.05
Roche Diagnostics Cobas Integra 800 | 115 | 9.8 | 0 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche Diagnostics | 4.9
Roche Diagnostics Cobas Integra 800 | 115 | 7.55 | 0.15 | 2.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche Diagnostics | 5.05
Roche Diagnostics Cobas Integra 800 | 118 | 5.39 | 0.19 | 3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche Diagnostics | 6.07
Roche Diagnostics Cobas Integra 800 | 118 | 8.74 | 0.04 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche Diagnostics | 4.94
Roche Diagnostics Cobas Integra 800 | 118 | 6.41 | 0.11 | 2.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche Diagnostics | 5.206
Roche Diagnostics Cobas Integra 800 | 111 | 8.55 | 0.05 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche Diagnostics | 6.518
Roche Diagnostics Cobas Integra 800 | 111 | 5.61 | 0.21 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche Diagnostics | 6.09
Roche Diagnostics Cobas Integra 800 | 111 | 6.57 | 0.17 | 3.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche Diagnostics | 6.246
Roche Diagnostics Cobas Integra 800 | 47 | 6.33 | 0.03 | 2.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche Diagnostics | 5.126
Roche Diagnostics Cobas Integra 800 | 47 | 7.71 | 0.11 | 2.8 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche Diagnostics | 5.598
Roche Diagnostics Cobas Integra 800 | 47 | 9.34 | 0.14 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche Diagnostics | 5.432
Roche Diagnostics Cobas Integra 800 | 144 | 5.65 | 0.05 | 2.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche Diagnostics | 4.166
Roche Diagnostics Cobas Integra 800 | 144 | 9.51 | 0.11 | 2.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche Diagnostics | 4.618
Roche Diagnostics Cobas Integra 800 | 144 | 7.24 | 0.04 | 2.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche Diagnostics | 4.352
Roche Diagnostics Cobas Integra 800 | 145 | 5.51 | 0.11 | 2.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche Diagnostics | 5.01
Roche Diagnostics Cobas Integra 800 | 145 | 8.26 | -0.04 | 2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche Diagnostics | 3.88
Roche Diagnostics Cobas Integra 800 | 145 | 5.65 | 0 | 2.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche Diagnostics | 4.704
Roche Diagnostics Cobas Integra 800 | 118 | 7.03 | -0.08 | 1.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche Diagnostics | 3.448
Roche Diagnostics Cobas Integra 800 | 118 | 9.35 | 0.03 | 2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche Diagnostics | 3.95
Roche Diagnostics Cobas Integra 800 | 118 | 6.08 | 0.01 | 2.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche Diagnostics | 4.126
Roche Diagnostics Cobas Integra 800 | 116 | 7.91 | -0.14 | 2.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche Diagnostics | 4.76
Roche Diagnostics Cobas Integra 800 | 116 | 5.39 | 0.09 | 2.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche Diagnostics | 5.382
Roche Diagnostics Cobas Integra 800 | 116 | 6.18 | 0.04 | 1.9 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche Diagnostics | 3.764
Roche Diagnostics Cobas Integra 800 | 135 | 6.68 | 0.1 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche Diagnostics | 5.588
Roche Diagnostics Cobas Integra 800 | 135 | 8.63 | 0.24 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche Diagnostics | 4.748
Roche Diagnostics Cobas Integra 800 | 135 | 5.76 | 0.11 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche Diagnostics | 5.206
Roche Diagnostics Cobas Integra 800 | 136 | 6.57 | 0.08 | 2.2 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche Diagnostics | 4.392
Roche Diagnostics Cobas Integra 800 | 136 | 7.05 | 0.08 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche Diagnostics | 4.392
Roche Diagnostics Cobas Integra 800 | 136 | 9.79 | 0.14 | 2.1 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche Diagnostics | 4.256
Roche Diagnostics Cobas Integra 800 | 125 | 6.8 | 0.01 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche Diagnostics | 5.302
Roche Diagnostics Cobas Integra 800 | 125 | 10.28 | 0 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche Diagnostics | 5.096
Roche Diagnostics Cobas Integra 800 | 125 | 6.81 | -0.01 | 2.2 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche Diagnostics | 4.302
Roche Diagnostics Cobas Integra 800 | 110 | 8.67 | 0.04 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche Diagnostics | 4.548
Roche Diagnostics Cobas Integra 800 | 110 | 5.48 | 0.16 | 2.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche Diagnostics | 5.06
Roche Diagnostics Cobas Integra 800 | 108 | 5.89 | 0.05 | 2.4 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche Diagnostics | 4.754
Roche Diagnostics Cobas Integra 800 | 108 | 11.75 | 0.04 | 2.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche Diagnostics | 4.352
Roche Diagnostics Cobas Integra 800 | 108 | 9.6 | 0.07 | 2.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche Diagnostics | 4.186
Roche Diagnostics Cobas Integra 800 | 108 | 5.23 | 0.19 | 2.4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche Diagnostics | 4.894
Roche Diagnostics Cobas Integra 800 | 108 | 7.36 | -0.02 | 2.2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche Diagnostics | 4.292
Roche Diagnostics Cobas Integra 800 | 106-120 | 11.76 | 0.07 | 2.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche Diagnostics | 4.97
Roche Diagnostics Cobas Integra 800 | 106-120 | 6.02 | 0.09 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche Diagnostics | 4.99
Roche Diagnostics Cobas Integra 800 | 106-120 | 5.36 | 0.19 | 2.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche Diagnostics | 5.286
Roche Diagnostics Cobas Integra 800 | 106-120 | 8.17 | 0.03 | 2.2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche Diagnostics | 4.342
Roche Diagnostics Cobas Integra 800 | 106-120 | 9.24 | -0.06 | 2.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche Diagnostics | 4.84
Roche Diagnostics Cobas Integra 800 | 105-120 | 5.4 | 0.08 | 2.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche Diagnostics | 5.176
Roche Diagnostics Cobas Integra 800 | 105-120 | 9.2 | 0.04 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche Diagnostics | 4.352
Roche Diagnostics Cobas Integra 800 | 105-120 | 5.41 | 0.1 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche Diagnostics | 4.412
Roche Diagnostics Cobas Integra 800 | 105-120 | 12.11 | 0.08 | 2.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche Diagnostics | 4.588
Roche Diagnostics Cobas Integra 800 | 105-120 | 6 | 0.06 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche Diagnostics | 4.568
Roche Diagnostics Cobas Integra 800 | 104 | 5.37 | 0.1 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche Diagnostics | 5.196
Roche Diagnostics Cobas Integra 800 | 104 | 10.64 | 0.05 | 1.7 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche Diagnostics | 3.382
Roche Diagnostics Cobas Integra 800 | 104 | 6.19 | -0.01 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche Diagnostics | 4.498
Roche Diagnostics Cobas Integra 800 | 104 | 12.27 | 0.04 | 1.9 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche Diagnostics | 3.764
Roche Diagnostics Cobas Integra 800 | 104 | 7.49 | -0.02 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche Diagnostics | 4.488
Roche Diagnostics Cobas Integra 800 | 104-118 | 9.14 | 0.03 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche Diagnostics | 4.538
Roche Diagnostics Cobas Integra 800 | 104-118 | 6.06 | 0.05 | 2.1 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche Diagnostics | 4.166
Roche Diagnostics Cobas Integra 800 | 104-118 | 11.81 | 0.1 | 1.9 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche Diagnostics | 3.824
Roche Diagnostics Cobas Integra 800 | 104-118 | 5.2 | 0.18 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche Diagnostics | 4.688
Roche Diagnostics Cobas Integra 800 | 104-118 | 7.61 | 0.03 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche Diagnostics | 4.146
Roche Diagnostics Cobas Integra 800 | 77-91 | 6.43 | 0.02 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche Diagnostics | 5.116
Roche Diagnostics Cobas Integra 800 | 77-91 | 9.61 | 0.08 | 2.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche Diagnostics | 4.196
Roche Diagnostics Cobas Integra 800 | 77-91 | 5.48 | 0.14 | 2.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche Diagnostics | 5.04
Roche Diagnostics Cobas Integra 800 | 77-91 | 8.59 | 0.08 | 1.9 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche Diagnostics | 3.804
Roche Diagnostics Cobas Integra 800 | 77-91 | 7.36 | 0.11 | 1.8 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 3.638
Roche Diagnostics Cobas Integra 800 | 61 | 7.42 | 0 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche Diagnostics | 5.096
Roche Diagnostics Cobas Integra 800 | 61 | 5.26 | 0.06 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche Diagnostics | 5.156
Roche Diagnostics Cobas Integra 800 | 61 | 8.36 | 0.05 | 2.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche Diagnostics | 4.166
Roche Diagnostics Cobas Integra 800 | 61 | 10.48 | 0.15 | 2.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche Diagnostics | 4.658
Roche Diagnostics Cobas Integra 800 | 61 | 5.94 | 0.07 | 2.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche Diagnostics | 5.362
Roche Diagnostics Cobas Integra 800 | 38 | 7.22 | -0.03 | 1.7 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche Diagnostics | 3.302
Roche Diagnostics Cobas Integra 800 | 38 | 6.79 | -0.02 | 2.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche Diagnostics | 5.076
Roche Diagnostics Cobas Integra 800 | 38 | 8.69 | 0.04 | 2.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche Diagnostics | 4.744
Roche Diagnostics Cobas Integra 800 | 38 | 9.6 | 0.1 | 2.8 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche Diagnostics | 5.588
Roche Diagnostics Cobas Integra 800 | 38 | 5.55 | 0.1 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche Diagnostics | 5.392
Roche Diagnostics Cobas Integra 800 | 14 | 7.01 | -0.14 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche Diagnostics | 3.192
Roche Diagnostics Cobas Integra 800 | 14 | 5.29 | 0.1 | 2.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche Diagnostics | 5.588
Roche Diagnostics Cobas Integra 800 | 14 | 8.34 | -0.08 | 1.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche Diagnostics | 3.056
Roche Diagnostics Cobas Integra 800 | 14 | 9.85 | 0.06 | 1.9 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche Diagnostics | 3.784
Roche Diagnostics Cobas Integra 800 | 14 | 6.12 | 0 | 1.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche Diagnostics | 3.332
Roche Diagnostics Cobas Integra 800 (non TQ) | 11 | 5.2 |  | 4.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics | 9.604
Roche Diagnostics Cobas Integra 800 (non TQ) | 11 | 8.5 |  | 3.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics | 6.076
Roche Diagnostics Cobas Integra 800 (non TQ) | 11 | 6.1 |  | 3.4 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra 800 Gen.2 | 116 | 5.3 |  | 3.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics | 7.644
Roche Diagnostics Cobas Integra 800 Gen.2 | 116 | 8.5 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas Integra 800 Gen.2 | 116 | 6.1 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra Gen.2 | 20 | 6.9 |  | 3.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche Diagnostics | 6.076
Roche Diagnostics Cobas Integra Gen.2 | 20 | 7 |  | 2.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics Cobas Integra Gen.2 | 20 | 11.6 |  | 4.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche Diagnostics | 8.036
Roche Diagnostics Cobas Integra Gen.2 | 72 | 5.5 |  | 3.3 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche Diagnostics | 6.468
Roche Diagnostics Cobas Integra Gen.2 | 72 | 11.7 |  | 2.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics Cobas Integra Gen.2 | 72 | 7.5 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche Diagnostics | 4.508
Roche Diagnostics Cobas Integra Gen.2 | 81 | 9 |  | 2.6 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche Diagnostics | 5.096
Roche Diagnostics Cobas Integra Gen.2 | 81 | 11 |  | 3.2 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche Diagnostics | 6.272
Roche Diagnostics Cobas Integra Gen.2 | 81 | 6.2 |  | 2.7 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche Diagnostics | 5.292
Roche Diagnostics Cobas Integra Gen.2 | 90 | 12.5 |  | 3.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics Cobas Integra Gen.2 | 90 | 5.4 |  | 3.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche Diagnostics | 6.272
Roche Diagnostics Cobas Integra Gen.2 | 90 | 8.8 |  | 2.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche Diagnostics | 5.684
Roche Diagnostics Cobas Integra Gen.2 | 101 | 11.6 |  | 3.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra Gen.2 | 101 | 6.8 |  | 3.4 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche Diagnostics | 6.664
Roche Diagnostics Cobas Integra Gen.2 | 101 | 5.2 |  | 2.8 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche Diagnostics | 5.488
Roche Diagnostics | 100 | 5.9839 | -0.4 | 5.2 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Roche Diagnostics | 9.792
Roche Diagnostics | 103 | 7.457 | 0.1 | 4.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Roche Diagnostics | 9.312
Roche Diagnostics | 288 | 5.8924 | -1.4 | 4.7 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Roche Diagnostics | 7.812
Roche Diagnostics | 287 | 7.4112 | -0.4 | 4.1 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Roche Diagnostics | 7.636
Roche Diagnostics | 135 | 6.3498 | 1.7 | 5 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Roche Diagnostics | 11.5
Roche Diagnostics | 135 | 7.5851 | 1.4 | 4.7 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Roche Diagnostics | 10.612
Roche Diagnostics | 319 | 6.176 | -0.2 | 3.8 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Roche Diagnostics | 7.248
Roche Diagnostics | 318 | 7.4936 | 0.4 | 3.5 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Roche Diagnostics | 7.26
Roche Diagnostics | 239 | 7.6491 | 2 | 4.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Roche Diagnostics | 11.604
Roche Diagnostics | 239 | 6.112 | 1.1 | 5.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Roche Diagnostics | 11.684
Roche Diagnostics | 597 | 7.4661 | 0 | 3.6 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Roche Diagnostics | 7.056
Roche Diagnostics | 598 | 5.993 | -0.2 | 4 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Roche Diagnostics | 7.64
Roche Diagnostics Cobas b 101 | 6 | 8.7288 | 1.8 | 3.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 9.248
Roche Diagnostics Cobas b 101 | 6 | 5.8374 | 2.1 | 4.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 10.136
Roche Diagnostics Cobas c 111 | 11 | 5.8283 | 3.7 | 6.1 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 15.656
Roche Diagnostics Cobas c 111 | 11 | 5.7917 | 3.3 | 4.3 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 11.728
Roche Diagnostics Cobas c 311 | 8 | 5.5904 | 1.1 | 9.3 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 19.328
Roche Diagnostics Cobas c 311 | 8 | 5.627 | 1.5 | 8.4 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 17.964
Roche Diagnostics Cobas c 311/c111 | 20 | 8.6831 | 1.3 | 3.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 8.356
Roche Diagnostics Cobas c 311/c111 | 20 | 5.5995 | -0.5 | 5.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 10.084
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 203 | 9.0308 | 5.1 | 3.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Roche Diagnostics | 12.352
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 199 | 5.7459 | 1.1 | 6.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Roche Diagnostics | 13.056
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 387 | 8.7014 | 1.5 | 3.5 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 8.36
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 387 | 5.6361 | -0.1 | 4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 7.74
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 214 | 5.691 | 2.2 | 6.2 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 14.352
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 213 | 5.6819 | 2.1 | 6.3 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 14.448
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 230 | 5.5629 | 0.8 | 4.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 8.836
Roche Diagnostics Cobas c 501/502 (part of Cobas 6000/8000) | 232 | 5.5629 | 0.8 | 4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 8.64
Roche Diagnostics Cobas c 503 (Cobas pro) | 7 | 5.7276 | 2.6 | 4.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 11.42
Roche Diagnostics Cobas c 503 (Cobas pro) | 7 | 5.7734 | 3.1 | 4.6 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 12.116
Roche Diagnostics Cobas c513 | 93 | 8.7197 | 1.7 | 2.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 6.992
Roche Diagnostics Cobas c513 | 93 | 5.6727 | 0.3 | 2.3 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 4.808
Roche Diagnostics Cobas c513 | 16 | 5.6819 | 2.1 | 2.8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 7.588
Roche Diagnostics Cobas c513 | 16 | 5.6544 | 1.8 | 2.8 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 7.288
Roche Diagnostics Cobas c513 | 38 | 5.5538 | 0.7 | 2.3 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 5.208
Roche Diagnostics Cobas c513 | 38 | 5.5355 | 0.5 | 2.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 4.616
Roche Diagnostics Cobas c311/c111 | 12 | 8.9667 | 4.4 | 6.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Roche Diagnostics | 16.748
Roche Diagnostics Cobas c311/c111 | 13 | 5.6361 | -0.1 | 10.3 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Roche Diagnostics | 20.088
Roche Diagnostics Cobas Integra | 39 | 8.9667 | 4.4 | 3.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Roche Diagnostics | 10.868
Roche Diagnostics Cobas Integra | 39 | 5.7459 | 1.1 | 7 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Roche Diagnostics | 14.82
Roche Diagnostics Cobas Integra | 98 | 8.7105 | 1.6 | 3.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 8.068
Roche Diagnostics Cobas Integra | 98 | 5.6178 | -0.3 | 3.5 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 6.56
Roche Diagnostics Cobas Integra | 51 | 5.6544 | 1.8 | 4.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 10.62
Roche Diagnostics Cobas Integra | 50 | 5.691 | 2.2 | 4.6 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 11.216
Roche Diagnostics Cobas Integra | 46 | 5.5355 | 0.5 | 4.2 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 8.732
Roche Diagnostics Cobas Integra | 46 | 5.5629 | 0.8 | 4.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 8.836
Roche Diagnostics not specified/other | 22 | 8.9484 | 4.2 | 3.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Roche Diagnostics | 11.844
Roche Diagnostics not specified/other | 23 | 5.8008 | 1.7 | 4 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Roche Diagnostics | 9.54
Roche Diagnostics not specified/other | 33 | 8.6648 | 1.1 | 3.3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Roche Diagnostics | 7.568
Roche Diagnostics not specified/other | 34 | 5.6178 | -0.3 | 3.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Roche Diagnostics | 5.776
Roche Diagnostics not specified/other | 93 | 5.4989 | 0.1 | 8.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 16.76
Roche Diagnostics not specified/other | 87 | 5.4623 | -0.3 | 9.2 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Roche Diagnostics | 17.732
Roche Diagnostics not specified/other | 384 | 5.5446 | 0.6 | 3.7 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 7.852
Roche Diagnostics not specified/other | 382 | 5.5446 | 0.6 | 4.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Roche Diagnostics | 8.636
Roche Diagnostics/Hitachi (Tina Quant II) | 62 | 10.6 |  | 3.3 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.468
Roche Diagnostics/Hitachi (Tina Quant II) | 62 | 5.4 |  | 4.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.82
Roche Diagnostics/Hitachi (Tina Quant II) | 62 | 8 |  | 4.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.212
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 6.8 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.996
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 6.8 |  | 4.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.408
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 11.5 |  | 3.8 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.448
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 6.8 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.996
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 6.8 |  | 4.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.408
Roche Diagnostics/Hitachi (Tina Quant II) | 53 | 11.5 |  | 3.8 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.448
Roche Diagnostics/Hitachi (Tina Quant II) | 63 | 5.5 |  | 5.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 10.192
Roche Diagnostics/Hitachi (Tina Quant II) | 63 | 11.8 |  | 4.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.232
Roche Diagnostics/Hitachi (Tina Quant II) | 63 | 7.5 |  | 3.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.86
Roche Diagnostics/Hitachi (Tina Quant II) | 44 | 9.1 |  | 3.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.86
Roche Diagnostics/Hitachi (Tina Quant II) | 44 | 11.2 |  | 4.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.82
Roche Diagnostics/Hitachi (Tina Quant II) | 44 | 6.3 |  | 5.1 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.996
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 13.2 |  | 4.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.82
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 5.4 |  | 3.6 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.056
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 8.6 |  | 4.2 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.232
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 11.7 |  | 3.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.076
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 6.8 |  | 4.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.016
Roche Diagnostics/Hitachi (Tina Quant II) | 41 | 5.2 |  | 4.5 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.82
Roche Diagnostics/Hitachi Modular P | 20 | 5.2 |  | 3.2 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.272
Roche Diagnostics/Hitachi Modular P | 20 | 8.3 |  | 3.2 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.272
Roche Diagnostics/Hitachi Modular P | 20 | 6 |  | 4.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 9.212
Roche Diagnostics/Hitachi Modular P | 15 | 6.61 | 0.01 | 3.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.87
Roche Diagnostics/Hitachi Modular P | 15 | 9.18 | -0.32 | 3.3 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.148
Roche Diagnostics/Hitachi Modular P | 15 | 7.27 | -0.13 | 4.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.494
Roche Diagnostics/Hitachi Modular P | 14 | 5.95 | 0.05 | 2.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 5.734
Roche Diagnostics/Hitachi Modular P | 14 | 9.63 | -0.17 | 4.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.866
Roche Diagnostics/Hitachi Modular P | 14 | 7.24 | -0.16 | 2.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 3.956
Roche Diagnostics/Hitachi Modular P | 17 | 5.29 | 0.09 | 4.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 8.518
Roche Diagnostics/Hitachi Modular P | 17 | 8.4 | -0.3 | 3.3 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.168
Roche Diagnostics/Hitachi Modular P | 17 | 6.22 | -0.08 | 3.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche Diagnostics/Hitachi | 6.976
Roche Diagnostics/Hitachi Modular P | 14 | 8.26 | -0.24 | 5.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 10.54
Roche Diagnostics/Hitachi Modular P | 14 | 5.41 | 0.01 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.654
Roche Diagnostics/Hitachi Modular P | 14 | 6.31 | -0.09 | 4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.75
Roche Diagnostics/Hitachi Modular P | 127 | 6.31 | 0.01 | 1.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 3.538
Roche Diagnostics/Hitachi Modular P | 127 | 7.68 | 0.08 | 2.1 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 4.196
Roche Diagnostics/Hitachi Modular P | 127 | 9.37 | 0.17 | 2.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche Diagnostics/Hitachi | 4.482
Roche Diagnostics/Hitachi Modular P | 14 | 5.61 | 0.01 | 2.9 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche Diagnostics/Hitachi | 5.694
Roche Diagnostics/Hitachi Modular P | 14 | 9.29 | -0.11 | 3.8 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche Diagnostics/Hitachi | 7.338
Roche Diagnostics/Hitachi Modular P | 14 | 7.21 | 0.01 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche Diagnostics/Hitachi | 4.714
Sebia Capillarys 2 | 53 | 7.1 | -0.15 | 1.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Sebia | 2.79
Sebia Capillarys 2 | 53 | 6.7 | -0.11 | 1.9 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Sebia | 3.614
Sebia Capillarys 2 | 53 | 8.5 | -0.15 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Sebia | 2.398
Sebia Capillarys 2 | 53 | 9.33 | -0.17 | 1.6 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Sebia | 2.966
Sebia Capillarys 2 | 53 | 5.34 | -0.11 | 2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Sebia | 3.81
Sebia Capillarys 2 | 53 | 7.09 | -0.06 | 1.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Sebia | 3.076
Sebia Capillarys 2 | 53 | 5.12 | -0.07 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Sebia | 3.85
Sebia Capillarys 2 | 53 | 8.3 | -0.12 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Sebia | 2.624
Sebia Capillarys 2 | 53 | 9.54 | -0.25 | 1.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Sebia | 2.886
Sebia Capillarys 2 | 53 | 6.03 | -0.09 | 1.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Sebia | 3.242
Sebia Capillarys 2 | 52 | 6.25 | -0.06 | 1.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Sebia | 3.664
Sebia Capillarys 2 | 52 | 8.94 | -0.17 | 1.3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Sebia | 2.378
Sebia Capillarys 2 | 52 | 5.23 | -0.08 | 1.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Sebia | 3.644
Sebia Capillarys 2 | 52 | 8.33 | -0.02 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Sebia | 3.508
Sebia Capillarys 2 | 52 | 7.61 | -0.01 | 1.4 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Sebia | 2.734
Sebia Capillarys 2 | 131 | 8.5641 | 0 | 2.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Sebia | 4.704
Sebia Capillarys 2 | 135 | 5.6178 | -0.3 | 3.8 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Sebia | 7.148
Sebia Capillarys 2 | 45 | 5.8832 | -1.5 | 2.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Sebia | 3.4
Sebia Capillarys 2 | 46 | 7.3197 | -1.4 | 2.5 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Sebia | 3.5
Sebia Capillarys 2 | 57 | 5.9381 | -0.9 | 2.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Sebia | 4.784
Sebia Capillarys 2 | 57 | 7.4478 | 0 | 2.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Sebia | 4.312
Sebia Capillarys 2 | 59 | 6.2767 | 0.9 | 3.7 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Sebia | 8.152
Sebia Capillarys 2 | 59 | 7.4021 | -0.6 | 2.6 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Sebia | 4.496
Sebia Capillarys 2 | 64 | 6.2035 | 0.1 | 2.9 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Sebia | 5.784
Sebia Capillarys 2 | 64 | 7.4295 | -0.3 | 2.4 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Sebia | 4.404
Sebia Capillarys 2 | 193 | 8.6648 | 1.1 | 3 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Sebia | 6.98
Sebia Capillarys 2 | 192 | 5.6178 | -0.3 | 3.7 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Sebia | 6.952
Sebia Capillarys 2 | 122 | 5.4623 | -0.3 | 3.7 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 6.952
Sebia Capillarys 2 | 120 | 5.4623 | -0.3 | 4.2 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 7.932
Sebia Capillarys 2 | 128 | 5.4806 | -0.1 | 3.9 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 7.544
Sebia Capillarys 2 | 129 | 5.4989 | 0.1 | 3.9 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 7.744
Sebia Capillarys 2 Flex Piercing | 73 | 5.59 | -0.07 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Sebia | 5.418
Sebia Capillarys 2 Flex Piercing | 73 | 9.28 | -0.03 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Sebia | 2.518
Sebia Capillarys 2 Flex Piercing | 73 | 5.18 | -0.1 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Sebia | 4.604
Sebia Capillarys 2 Flex Piercing | 73 | 7.33 | -0.08 | 1.4 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Sebia | 2.664
Sebia Capillarys 2 Flex Piercing | 71 | 9.21 | -0.16 | 1.6 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Sebia | 2.976
Sebia Capillarys 2 Flex Piercing | 71 | 4.93 | 0.01 | 2.2 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Sebia | 4.322
Sebia Capillarys 2 Flex Piercing | 71 | 6.07 | 0.03 | 1.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Sebia | 2.97
Sebia Capillarys 2 Flex Piercing | 71 | 5.26 | -0.02 | 1.9 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Sebia | 3.704
Sebia Capillarys 2 Flex Piercing | 71 | 8.03 | -0.05 | 2.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Sebia | 4.066
Sebia Capillarys 2 Flex Piercing | 63 | 6.39 | -0.02 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Sebia | 3.704
Sebia Capillarys 2 Flex Piercing | 63 | 8.58 | -0.02 | 1.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Sebia | 2.528
Sebia Capillarys 2 Flex Piercing | 63 | 5.36 | -0.06 | 1.9 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Sebia | 3.664
Sebia Capillarys 2 Flex Piercing | 63 | 7.38 | 0 | 1.3 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Sebia | 2.548
Sebia Capillarys 2 Flex Piercing | 63 | 9.52 | -0.23 | 1.5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Sebia | 2.71
Sebia Capillarys 2 Flex Piercing | 67 | 5.18 | -0.07 | 1.8 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Sebia | 3.458
Sebia Capillarys 2 Flex Piercing | 67 | 6.33 | -0.08 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Sebia | 3.056
Sebia Capillarys 2 Flex Piercing | 67 | 8.12 | -0.09 | 1.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Sebia | 2.654
Sebia Capillarys 2 Flex Piercing | 67 | 5.43 | -0.04 | 1.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Sebia | 3.292
Sebia Capillarys 2 Flex Piercing | 67 | 9.46 | -0.02 | 1.5 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Sebia | 2.92
Sebia Capillarys 2 Flex Piercing | 62 | 7.91 | -0.06 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Sebia | 3.86
Sebia Capillarys 2 Flex Piercing | 62 | 5.87 | -0.02 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Sebia | 3.9
Sebia Capillarys 2 Flex Piercing | 62 | 5.1 | -0.03 | 2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Sebia | 3.89
Sebia Capillarys 2 Flex Piercing | 62 | 7.24 | -0.16 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Sebia | 3.564
Sebia Capillarys 2 Flex Piercing | 62 | 9.14 | -0.03 | 1.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Sebia | 2.518
Sebia Capillarys 2 Flex Piercing | 61 | 7.77 | -0.04 | 1.5 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Sebia | 2.9
Sebia Capillarys 2 Flex Piercing | 61 | 5.82 | -0.09 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Sebia | 3.438
Sebia Capillarys 2 Flex Piercing | 61 | 8.6 | -0.09 | 1.5 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Sebia | 2.85
Sebia Capillarys 2 Flex Piercing | 61 | 5.82 | -0.04 | 1.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Sebia | 3.096
Sebia Capillarys 2 Flex Piercing | 61 | 9.38 | -0.18 | 1.2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Sebia | 2.172
Sebia Capillarys 2 Flex Piercing | 61 | 5.4 | -0.08 | 2.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Sebia | 4.232
Sebia Capillarys 2 Flex Piercing | 61 | 8.2 | -0.06 | 1.5 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Sebia | 2.88
Sebia Capillarys 2 Flex Piercing | 61 | 5.09 | -0.06 | 2.1 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Sebia | 4.056
Sebia Capillarys 2 Flex Piercing | 61 | 9.95 | -0.1 | 1.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Sebia | 2.252
Sebia Capillarys 2 Flex Piercing | 61 | 5.94 | -0.05 | 1.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Sebia | 2.302
Sebia Capillarys 2 Flex Piercing | 60 | 8.79 | -0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Sebia | 3.262
Sebia Capillarys 2 Flex Piercing | 60 | 6.69 | -0.15 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Sebia | 3.574
Sebia Capillarys 2 Flex Piercing | 60 | 9.24 | -0.15 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Sebia | 3.378
Sebia Capillarys 2 Flex Piercing | 60 | 5.95 | -0.18 | 2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Sebia | 3.74
Sebia Capillarys 2 Flex Piercing | 60 | 7.37 | -0.15 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Sebia | 2.986
Sebia Capillarys 2 Flex Piercing | 53 | 5.72 | -0.08 | 2.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Sebia | 4.036
Sebia Capillarys 2 Flex Piercing | 53 | 8.93 | -0.19 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Sebia | 2.75
Sebia Capillarys 2 Flex Piercing | 53 | 5.43 | -0.04 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Sebia | 4.468
Sebia Capillarys 2 Flex Piercing | 53 | 7.07 | -0.18 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Sebia | 2.368
Sebia Capillarys 2 Flex Piercing | 53 | 9.13 | -0.13 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Sebia | 2.614
Sebia Capillarys 2 Flex Piercing | 55 | 5.13 | -0.08 | 2.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Sebia | 4.428
Sebia Capillarys 2 Flex Piercing | 55 | 6.7 | 0.02 | 1.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Sebia | 3.352
Sebia Capillarys 2 Flex Piercing | 55 | 8.05 |  | 1.6 | GH5-13 |  | 2021 | CAP | Fresh | Sebia | 3.136
Sebia Capillarys 2 Flex Piercing | 55 | 5.98 | 0 | 1.7 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Sebia | 3.332
Sebia Capillarys 2 Flex Piercing | 55 | 10.03 | -0.09 | 1.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Sebia | 3.046
Sebia Capillarys 2 Flex Piercing | 73 | 5.26 | -0.2 | 3.1 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Sebia | 5.876
Sebia Capillarys 2 Flex Piercing | 36 | 7.29 | -0.13 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Sebia | 3.202
Sebia Capillarys 2 Flex Piercing | 36 | 5.05 | -0.15 | 2.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Sebia | 4.162
Sebia Capillarys 2 Flex Piercing | 36 | 8.28 | -0.03 | 1.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Sebia | 2.714
Sebia Capillarys 2 Flex Piercing | 36 | 10.26 | -0.07 | 1.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Sebia | 2.478
Sebia Capillarys 2 Flex Piercing | 36 | 5.74 | -0.13 | 2.1 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Sebia | 3.986
Sebia Capillarys 2 Flex Piercing | 32-44 | 6.34 | -0.07 | 2.2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Sebia | 4.242
Sebia Capillarys 2 Flex Piercing | 32-44 | 9.44 | -0.09 | 1.2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Sebia | 2.262
Sebia Capillarys 2 Flex Piercing | 32-44 | 5.24 | -0.1 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Sebia | 3.232
Sebia Capillarys 2 Flex Piercing | 32-44 | 8.52 | 0.01 | 1.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Sebia | 2.754
Sebia Capillarys 2 Flex Piercing | 32-44 | 7.14 | -0.11 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Sebia | 4.398
SEBIA Capillarys 2 flex-piercing | 56 | 7.4936 | 0.3 | 2.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Sebia | 5.788
SEBIA Capillarys 2 flex-piercing | 57 | 6.0296 | 0.2 | 3.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Sebia | 7.844
SEBIA Capillarys 2 flex-piercing | 103 | 7.4753 | 0.1 | 2.6 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Sebia | 5.196
SEBIA Capillarys 2 flex-piercing | 102 | 6.0296 | 0.2 | 3.2 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Sebia | 6.472
Sebia Capillarys 2/ Minicap | 15 | 6.43 | -0.15 | 2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Sebia | 3.77
Sebia Capillarys 2/ Minicap | 15 | 8.32 | -0.07 | 1.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Sebia | 2.478
Sebia Capillarys 2/ Minicap | 15 | 5.43 | -0.22 | 1.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Sebia | 1.936
Sebia Capillarys 2/ Minicap | 12 | 6.38 | -0.11 | 1.4 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Sebia | 2.634
Sebia Capillarys 2/ Minicap | 12 | 6.86 | -0.11 | 1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Sebia | 1.85
Sebia Capillarys 2/ Minicap | 12 | 9.55 | -0.1 | 0.5 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Sebia | 0.88
Sebia Capillarys 2/ Minicap | 20 | 6.62 | -0.17 | 1.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Sebia | 2.966
Sebia Capillarys 2/ Minicap | 20 | 10.04 | -0.24 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Sebia | 4.66
Sebia Capillarys 2/ Minicap | 20 | 6.6 | -0.22 | 2.1 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Sebia | 3.896
Sebia Capillarys 2/ Minicap | 12 | 8.58 | -0.05 | 1.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Sebia | 3.282
Sebia Capillarys 2/ Minicap | 12 | 5.14 | -0.18 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Sebia | 5.112
Sebia Capillarys 2/ Minicap | 17 | 5.66 | -0.18 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Sebia | 5.112
Sebia Capillarys 2/ Minicap | 17 | 11.6 | -0.11 | 1.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Sebia | 3.418
Sebia Capillarys 2/ Minicap | 17 | 9.41 | -0.12 | 1.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Sebia | 2.82
Sebia Capillarys 2/ Minicap | 17 | 4.82 | -0.22 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Sebia | 5.268
Sebia Capillarys 2/ Minicap | 17 | 7.18 | -0.2 | 1.9 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Sebia | 3.524
Sebia Capillarys 2/ Minicap | 17-28 | 11.59 | -0.1 | 2.2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Sebia | 4.212
Sebia Capillarys 2/ Minicap | 17-28 | 5.83 | -0.1 | 1.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Sebia | 3.428
Sebia Capillarys 2/ Minicap | 17-28 | 5.03 | -0.14 | 1.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Sebia | 2.996
Sebia Capillarys 2/ Minicap | 17-28 | 8.08 | -0.06 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Sebia | 3.076
Sebia Capillarys 2/ Minicap | 17-28 | 9.14 | -0.16 | 1.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Sebia | 3.172
Sebia Capillarys 2/ Minicap | 19-31 | 5.16 | -0.16 | 2.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Sebia | 4.152
Sebia Capillarys 2/ Minicap | 19-31 | 9.17 | 0.01 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Sebia | 3.342
Sebia Capillarys 2/ Minicap | 19-31 | 5.15 | -0.16 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Sebia | 4.74
Sebia Capillarys 2/ Minicap | 19-31 | 11.94 | -0.09 | 1.9 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Sebia | 3.634
Sebia Capillarys 2/ Minicap | 19-31 | 5.83 | -0.11 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Sebia | 3.026
Sebia Capillarys 2/ Minicap | 25 | 5.14 | -0.13 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Sebia | 4.966
Sebia Capillarys 2/ Minicap | 25 | 10.52 | -0.07 | 1.3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Sebia | 2.478
Sebia Capillarys 2/ Minicap | 25 | 6 | -0.2 | 1.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Sebia | 3.132
Sebia Capillarys 2/ Minicap | 25 | 12.14 | -0.09 | 1.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Sebia | 3.046
Sebia Capillarys 2/ Minicap | 25 | 7.44 | -0.07 | 1.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Sebia | 2.478
Sebia Capillarys 2/ Minicap | 28-43 | 9.02 | -0.09 | 1.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Sebia | 2.85
Sebia Capillarys 2/ Minicap | 28-43 | 5.89 | -0.12 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Sebia | 4.388
Sebia Capillarys 2/ Minicap | 28-43 | 11.53 | -0.18 | 1.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Sebia | 3.348
Sebia Capillarys 2/ Minicap | 28-43 | 4.9 | -0.12 | 2 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Sebia | 3.8
Sebia Capillarys 2/ Minicap | 28-43 | 7.56 | -0.02 | 2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Sebia | 3.9
Sebia Capillarys 3 | 101 | 8.6282 | 0.7 | 2.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Sebia | 5.796
Sebia Capillarys 3 | 104 | 5.6636 | 0.2 | 2.9 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Sebia | 5.884
Sebia Capillarys 3 | 9 | 5.929 | -1 | 2.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Sebia | 4.096
Sebia Capillarys 3 | 9 | 7.2923 | -1.7 | 1.6 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Sebia | 1.436
Sebia Capillarys 3 | 8 | 5.9839 | -0.4 | 2.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Sebia | 4.696
Sebia Capillarys 3 | 8 | 7.4753 | 0.3 | 2 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Sebia | 4.22
Sebia Capillarys 3 | 8 | 6.3041 | 1.2 | 2.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Sebia | 5.316
Sebia Capillarys 3 | 8 | 7.4112 | -0.5 | 1.4 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Sebia | 2.244
Sebia Capillarys 3 | 13 | 6.2401 | 0.5 | 2 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Sebia | 4.42
Sebia Capillarys 3 | 13 | 7.457 | 0 | 1.5 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Sebia | 2.94
Sebia Capillarys 3 | 112 | 8.6282 | 0.7 | 2.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Sebia | 5.404
Sebia Capillarys 3 | 110 | 5.5904 | -0.6 | 3.3 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Sebia | 5.868
Sebia Capillarys 3 | 127 | 5.5263 | 0.4 | 3.3 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 6.868
Sebia Capillarys 3 | 122 | 5.5263 | 0.4 | 2.9 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 6.084
Sebia Capillarys 3 | 141 | 5.4806 | -0.1 | 3.3 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 6.368
Sebia Capillarys 3 | 142 | 5.4897 | 0 | 3.3 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 6.468
Sebia Capillarys 3 Tera | 14 | 7.93 | -0.04 | 2.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Sebia | 5.448
Sebia Capillarys 3 Tera | 14 | 5.94 | 0.05 | 1.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Sebia | 3.774
Sebia Capillarys 3 Tera | 14 | 5.1 | -0.03 | 1.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Sebia | 2.91
Sebia Capillarys 3 Tera | 14 | 7.25 | -0.15 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Sebia | 4.162
Sebia Capillarys 3 Tera | 14 | 9.16 | -0.01 | 1.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Sebia | 2.734
Sebia Capillarys 3 Tera | 16 | 7.74 | -0.07 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Sebia | 2.674
Sebia Capillarys 3 Tera | 16 | 5.81 | -0.1 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Sebia | 3.428
Sebia Capillarys 3 Tera | 16 | 8.59 | -0.1 | 0.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Sebia | 1.468
Sebia Capillarys 3 Tera | 16 | 5.78 | -0.08 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Sebia | 2.468
Sebia Capillarys 3 Tera | 16 | 9.39 | -0.17 | 1.3 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Sebia | 2.378
Sebia Capillarys 3 Tera/Octa | 22 | 8.82 | -0.04 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Sebia | 2.704
Sebia Capillarys 3 Tera/Octa | 22 | 6.72 | -0.12 | 1.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Sebia | 3.408
Sebia Capillarys 3 Tera/Octa | 22 | 9.28 | -0.11 | 1.6 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Sebia | 3.026
Sebia Capillarys 3 Tera/Octa | 22 | 5.95 | -0.18 | 1.7 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Sebia | 3.152
Sebia Capillarys 3 Tera/Octa | 22 | 7.39 | -0.13 | 1.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Sebia | 2.026
Sebia Capillarys 3 Tera/Octa | 22 | 5.39 | -0.09 | 1.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Sebia | 2.85
Sebia Capillarys 3 Tera/Octa | 22 | 8.2 | -0.06 | 1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Sebia | 1.9
Sebia Capillarys 3 Tera/Octa | 22 | 5.09 | -0.06 | 1.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Sebia | 3.076
Sebia Capillarys 3 Tera/Octa | 22 | 9.91 | -0.14 | 1.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Sebia | 2.604
Sebia Capillarys 3 Tera/Octa | 22 | 5.97 | -0.02 | 1.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Sebia | 2.528
Sebia Capillarys 3 Tera/Octa | 22 | 5.74 | -0.06 | 1.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Sebia | 3.076
Sebia Capillarys 3 Tera/Octa | 22 | 8.99 | -0.13 | 1.2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Sebia | 2.222
Sebia Capillarys 3 Tera/Octa | 22 | 5.43 | -0.04 | 1.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Sebia | 2.9
Sebia Capillarys 3 Tera/Octa | 22 | 7.06 | -0.19 | 1.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Sebia | 1.966
Sebia Capillarys 3 Tera/Octa | 22 | 9.17 | -0.09 | 1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Sebia | 1.87
Sebia Capillarys 3 Tera/Octa | 26 | 5.2 | -0.01 | 2.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Sebia | 4.498
Sebia Capillarys 3 Tera/Octa | 26 | 6.69 | 0.01 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Sebia | 3.146
Sebia Capillarys 3 Tera/Octa | 26 | 8.06 |  | 1.5 | GH5-13 |  | 2021 | CAP | Fresh | Sebia | 2.94
Sebia Capillarys 3 Tera/Octa | 26 | 6.01 | 0.03 | 1.3 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Sebia | 2.578
Sebia Capillarys 3 Tera/Octa | 26 | 10.1 | -0.02 | 1.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Sebia | 2.332
Sebia Capillarys 3 Tera | 36 | 7.4478 | -0.2 | 2 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Sebia | 3.72
Sebia Capillarys 3 Tera | 36 | 6.0479 | 0.4 | 2 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Sebia | 4.32
Sebia Capillarys 3 Tera | 26 | 7.4753 | 0.1 | 1.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Sebia | 3.824
Sebia Capillarys 3 Tera | 26 | 6.0296 | 0.2 | 2.1 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Sebia | 4.316
Sebia MINICAP | 23 | 8.5733 | 0.1 | 3.2 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Sebia | 6.372
Sebia MINICAP | 20 | 5.627 | -0.2 | 3.8 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Sebia | 7.248
Sebia MINICAP | 10 | 5.9198 | -1.1 | 2.8 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Sebia | 4.388
Sebia MINICAP | 11 | 7.3929 | -0.6 | 2.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Sebia | 3.712
Sebia MINICAP | 13 | 6.2126 | 0.2 | 3.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Sebia | 6.276
Sebia MINICAP | 13 | 7.4387 | -0.2 | 2 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Sebia | 3.72
Sebia MINICAP | 23 | 8.555 | -0.1 | 2.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Sebia | 5.388
Sebia MINICAP | 23 | 5.5629 | -0.9 | 3.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Sebia | 5.176
Sebia MINICAP | 25 | 5.4623 | -0.3 | 6.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 12.44
Sebia MINICAP | 23 | 5.5446 | 0.6 | 4.9 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 10.204
Sebia MINICAP | 25 | 5.4348 | -0.6 | 4.9 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 9.004
Sebia MINICAP | 24 | 5.4531 | -0.4 | 4.8 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 9.008
Sebia MINICAP flex-piercing | 13 | 7.3563 | -1.2 | 2.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Sebia | 4.484
Sebia MINICAP flex-piercing | 13 | 5.9656 | -0.5 | 3.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Sebia | 6.164
Sebia MINICAP flex-piercing | 17 | 7.4295 | -0.4 | 4.1 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Sebia | 7.636
Sebia MINICAP flex-piercing | 17 | 5.9564 | -0.6 | 3 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Sebia | 5.28
Sebia not specified/other | 10 | 8.6648 | 1.1 | 2.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Sebia | 6.784
Sebia not specified/other | 10 | 5.627 | -0.2 | 4.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Sebia | 7.836
Sebia not specified/other | 7 | 5.4989 | 0.1 | 2.7 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 5.392
Sebia not specified/other | 7 | 5.5629 | 0.8 | 3.7 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Sebia | 8.052
Sebia not specified/other | 60 | 5.4897 | 0 | 3.4 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 6.664
Sebia not specified/other | 60 | 5.4989 | 0.1 | 3.6 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Sebia | 7.156
Siemens Advia | 47 | 5.86 | -0.04 | 5.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 11.132
Siemens Advia | 47 | 9.73 | -0.07 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 8.358
Siemens Advia | 47 | 7.49 | 0.09 | 4.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 8.91
Siemens Advia | 52 | 5.3 | 0.1 | 4.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 8.332
Siemens Advia | 52 | 8.74 | 0.04 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 4.94
Siemens Advia | 52 | 6.43 | 0.13 | 4.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 8.166
Siemens Advia | 48 | 8.51 | 0.01 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 6.478
Siemens Advia | 48 | 5.49 | 0.09 | 3.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 7.538
Siemens Advia | 48 | 6.54 | 0.14 | 3.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 7.784
Siemens Advia | 32 | 6.52 | -0.06 | 4.3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 8.368
Siemens Advia | 32 | 8.33 | -0.06 | 4.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 8.172
Siemens Advia | 32 | 5.5 | -0.15 | 6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 11.61
Siemens Advia | 40 | 6.78 | 0.29 | 3.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 7.738
Siemens Advia | 40 | 7.27 | 0.3 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 6.376
Siemens Advia | 40 | 9.72 | 0.07 | 3.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 7.322
Siemens Advia | 29 | 6.93 | 0.14 | 4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 7.98
Siemens Advia | 29 | 10.09 | -0.19 | 4.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 9.414
Siemens Advia | 29 | 6.88 | 0.06 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 5.744
Siemens Advia | 25 | 8.64 | 0.01 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.694
Siemens Advia | 25 | 5.36 | 0.04 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 7.292
Siemens Advia | 24 | 6.02 | 0.18 | 5.1 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 10.176
Siemens Advia | 24 | 11.72 | 0.01 | 4.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 8.83
Siemens Advia | 24 | 9.71 | 0.18 | 5.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 10.372
Siemens Advia | 24 | 5.14 | 0.1 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 5.392
Siemens Advia | 24 | 7.74 | 0.36 | 5.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 11.14
Siemens Advia | 22-24 | 11.53 | -0.16 | 3.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 6.504
Siemens Advia | 22-24 | 5.98 | 0.05 | 4.1 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 8.086
Siemens Advia | 22-24 | 5.16 | -0.01 | 4.1 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 8.026
Siemens Advia | 22-24 | 8.17 | 0.03 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 6.89
Siemens Advia | 22-24 | 8.78 | -0.52 | 3.9 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 7.124
Siemens Advia | 25-27 | 5.42 | 0.1 | 4.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 8.332
Siemens Advia | 25-27 | 9.31 | 0.15 | 3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 6.03
Siemens Advia | 25-27 | 5.39 | 0.08 | 4.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 8.116
Siemens Advia | 25-27 | 11.79 | -0.24 | 3.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 6.228
Siemens Advia | 25-27 | 6.03 | 0.09 | 3.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 7.342
Siemens Advia | 25 | 5.27 | 0 | 2.8 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 5.488
Siemens Advia | 25 | 10.5 | -0.09 | 3.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 6.182
Siemens Advia | 25 | 6.12 | -0.08 | 3.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 7.172
Siemens Advia | 25 | 11.88 | -0.35 | 3.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 6.314
Siemens Advia | 25 | 7.56 | 0.05 | 3.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 6.91
Siemens Advia | 21-24 | 9.08 | -0.03 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 4.478
Siemens Advia | 21-24 | 6.09 | 0.08 | 2.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 4.98
Siemens Advia | 21-24 | 11.42 | -0.29 | 3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 5.59
Siemens Advia | 21-24 | 4.98 | -0.04 | 3.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 7.408
Siemens Advia | 21-24 | 7.6 | 0.02 | 2.6 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 5.116
Siemens Advia | 18-19 | 6.57 | 0.16 | 2.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 5.452
Siemens Advia | 18-19 | 9.49 | -0.04 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 5.448
Siemens Advia | 18-19 | 5.46 | 0.12 | 2.3 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 4.628
Siemens Advia | 18-19 | 8.63 | 0.12 | 2.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 4.236
Siemens Advia | 18-19 | 7.65 | 0.4 | 3.2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 6.672
Siemens Advia | 15 | 7.19 | 0.04 | 2.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 5.528
Siemens Advia | 15 | 5.31 | 0.12 | 3.9 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 7.764
Siemens Advia | 15 | 8.32 | -0.1 | 3.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 5.976
Siemens Advia | 15 | 9.6 | -0.19 | 3.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 6.278
Siemens Advia | 15 | 6.16 | 0.04 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 4.744
Siemens Advia | 12 | 6.43 | 0.12 | 2.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 4.236
Siemens Advia | 12 | 9.17 | 0.06 | 3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 5.94
Siemens Advia | 12 | 5.35 | 0.04 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 5.528
Siemens Advia | 12 | 8.29 | -0.06 | 3.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 6.016
Siemens Advia | 12 | 7.77 | 0.15 | 3.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 7.402
Siemens Advia | 16 | 7.38 | -0.04 | 1.8 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 3.488
Siemens Advia | 16 | 5.16 | -0.04 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 5.056
Siemens Advia | 16 | 8.24 | -0.07 | 2.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 5.614
Siemens Advia | 16 | 9.95 | -0.38 | 2.6 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 4.716
Siemens Advia | 16 | 5.89 | 0.02 | 2.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 4.92
Siemens Advia | 16 | 7.29 | 0.04 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 5.528
Siemens Advia | 16 | 6.79 | -0.02 | 2.8 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 5.468
Siemens Advia | 16 | 8.59 | -0.06 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 4.448
Siemens Advia | 16 | 9.31 | -0.19 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 4.318
Siemens Advia | 16 | 5.45 | 0 | 2.8 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 5.488
Siemens Advia (immunoassay) | 7 | 8.7837 | 2.4 | 6.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Siemens | 15.532
Siemens Advia (immunoassay) | 7 | 5.7368 | 1 | 5.2 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Siemens | 11.192
Siemens Advia 1650 | 16 | 5.1 |  | 4.1 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 8.036
Siemens Advia 1650 | 16 | 8.3 |  | 3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 5.88
Siemens Advia 1650 | 16 | 6.1 |  | 2.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 5.292
Siemens Advia 1650/2400 | 34 | 13.6 |  | 7.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Siemens | 13.916
Siemens Advia 1650/2400 | 34 | 5.6 |  | 4.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Siemens | 8.232
Siemens Advia 1650/2400 | 34 | 9 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Siemens | 5.88
Siemens Advia 1650/2400 | 42 | 12 |  | 5.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 9.996
Siemens Advia 1650/2400 | 42 | 7 |  | 3.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 7.056
Siemens Advia 1650/2400 | 42 | 5.3 |  | 4.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 8.232
Siemens Advia 1800 | 14 | 5.1 |  | 5.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 10.584
Siemens Advia 1800 | 14 | 8.4 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 8.036
Siemens Advia 1800 | 14 | 6.1 |  | 5.1 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 9.996
Siemens Advia A1c_3 Reagent | 59 | 5.56 | 0.16 | 5.7 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 11.332
Siemens Advia A1c_3 Reagent | 59 | 8.54 | 0.24 | 5.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 11.216
Siemens Advia A1c_3 Reagent | 59 | 5.9 | 0.25 | 6.9 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 13.774
Siemens Advia New Reagent | 21 | 6.41 | -0.19 | 7.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 14.706
Siemens Advia New Reagent | 21 | 9 | -0.5 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 6.164
Siemens Advia New Reagent | 21 | 6.98 | -0.42 | 9.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 18.592
Siemens Advia Original Reagent | 26 | 6.79 | 0.19 | 3.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 7.638
Siemens Advia Original Reagent | 26 | 9.31 | -0.19 | 4.1 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 7.846
Siemens Advia Original Reagent | 26 | 7.54 | 0.14 | 4.1 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 8.176
Siemens Advia Chemistry Systems | 43 | 7.19 | 0.08 | 4.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 9.292
Siemens Advia Chemistry Systems | 43 | 9.28 | -0.04 | 3.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 7.408
Siemens Advia Chemistry Systems | 43 | 6.33 | 0.26 | 4.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 8.884
Siemens Advia Chemistry Systems | 44 | 8.22 | 0.17 | 4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 8.01
Siemens Advia Chemistry Systems | 44 | 5.5 | 0.2 | 3.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 7.256
Siemens Advia Chemistry Systems | 44 | 6.42 | 0.28 | 3.8 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 7.728
Siemens Advia Chemistry Systems | 13 | 9.25 | -0.12 | 3.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 7.132
Siemens Advia Chemistry Systems | 13 | 5.02 | 0.1 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 6.568
Siemens Advia Chemistry Systems | 13 | 6.22 | 0.18 | 3.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 7.236
Siemens Advia Chemistry Systems | 13 | 5.4 | 0.12 | 2.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 4.432
Siemens Advia Chemistry Systems | 13 | 7.93 | -0.15 | 3.8 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 7.298
Siemens Advia New Reagent | 21 | 5.63 | -0.27 | 4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 7.57
Siemens Advia New Reagent | 21 | 9.61 | -0.19 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 7.65
Siemens Advia New Reagent | 21 | 7.31 | -0.09 | 4.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 7.946
Siemens Advia New Reagent | 24 | 5.22 | 0.02 | 4.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 8.252
Siemens Advia New Reagent | 24 | 8.73 | 0.03 | 2.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 4.734
Siemens Advia New Reagent | 24 | 6.33 | 0.03 | 3.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 7.674
Siemens Advia New Reagent | 25 | 8.48 | -0.02 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 6.448
Siemens Advia New Reagent | 25 | 5.42 | 0.02 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 7.664
Siemens Advia New Reagent | 25 | 6.47 | 0.07 | 4.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 8.302
Siemens Advia New Reagent | 10 | 6.32 | 0.02 | 4.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 9.036
Siemens Advia New Reagent | 10 | 7.49 | -0.11 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 7.73
Siemens Advia New Reagent | 10 | 9.13 | -0.07 | 5.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 10.514
Siemens Advia New Reagent | 41 | 5.35 | -0.25 | 5.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 9.746
Siemens Advia New Reagent | 41 | 9.01 | -0.39 | 6.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 12.742
Siemens Advia New Reagent | 41 | 7.02 | -0.18 | 5.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 10.992
Siemens Advia | 8 | 8.8661 | 3.3 | 6.1 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Siemens | 15.256
Siemens Advia | 8 | 5.7459 | 1.1 | 11 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Siemens | 22.66
Siemens Advia | 11 | 5.5721 | 0.9 | 5.6 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 11.876
Siemens Advia | 11 | 5.5629 | 0.8 | 6 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 12.56
Siemens Advia Original Reagent | 13 | 6.03 | 0.13 | 2.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 5.422
Siemens Advia Original Reagent | 13 | 9.65 | -0.15 | 4.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 7.886
Siemens Advia Original Reagent | 13 | 7.51 | 0.11 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 6.97
Siemens Advia Original Reagent | 13 | 5.42 | 0.22 | 3.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 6.688
Siemens Advia Original Reagent | 13 | 8.83 | 0.13 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 5.03
Siemens Advia Original Reagent | 13 | 6.55 | 0.25 | 2.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 5.542
Siemens Advia Original Reagent | 21 | 8.54 | 0.04 | 2.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.528
Siemens Advia Original Reagent | 21 | 5.57 | 0.17 | 3.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 6.246
Siemens Advia Original Reagent | 21 | 6.62 | 0.22 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 6.688
Siemens Advia Original Reagent | 38 | 6.29 | -0.01 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 7.438
Siemens Advia Original Reagent | 38 | 7.54 | -0.06 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 7.78
Siemens Advia Original Reagent | 38 | 9 | -0.2 | 4.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 9.012
Siemens Advia Original Reagent | 20 | 5.8 | 0.2 | 3.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 6.668
Siemens Advia Original Reagent | 20 | 9.69 | 0.29 | 5.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 10.286
Siemens Advia Original Reagent | 20 | 7.53 | 0.33 | 4.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 8.562
Siemens Advia | 15 | 6.2767 | 2.8 | 5.6 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Siemens | 13.776
Siemens Advia | 15 | 7.8322 | 4.2 | 3.9 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Siemens | 11.844
Siemens Advia | 13 | 6.4413 | 2.7 | 6.8 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Siemens | 16.028
Siemens Advia | 14 | 7.7315 | 3 | 8.6 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Siemens | 19.856
Siemens Advia | 21 | 7.6674 | 2.2 | 7.3 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Siemens | 16.508
Siemens Advia | 20 | 6.1669 | 1.7 | 7.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Siemens | 16.988
Siemens Atellica CH- Enzymatic | 25 | 7.86 | -0.11 | 2.4 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 4.594
Siemens Atellica CH- Enzymatic | 25 | 5.68 | 0.21 | 3.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 6.482
Siemens Atellica CH- Enzymatic | 25 | 4.92 | -0.21 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 5.474
Siemens Atellica CH- Enzymatic | 25 | 7.15 | -0.25 | 1.6 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 2.886
Siemens Atellica CH- Enzymatic | 25 | 8.96 | -0.21 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 4.298
Siemens Atellica CH- Enzymatic | 34 | 7.63 | -0.18 | 3.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 5.896
Siemens Atellica CH- Enzymatic | 34 | 5.65 | -0.26 | 2.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 5.424
Siemens Atellica CH- Enzymatic | 34 | 8.48 | -0.21 | 2.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 5.474
Siemens Atellica CH- Enzymatic | 34 | 5.64 | -0.22 | 2.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 4.092
Siemens Atellica CH- Enzymatic | 34 | 9.36 | -0.2 | 3.2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 6.072
Siemens Atellica CH- Enzymatic | 75 | 8.82 | -0.04 | 1.9 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 3.684
Siemens Atellica CH- Enzymatic | 75 | 6.67 | -0.17 | 2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 3.75
Siemens Atellica CH- Enzymatic | 75 | 9.36 | -0.03 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 3.89
Siemens Atellica CH- Enzymatic | 75 | 5.91 | -0.22 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 3.308
Siemens Atellica CH- Enzymatic | 75 | 7.41 | -0.11 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 3.418
Siemens Atellica CH- Enzymatic | 58 | 5.45 | -0.03 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 5.458
Siemens Atellica CH- Enzymatic | 58 | 8.34 | 0.08 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 3.804
Siemens Atellica CH- Enzymatic | 58 | 5.11 | -0.04 | 3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 5.84
Siemens Atellica CH- Enzymatic | 58 | 10.23 | 0.18 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 3.512
Siemens Atellica CH- Enzymatic | 58 | 5.97 | -0.02 | 2.7 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 5.272
Siemens Atellica CH- Enzymatic | 71 | 5.57 | -0.23 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 3.102
Siemens Atellica CH- Enzymatic | 71 | 9.04 | -0.08 | 1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 1.88
Siemens Atellica CH- Enzymatic | 71 | 5.28 | -0.19 | 2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 3.73
Siemens Atellica CH- Enzymatic | 71 | 7.15 | -0.1 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 2.448
Siemens Atellica CH- Enzymatic | 71 | 9.18 | -0.08 | 1.3 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 2.468
Siemens Atellica CH- Enzymatic | 83 | 5.07 | -0.14 | 2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 3.78
Siemens Atellica CH- Enzymatic | 83 | 6.55 | -0.13 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 2.614
Siemens Atellica CH- Enzymatic | 83 | 8.21 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 3.528
Siemens Atellica CH- Enzymatic | 83 | 5.9 | -0.08 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 3.448
Siemens Atellica CH- Enzymatic | 83 | 10.27 | 0.15 | 1.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 2.502
Siemens Atellica CH- Immunoassay | 17 | 8.1 | 0.13 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 4.442
Siemens Atellica CH- Immunoassay | 17 | 6 | 0.11 | 2.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 5.794
Siemens Atellica CH- Immunoassay | 17 | 5.13 | 0 | 3.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 6.272
Siemens Atellica CH- Immunoassay | 17 | 7.59 | 0.19 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 6.07
Siemens Atellica CH- Immunoassay | 17 | 9.23 | 0.06 | 2.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 4.764
Siemens Atellica CH- Immunoassay | 11 | 7.86 | 0.05 | 3.6 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 7.106
Siemens Atellica CH- Immunoassay | 11 | 5.85 | -0.06 | 3.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 6.212
Siemens Atellica CH- Immunoassay | 11 | 8.71 | 0.02 | 2.6 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 5.116
Siemens Atellica CH- Immunoassay | 11 | 5.92 | 0.06 | 4.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 8.292
Siemens Atellica CH- Immunoassay | 11 | 9.65 | 0.09 | 3.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 6.754
Siemens Atellica CH not specified/other | 19 | 5.4989 | 0.1 | 5.6 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 11.076
Siemens Atellica CH not specified/other | 19 | 5.4897 | 0 | 4.3 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 8.428
Siemens DCA 2000 | 187 | 13.7 |  | 4 | GH2-01 | 13 | 2008 | CAP | Fresh | Siemens | 7.84
Siemens DCA 2000 | 187 | 5.2 |  | 2.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Siemens | 5.488
Siemens DCA 2000 | 187 | 8.5 |  | 2.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Siemens | 5.684
Siemens DCA 2000 | 178 | 12 |  | 4.9 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 9.604
Siemens DCA 2000 | 178 | 6.8 |  | 2.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 5.096
Siemens DCA 2000 | 178 | 5.1 |  | 2.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 5.096
Siemens DCA 2000/200+ | 134 | 5.1 |  | 2.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 5.684
Siemens DCA 2000/200+ | 134 | 8.4 |  | 3.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 6.468
Siemens DCA 2000/200+ | 134 | 6 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 6.468
Siemens DCA 2000/200+ | 106 | 6.81 | 0.21 | 2.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 5.11
Siemens DCA 2000/200+ | 106 | 9.48 | -0.02 | 3 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 5.86
Siemens DCA 2000/200+ | 106 | 7.56 | 0.16 | 3.1 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 6.236
Siemens DCA 2000/200+ | 84 | 6.07 | 0.17 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 5.658
Siemens DCA 2000/200+ | 84 | 9.66 | -0.14 | 3.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 6.132
Siemens DCA 2000/200+ | 84 | 7.5 | 0.1 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 4.804
Siemens DCA 2000/200+ | 75 | 5.29 | 0.09 | 2.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 5.186
Siemens DCA 2000/200+ | 75 | 8.71 | 0.01 | 3.2 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 6.282
Siemens DCA 2000/200+ | 75 | 6.45 | 0.15 | 2.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 5.246
Siemens DCA 2000/2000+ | 66 | 8.5 | 0 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 6.076
Siemens DCA 2000/2000+ | 66 | 5.47 | 0.07 | 2.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.754
Siemens DCA 2000/2000+ | 66 | 6.54 | 0.14 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.432
Siemens DCA 2000/2000+ | 19 | 6.58 | 0.28 | 1.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.808
Siemens DCA 2000/2000+ | 19 | 7.95 | 0.35 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 5.446
Siemens DCA 2000/2000+ | 19 | 9.52 | 0.32 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.024
Siemens DCA 2000/2000+ | 71 | 5.61 | 0.01 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 5.302
Siemens DCA 2000/2000+ | 71 | 9.25 | -0.15 | 3.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 6.122
Siemens DCA 2000/2000+ | 71 | 7.19 | -0.01 | 3.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 6.458
Siemens DCA 2000/2000+ | 58 | 5.46 | 0.06 | 2.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 5.548
Siemens DCA 2000/2000+ | 58 | 8.28 | -0.02 | 3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 5.86
Siemens DCA 2000/2000+ | 58 | 5.77 | 0.12 | 3.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 6.196
Siemens DCA 2000/2000+ | 39 | 7.08 | -0.03 | 3.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 7.026
Siemens DCA 2000/2000+ | 39 | 9.41 | 0.09 | 3.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 6.166
Siemens DCA 2000/2000+ | 39 | 6.08 | 0.01 | 3.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 6.086
Siemens DCA 2000/2000+ | 37 | 7.92 | -0.13 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 5.162
Siemens DCA 2000/2000+ | 37 | 5.24 | -0.06 | 2.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 4.84
Siemens DCA 2000/2000+ | 37 | 6.09 | -0.05 | 2.3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 4.458
Siemens DCA 2000/2000+ | 29 | 6.51 | -0.07 | 3.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 6.202
Siemens DCA 2000/2000+ | 29 | 8.21 | -0.18 | 3.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 6.288
Siemens DCA 2000/2000+ | 29 | 5.52 | -0.13 | 3.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 6.534
Siemens DCA 2000/2000+ | 30 | 6.6 | 0.11 | 3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 5.99
Siemens DCA 2000/2000+ | 30 | 7.08 | 0.11 | 2.7 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 5.402
Siemens DCA 2000/2000+ | 30 | 9.68 | 0.03 | 3.5 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 6.89
Siemens DCA 2000/2000+ | 26 | 6.67 | -0.12 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 5.172
Siemens DCA 2000/2000+ | 26 | 10.21 | -0.07 | 3.2 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 6.202
Siemens DCA 2000/2000+ | 26 | 6.67 | -0.15 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 5.534
Siemens DCA 2000/2000+ | <10 |  |  |  | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | <10 |  |  |  | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 6-20 | 12.25 | 0.56 | 6.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 14.084
Siemens DCA 2000/2000+ | 6-20 | 5.83 | -0.1 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 4.8
Siemens DCA 2000/2000+ | 6-20 | 5.16 | -0.01 | 2.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 5.282
Siemens DCA 2000/2000+ | 6-20 |  |  |  | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 6-20 |  |  |  | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 8-26 | 5.33 | 0.01 | 3.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 6.478
Siemens DCA 2000/2000+ | 8-26 | 9.2 | 0.04 | 2.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 5.724
Siemens DCA 2000/2000+ | 8-26 | 5.29 | -0.02 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 4.292
Siemens DCA 2000/2000+ | 8-26 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 8-26 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 5-21 | 9.22 | 0.11 | 3.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 6.774
Siemens DCA 2000/2000+ | 5-21 | 6.12 | 0.11 | 4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 7.95
Siemens DCA 2000/2000+ | 5-21 | 12.17 | 0.46 | 3.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 7.516
Siemens DCA 2000/2000+ | 5-21 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 5-21 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 0-14 | 6.54 | 0.13 | 3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 6.01
Siemens DCA 2000/2000+ | 0-14 | 9.59 | 0.06 | 3.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 7.312
Siemens DCA 2000/2000+ | 0-14 | 5.4 | 0.06 | 2.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 4.764
Siemens DCA 2000/2000+ | 0-14 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 0-14 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 10 | 7.43 | 0.18 | 3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 6.06
Siemens DCA 2000/2000+ | 10 | 7.08 | 0.27 | 3.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 6.542
Siemens DCA 2000/2000+ | 10 | 8.96 | 0.31 | 3.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 6.386
Siemens DCA 2000/2000+ | 10 |  |  |  | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 10 |  |  |  | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 11 | 7.17 | 0.02 | 4.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 8.252
Siemens DCA 2000/2000+ | 11 | 5.27 | 0.08 | 4.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 8.9
Siemens DCA 2000/2000+ | 11 | 8.47 | 0.05 | 2.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 5.734
Siemens DCA 2000/2000+ | 11 |  |  |  | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 11 |  |  |  | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/Vantage | 61 | 9.0948 | 5.8 | 4.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Siemens | 14.424
Siemens DCA 2000/Vantage | 58 | 6.0479 | 4.4 | 5.2 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Siemens | 14.592
Siemens DCA 2000/Vantage | 9 | 7.6949 | 2.5 | 3.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Siemens | 9.36
Siemens DCA 2000/Vantage | 8 | 6.2126 | 2.2 | 4.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Siemens | 10.824
Siemens DCA 2000/Vantage | 322 | 7.3838 | -0.9 | 3.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Siemens | 5.96
Siemens DCA 2000/Vantage | 321 | 5.9747 | -0.4 | 3.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Siemens | 7.048
Siemens DCA 2000/Vantage | 262 | 8.4452 | -1.3 | 3.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Siemens | 6.148
Siemens DCA 2000/Vantage | 263 | 5.4714 | -1.9 | 3.8 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Siemens | 5.548
Siemens DCA 2000/Vantage | 38 | 6.0296 | 5.9 | 5.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Siemens | 16.68
Siemens DCA 2000/Vantage | 27 | 6.0205 | 5.8 | 4.5 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Siemens | 14.62
Siemens DCA 2000/Vantage | 208 | 5.508 | 0.2 | 3.8 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 7.648
Siemens DCA 2000/Vantage | 207 | 5.5263 | 0.4 | 3.5 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 7.26
Siemens DCA 2000/Vantage | 6 | 6.3956 | 4.1 | 3.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Siemens | 11.744
Siemens DCA 2000/Vantage | 5 | 7.8139 | 4 | 3.3 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Siemens | 10.468
Siemens DCA 2000/Vantage | 158 | 6.0662 | 0.5 | 3.4 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Siemens | 7.164
Siemens DCA 2000/Vantage | 160 | 7.5119 | 0.7 | 3.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Siemens | 7.952
Siemens DCA 2000/Vantage | 9 | 6.6975 | 5.5 | 5.1 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Siemens | 15.496
Siemens DCA 2000/Vantage | 9 | 7.9054 | 4.9 | 7.3 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Siemens | 19.208
Siemens DCA 2000/Vantage | 217 | 6.295 | 1.1 | 4.3 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Siemens | 9.528
Siemens DCA 2000/Vantage | 217 | 7.6034 | 1.6 | 4.2 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Siemens | 9.832
Siemens DCA Vantage | 59 | 5.1 |  | 2.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 4.508
Siemens DCA Vantage | 59 | 8.4 |  | 2.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 5.096
Siemens DCA Vantage | 59 | 6 |  | 2.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 5.684
Siemens DCA Vantage | 97 | 6.78 | 0.18 | 2.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 5.668
Siemens DCA Vantage | 97 | 9.39 | -0.11 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 6.554
Siemens DCA Vantage | 97 | 7.55 | 0.15 | 2.5 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 5.05
Siemens DCA Vantage | 118 | 6.07 | 0.17 | 2.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 4.482
Siemens DCA Vantage | 118 | 9.62 | -0.18 | 3.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 6.092
Siemens DCA Vantage | 118 | 7.44 | 0.04 | 2.6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 5.136
Siemens DCA Vantage | 115 | 5.33 | 0.13 | 2.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 5.422
Siemens DCA Vantage | 115 | 8.63 | -0.07 | 2.8 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 5.418
Siemens DCA Vantage | 115 | 6.4 | 0.1 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 4.412
Siemens DCA Vantage | 162 | 8.39 | -0.11 | 2.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.182
Siemens DCA Vantage | 162 | 5.48 | 0.08 | 2.4 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 4.784
Siemens DCA Vantage | 162 | 6.52 | 0.12 | 2.5 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.02
Siemens DCA Vantage | 59 | 6.26 | -0.04 | 3.1 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 6.036
Siemens DCA Vantage | 59 | 7.66 | 0.06 | 3.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 6.92
Siemens DCA Vantage | 59 | 9.17 | -0.03 | 2.9 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.654
Siemens DCA Vantage | 253 | 5.56 | -0.04 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 5.252
Siemens DCA Vantage | 253 | 9.18 | -0.22 | 3.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 5.856
Siemens DCA Vantage | 253 | 7.14 | -0.06 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 5.232
Siemens DCA Vantage | 261 | 5.44 | 0.04 | 2.6 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 5.136
Siemens DCA Vantage | 261 | 8.17 | -0.13 | 2.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 4.966
Siemens DCA Vantage | 261 | 5.78 | 0.13 | 2.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 5.618
Siemens DCA Vantage | 235 | 6.99 | -0.12 | 2.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 5.368
Siemens DCA Vantage | 235 | 9.26 | -0.06 | 3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 5.82
Siemens DCA Vantage | 235 | 6.05 | -0.02 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 5.076
Siemens DCA Vantage | 258 | 7.84 | -0.21 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 5.082
Siemens DCA Vantage | 258 | 5.21 | -0.09 | 2.3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 4.418
Siemens DCA Vantage | 258 | 6.03 | -0.11 | 2.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 4.79
Siemens DCA Vantage | 307 | 6.44 | -0.14 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 5.348
Siemens DCA Vantage | 307 | 8.14 | -0.25 | 3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 5.63
Siemens DCA Vantage | 307 | 5.45 | -0.2 | 3.2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 6.072
Siemens DCA Vantage | 305 | 6.46 | -0.03 | 2.6 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 5.066
Siemens DCA Vantage | 305 | 6.95 | -0.02 | 2.8 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 5.468
Siemens DCA Vantage | 305 | 9.61 | -0.04 | 3.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 6.232
Siemens DCA Vantage | 340 | 6.6 | -0.19 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 5.102
Siemens DCA Vantage | 340 | 9.9 | -0.38 | 3.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 6.088
Siemens DCA Vantage | 340 | 6.61 | -0.21 | 2.7 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 5.082
Siemens DCA Vantage | 145 | 8.31 | -0.32 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.168
Siemens DCA Vantage | 145 | 5.22 | -0.1 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 5.192
Siemens DCA Vantage | 168 | 5.74 | -0.1 | 2.3 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 4.408
Siemens DCA Vantage | 168 | 11.91 | 0.2 | 5.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 10.392
Siemens DCA Vantage | 168 | 9.2 | -0.33 | 3.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 5.942
Siemens DCA Vantage | 168 | 4.97 | -0.07 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 5.222
Siemens DCA Vantage | 168 | 7.13 | -0.25 | 2.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 4.846
Siemens DCA Vantage | 171-360 | 11.94 | 0.25 | 6.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 12.206
Siemens DCA Vantage | 171-360 | 5.77 | -0.16 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 4.74
Siemens DCA Vantage | 171-360 | 5.1 | -0.07 | 2.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 5.222
Siemens DCA Vantage | 171-360 | 7.89 | -0.25 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 4.846
Siemens DCA Vantage | 171-360 | 8.99 | -0.31 | 3.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 6.55
Siemens DCA Vantage | 189-473 | 5.25 | -0.07 | 2.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 4.242
Siemens DCA Vantage | 189-473 | 9.07 | -0.09 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 5.398
Siemens DCA Vantage | 189-473 | 5.27 | -0.04 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 4.86
Siemens DCA Vantage | 189-473 | 12.26 | 0.23 | 4.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 9.442
Siemens DCA Vantage | 189-473 | 5.86 | -0.08 | 2 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 3.84
Siemens DCA Vantage | 186 | 5.2 | -0.07 | 2.2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 4.242
Siemens DCA Vantage | 186 | 10.44 | -0.15 | 4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 7.69
Siemens DCA Vantage | 186 | 6.06 | -0.14 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 4.172
Siemens DCA Vantage | 186 | 12.57 | 0.34 | 5.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 11.316
Siemens DCA Vantage | 186 | 7.3 | -0.21 | 2.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 4.69
Siemens DCA Vantage | 176-460 | 9.05 | -0.06 | 2.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 176-460 | 5.96 | -0.05 | 2.2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 4.262
Siemens DCA Vantage | 176-460 | 11.92 | 0.21 | 4.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 8.638
Siemens DCA Vantage | 176-460 | 5 | -0.02 | 2.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 4.88
Siemens DCA Vantage | 176-460 | 7.55 | -0.03 | 2.6 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 5.066
Siemens DCA Vantage | 150-483 | 6.45 | 0.04 | 2.3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 4.548
Siemens DCA Vantage | 150-483 | 9.46 | -0.07 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 5.418
Siemens DCA Vantage | 150-483 | 5.37 | 0.03 | 2.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 4.93
Siemens DCA Vantage | 150-483 | 8.49 | -0.02 | 2.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 4.684
Siemens DCA Vantage | 150-483 | 7.31 | 0.06 | 2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 3.98
Siemens DCA Vantage | 148 | 7.6 | 0.18 | 2.8 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 5.668
Siemens DCA Vantage | 148 | 5.23 | 0.03 | 2.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 5.322
Siemens DCA Vantage | 148 | 8.48 | 0.17 | 2.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 5.462
Siemens DCA Vantage | 148 | 10.61 | 0.28 | 3.6 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 7.336
Siemens DCA Vantage | 148 | 6.02 | 0.15 | 2.6 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 5.246
Siemens DCA Vantage | 435 | 7.31 | 0.06 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 5.548
Siemens DCA Vantage | 435 | 6.93 | 0.12 | 2.8 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 5.608
Siemens DCA Vantage | 435 | 8.73 | 0.08 | 2.9 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 5.764
Siemens DCA Vantage | 435 | 9.74 | 0.24 | 3.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 7.1
Siemens DCA Vantage | 435 | 5.48 | 0.03 | 2.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 4.538
Siemens DCA Vantage | 429 | 7.17 | 0.02 | 2.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 4.724
Siemens DCA Vantage | 429 | 5.29 | 0.1 | 2.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 5.392
Siemens DCA Vantage | 429 | 8.43 | 0.01 | 2.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 5.498
Siemens DCA Vantage | 429 | 9.87 | 0.08 | 3.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 6.548
Siemens DCA Vantage | 429 | 6.16 | 0.04 | 2.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 5.136
Siemens DCA Vantage | 129 | 6.37 | 0.06 | 2.8 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 5.548
Siemens DCA Vantage | 129 | 9.24 | 0.13 | 3.4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 6.794
Siemens DCA Vantage | 129 | 5.37 | 0.06 | 3.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 6.332
Siemens DCA Vantage | 129 | 8.48 | 0.13 | 2.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 5.422
Siemens DCA Vantage | 129 | 7.73 | 0.11 | 2.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 5.402
Siemens DCA Vantage | 410 | 9.29 | -0.08 | 3.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 5.996
Siemens DCA Vantage | 410 | 5.03 | 0.11 | 2.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 5.01
Siemens DCA Vantage | 410 | 6.11 | 0.07 | 2.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 4.97
Siemens DCA Vantage | 410 | 5.29 | 0.01 | 2.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 4.322
Siemens DCA Vantage | 410 | 8.08 | 0 | 2.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 4.508
Siemens DCA Vantage | 392 | 5.46 | 0 | 3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 5.88
Siemens DCA Vantage | 392 | 5.98 | 0.32 | 2.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 6.004
Siemens DCA Vantage | 392 | 9.36 | 0.05 | 3.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 6.518
Siemens DCA Vantage | 392 | 5.22 | -0.06 | 2.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 392 | 7.4 | -0.01 | 2.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 4.89
Siemens DCA Vantage | 132 | 6.37 | -0.04 | 2.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 4.664
Siemens DCA Vantage | 132 | 8.52 | -0.08 | 3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 5.8
Siemens DCA Vantage | 132 | 5.37 | -0.05 | 2.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 5.046
Siemens DCA Vantage | 132 | 7.41 | 0.03 | 2.6 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 5.126
Siemens DCA Vantage | 132 | 9.74 | -0.01 | 3.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 6.262
Siemens DCA Vantage | 395 | 5.16 | -0.09 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 4.614
Siemens DCA Vantage | 395 | 6.38 | -0.03 | 2.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 4.674
Siemens DCA Vantage | 395 | 8.23 | 0.02 | 2.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 4.724
Siemens DCA Vantage | 395 | 5.43 | -0.04 | 2.6 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 5.056
Siemens DCA Vantage | 395 | 9.48 | 0 | 3.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 6.468
Siemens DCA Vantage | 305 | 8.13 | 0.16 | 3.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 6.236
Siemens DCA Vantage | 305 | 5.91 | 0.02 | 2.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 4.92
Siemens DCA Vantage | 305 | 5.11 | -0.02 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 4.684
Siemens DCA Vantage | 305 | 7.5 | 0.1 | 2.4 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 4.804
Siemens DCA Vantage | 305 | 9.07 | -0.1 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 5.192
Siemens DCA Vantage | 72 | 7.74 | -0.07 | 2.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 5.418
Siemens DCA Vantage | 72 | 5.88 | -0.03 | 2.3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 4.478
Siemens DCA Vantage | 72 | 8.63 | -0.06 | 2.6 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 72 | 5.8 | -0.06 | 2.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 72 | 9.45 | -0.11 | 3.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 6.946
Siemens DCA Vantage | 331 | 5.43 | -0.05 | 2.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 4.85
Siemens DCA Vantage | 331 | 8.31 | 0.05 | 2.7 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 5.342
Siemens DCA Vantage | 331 | 5.09 | -0.06 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 331 | 10.19 | 0.14 | 3.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 7.784
Siemens DCA Vantage | 331 | 5.93 | -0.06 | 2.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 5.036
Siemens DCA Vantage | 321 | 8.88 | 0.02 | 2.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 5.508
Siemens DCA Vantage | 321 | 6.84 | 0 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 4.508
Siemens DCA Vantage | 321 | 9.46 | 0.07 | 3.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 6.146
Siemens DCA Vantage | 321 | 6.09 | -0.04 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 4.86
Siemens DCA Vantage | 321 | 7.47 | -0.05 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 4.654
Siemens DCA Vantage | 80 | 5.76 | -0.04 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 4.272
Siemens DCA Vantage | 80 | 9.22 | 0.1 | 2.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 5.588
Siemens DCA Vantage | 80 | 5.41 | -0.06 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 4.252
Siemens DCA Vantage | 80 | 7.33 | 0.08 | 2.2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 4.392
Siemens DCA Vantage | 80 | 9.31 | 0.05 | 2.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 4.95
Siemens DCA Vantage | 306 | 5.13 | -0.08 | 2.8 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 5.408
Siemens DCA Vantage | 306 | 6.64 | -0.04 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 5.056
Siemens DCA Vantage | 306 | 8.26 |  | 3.3 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 6.468
Siemens DCA Vantage | 306 | 5.97 | -0.01 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 4.694
Siemens DCA Vantage | 306 | 10.11 | -0.01 | 3.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 7.242
Siemens Dimension ExL | 214 | 5.62 | 0.16 | 4.1 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 8.196
Siemens Dimension ExL | 214 | 5.94 | 0.28 | 3.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 7.532
Siemens Dimension ExL | 214 | 9.03 | -0.28 | 2.2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 4.032
Siemens Dimension ExL | 214 | 5.41 | 0.13 | 4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 7.97
Siemens Dimension ExL | 214 | 7.43 | 0.02 | 2.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 5.704
Siemens Dimension ExL | 12 | 5.4 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 7.252
Siemens Dimension ExL | 12 | 8.4 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 7.056
Siemens Dimension ExL | 12 | 6.2 |  | 3.1 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 6.076
Siemens Dimension ExL | 25 | 6.62 | 0.02 | 4.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 9.036
Siemens Dimension ExL | 25 | 9.32 | -0.18 | 5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 9.62
Siemens Dimension ExL | 25 | 7.28 | -0.12 | 5.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 10.072
Siemens Dimension ExL | 39 | 6.03 | 0.13 | 2.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 4.442
Siemens Dimension ExL | 39 | 9.6 | -0.2 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 7.64
Siemens Dimension ExL | 39 | 7.19 | -0.21 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 4.494
Siemens Dimension ExL | 51 | 5.48 | 0.28 | 3.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 7.14
Siemens Dimension ExL | 51 | 8.42 | -0.28 | 3.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 6.972
Siemens Dimension ExL | 51 | 6.33 | 0.03 | 3.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 6.106
Siemens Dimension ExL | 59 | 8.55 | 0.05 | 3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.93
Siemens Dimension ExL | 59 | 5.72 | 0.32 | 2.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.22
Siemens Dimension ExL | 59 | 6.55 | 0.15 | 2.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.834
Siemens Dimension ExL | 131 | 7.14 | 0.03 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 4.342
Siemens Dimension ExL | 131 | 9.44 | 0.12 | 2.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 5.216
Siemens Dimension ExL | 131 | 6.29 | 0.22 | 2.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 5.904
Siemens Dimension ExL | 173 | 8.14 | 0.09 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 5.97
Siemens Dimension ExL | 173 | 5.46 | 0.16 | 3.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 7.412
Siemens Dimension ExL | 173 | 6.39 | 0.25 | 3.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 6.522
Siemens Dimension ExL | 185 | 6.76 | 0.18 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 5.276
Siemens Dimension ExL | 185 | 8.35 | -0.04 | 2.9 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 5.644
Siemens Dimension ExL | 185 | 5.78 | 0.13 | 3.5 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 6.99
Siemens Dimension ExL | 188 | 6.85 | 0.36 | 3.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 6.828
Siemens Dimension ExL | 188 | 7.25 | 0.28 | 3.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 6.944
Siemens Dimension ExL | 188 | 9.64 | -0.01 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 4.694
Siemens Dimension ExL | 193 | 6.89 | 0.1 | 3.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 6.96
Siemens Dimension ExL | 193 | 10.23 | -0.05 | 3.4 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 6.614
Siemens Dimension ExL | 193 | 6.92 | 0.1 | 3.2 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 6.372
Siemens Dimension ExL | 128 | 8.65 | 0.02 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.704
Siemens Dimension ExL | 128 | 5.43 | 0.11 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 7.362
Siemens Dimension ExL | 126 | 5.94 | 0.1 | 3.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 7.352
Siemens Dimension ExL | 126 | 11.35 | -0.36 | 3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 5.52
Siemens Dimension ExL | 126 | 9.34 | -0.19 | 3.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 6.082
Siemens Dimension ExL | 126 | 5.16 | 0.12 | 4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 7.96
Siemens Dimension ExL | 126 | 7.25 | -0.13 | 4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 7.71
Siemens Dimension ExL | 134-205 | 11.26 | -0.43 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 5.646
Siemens Dimension ExL | 134-205 | 6.03 | 0.1 | 3.2 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 6.372
Siemens Dimension ExL | 134-205 | 5.3 | 0.13 | 3.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 7.382
Siemens Dimension ExL | 134-205 | 7.97 | -0.17 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 4.73
Siemens Dimension ExL | 134-205 | 9.02 | -0.28 | 2.2 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 4.032
Siemens Dimension ExL | 132-205 | 5.41 | 0.09 | 3.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 6.166
Siemens Dimension ExL | 132-205 | 9.12 | -0.04 | 2.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 5.252
Siemens Dimension ExL | 132-205 | 5.41 | 0.1 | 3.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 6.372
Siemens Dimension ExL | 132-205 | 11.8 | -0.23 | 3.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 5.846
Siemens Dimension ExL | 132-205 | 5.98 | 0.04 | 3.4 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 6.704
Siemens Dimension ExL | 134 | 5.31 | 0.04 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 5.724
Siemens Dimension ExL | 134 | 10.24 | -0.35 | 2.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 3.962
Siemens Dimension ExL | 134 | 6.17 | -0.03 | 3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 5.85
Siemens Dimension ExL | 134 | 11.64 | -0.59 | 2.7 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 4.702
Siemens Dimension ExL | 134 | 7.44 | -0.07 | 2.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 5.026
Siemens Dimension ExL | 142-209 | 9.02 | -0.09 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 4.81
Siemens Dimension ExL | 142-209 | 6.06 | 0.05 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 4.558
Siemens Dimension ExL | 142-209 | 11.39 | -0.32 | 2.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 4.188
Siemens Dimension ExL | 142-209 | 5.1 | 0.08 | 3.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 6.744
Siemens Dimension ExL | 142-209 | 7.53 | -0.05 | 3.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 6.026
Siemens Dimension ExL | 140-209 | 6.62 | 0.21 | 3.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 7.266
Siemens Dimension ExL | 140-209 | 9.58 | 0.05 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 5.342
Siemens Dimension ExL | 140-209 | 5.46 | 0.12 | 3.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 7.372
Siemens Dimension ExL | 140-209 | 8.65 | 0.14 | 2.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 5.628
Siemens Dimension ExL | 140-209 | 7.35 | 0.1 | 2.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 5.392
Siemens Dimension ExL | 207 | 7.19 | 0.04 | 2.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 5.724
Siemens Dimension ExL | 207 | 5.26 | 0.07 | 3.6 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 7.126
Siemens Dimension ExL | 207 | 8.16 | -0.26 | 2.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 4.836
Siemens Dimension ExL | 207 | 9.47 | -0.32 | 2.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 4.776
Siemens Dimension ExL | 207 | 6.16 | 0.04 | 3.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 6.704
Siemens Dimension ExL | 148 | 6.39 | 0.08 | 3.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 6.156
Siemens Dimension ExL | 148 | 8.89 | -0.22 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 4.876
Siemens Dimension ExL | 148 | 5.4 | 0.09 | 2.7 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 5.382
Siemens Dimension ExL | 148 | 8.37 | 0.02 | 2.5 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 4.92
Siemens Dimension ExL | 148 | 7.47 | -0.15 | 2.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 4.75
Siemens Dimension ExL | 210 | 9.2 | -0.17 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 3.946
Siemens Dimension ExL | 210 | 5.1 | 0.18 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 6.256
Siemens Dimension ExL | 210 | 6.21 | 0.17 | 2.7 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 5.462
Siemens Dimension ExL | 210 | 5.46 | 0.18 | 2.8 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 5.668
Siemens Dimension ExL | 210 | 8.08 | 0 | 2.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 5.096
Siemens Dimension ExL | 149 | 6.48 | 0.07 | 2.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 5.558
Siemens Dimension ExL | 149 | 8.5 | -0.1 | 2.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 5.192
Siemens Dimension ExL | 149 | 5.51 | 0.09 | 2.5 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 4.99
Siemens Dimension ExL | 149 | 7.39 | 0.01 | 2.9 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 5.694
Siemens Dimension ExL | 149 | 9.58 | -0.17 | 2.6 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 4.926
Siemens Dimension ExL | 197 | 5.31 | 0.06 | 2.8 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 5.548
Siemens Dimension ExL | 197 | 6.57 | 0.16 | 2.5 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 5.06
Siemens Dimension ExL | 197 | 8.08 | -0.13 | 2.2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 4.182
Siemens Dimension ExL | 197 | 5.54 | 0.07 | 2.8 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 5.558
Siemens Dimension ExL | 197 | 9.28 | -0.2 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 4.308
Siemens Dimension ExL | 189 | 8.03 | 0.06 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 3.98
Siemens Dimension ExL | 189 | 6.07 | 0.18 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 4.884
Siemens Dimension ExL | 189 | 5.27 | 0.14 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 5.824
Siemens Dimension ExL | 189 | 7.44 | 0.04 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 4.352
Siemens Dimension ExL | 189 | 9.02 | -0.15 | 2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 3.77
Siemens Dimension ExL | 144 | 7.65 | -0.16 | 2.3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 4.348
Siemens Dimension ExL | 144 | 6 | 0.09 | 2.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 5.774
Siemens Dimension ExL | 144 | 8.48 | -0.21 | 2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 3.71
Siemens Dimension ExL | 144 | 6.04 | 0.18 | 3.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 6.452
Siemens Dimension ExL | 144 | 9.31 | -0.25 | 2.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 4.454
Siemens Dimension ExL | 189 | 5.54 | 0.06 | 2.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 4.96
Siemens Dimension ExL | 189 | 8.06 | -0.2 | 2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 3.72
Siemens Dimension ExL | 189 | 5.19 | 0.04 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 5.724
Siemens Dimension ExL | 189 | 9.83 | -0.22 | 2.1 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 3.896
Siemens Dimension ExL | 189 | 6.04 | 0.05 | 2.4 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 4.754
Siemens Dimension ExL | 185 | 8.86 | 0 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 4.312
Siemens Dimension ExL | 185 | 6.88 | 0.04 | 2.2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 4.352
Siemens Dimension ExL | 185 | 9.24 | -0.15 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 4.75
Siemens Dimension ExL | 185 | 6.17 | 0.04 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 4.352
Siemens Dimension ExL | 185 | 7.52 | 0 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 4.704
Siemens Dimension ExL | 136 | 5.91 | 0.11 | 3.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 6.186
Siemens Dimension ExL | 136 | 8.97 | -0.15 | 2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 3.77
Siemens Dimension ExL | 136 | 5.54 | 0.07 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 4.382
Siemens Dimension ExL | 136 | 7.31 | 0.06 | 2.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 4.764
Siemens Dimension ExL | 136 | 9.15 | -0.11 | 2.1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 4.006
Siemens Dimension ExL | 185 | 5.27 | 0.06 | 2.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 4.96
Siemens Dimension ExL | 185 | 6.7 | 0.02 | 1.9 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 3.744
Siemens Dimension ExL | 185 | 8.18 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 3.528
Siemens Dimension ExL | 185 | 6.1 | 0.12 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 3.844
Siemens Dimension ExL | 185 | 9.92 | -0.2 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 3.524
Siemens Dimension ExL | 144 | 7.61 | 0.19 | 3 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 6.07
Siemens Dimension ExL | 144 | 5.3 | 0.1 | 3.1 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 6.176
Siemens Dimension ExL | 144 | 8.28 | -0.03 | 2.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 4.674
Siemens Dimension ExL | 144 | 10.48 | 0.15 | 2.4 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 4.854
Siemens Dimension ExL | 144 | 6.03 | 0.16 | 3.6 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 7.216
Siemens Dimension ExL | 206 | 7.29 | 0.04 | 2.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 4.94
Siemens Dimension ExL | 206 | 6.96 | 0.15 | 3.1 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 6.226
Siemens Dimension ExL | 206 | 8.53 | -0.12 | 2.9 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 5.564
Siemens Dimension ExL | 206 | 9.49 | -0.01 | 2.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 4.89
Siemens Dimension ExL | 206 | 5.51 | 0.06 | 3.1 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 6.136
Siemens Dimension ExL new reagent | 196 | 6.25 | -0.05 | 2.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 4.85
Siemens Dimension ExL new reagent | 196 | 7.53 | -0.07 | 3.1 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 6.006
Siemens Dimension ExL new reagent | 196 | 9.02 | -0.18 | 3.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 6.092
Siemens Dimension ExL new reagent | 62 | 5.77 | 0.17 | 3.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 6.638
Siemens Dimension ExL new reagent | 62 | 9.57 | 0.17 | 3.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 6.638
Siemens Dimension ExL new reagent | 62 | 7.34 | 0.14 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 5.432
Siemens Dimension ExL new reagent | 113 | 5.67 | 0.27 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 6.738
Siemens Dimension ExL new reagent | 113 | 8.08 | -0.22 | 3.1 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 5.856
Siemens Dimension ExL new reagent | 113 | 5.9 | 0.25 | 3.7 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 7.502
Siemens Dimension ExL original reagent | 44 | 8.58 | 0.08 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 6.156
Siemens Dimension ExL original reagent | 44 | 5.71 | 0.31 | 2.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 4.818
Siemens Dimension ExL original reagent | 44 | 6.55 | 0.15 | 2.6 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.246
Siemens Dimension ExL original reagent | 29 | 5.75 | 0.15 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 5.246
Siemens Dimension ExL original reagent | 29 | 9.22 | -0.18 | 3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 5.7
Siemens Dimension ExL original reagent | 29 | 7.08 | -0.12 | 4.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 8.112
Siemens Dimension ExL original reagent | 32 | 6.39 | 0.09 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 7.734
Siemens Dimension ExL original reagent | 32 | 7.67 | 0.07 | 2.4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 4.774
Siemens Dimension ExL original reagent | 32 | 9.26 | 0.06 | 3.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 6.332
Siemens Dimension RxL | 10 | 5.71 | 0.25 | 3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 6.13
Siemens Dimension RxL | 10 | 5.98 | 0.32 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 5.612
Siemens Dimension RxL | 10 | 9.06 | -0.25 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 3.67
Siemens Dimension RxL | 10 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension RxL | 10 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension RxL | 346 | 5.4 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 7.252
Siemens Dimension RxL | 346 | 8.4 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 8.036
Siemens Dimension RxL | 346 | 6.1 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 7.644
Siemens Dimension RxL | 349 | 6.67 | 0.07 | 3.3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 6.538
Siemens Dimension RxL | 349 | 9.41 | -0.09 | 4.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 8.534
Siemens Dimension RxL | 349 | 7.29 | -0.11 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 7.142
Siemens Dimension RxL | 294 | 6.03 | 0.13 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 6.206
Siemens Dimension RxL | 294 | 9.53 | -0.27 | 4.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 7.962
Siemens Dimension RxL | 294 | 7.21 | -0.19 | 3.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 5.886
Siemens Dimension RxL | 252 | 5.45 | 0.25 | 3.1 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 6.326
Siemens Dimension RxL | 252 | 8.43 | -0.27 | 3.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 6.394
Siemens Dimension RxL | 252 | 6.34 | 0.04 | 3.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 6.116
Siemens Dimension RxL | 236 | 8.47 | -0.03 | 2.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.458
Siemens Dimension RxL | 236 | 5.69 | 0.29 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.582
Siemens Dimension RxL | 236 | 6.53 | 0.13 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.422
Siemens Dimension RxL | 118 | 7.17 | 0.06 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 4.372
Siemens Dimension RxL | 118 | 9.38 | 0.06 | 2.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 4.764
Siemens Dimension RxL | 118 | 6.33 | 0.26 | 2.7 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 5.552
Siemens Dimension RxL | 86 | 8.16 | 0.11 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 5.99
Siemens Dimension RxL | 86 | 5.47 | 0.17 | 3.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 7.422
Siemens Dimension RxL | 86 | 6.4 | 0.26 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 6.14
Siemens Dimension RxL | 67 | 6.77 | 0.19 | 3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 6.07
Siemens Dimension RxL | 67 | 8.34 | -0.05 | 2.5 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 4.85
Siemens Dimension RxL | 67 | 5.8 | 0.15 | 3.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 6.226
Siemens Dimension RxL | 86 | 6.88 | 0.39 | 3.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 7.642
Siemens Dimension RxL | 86 | 7.27 | 0.3 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 6.376
Siemens Dimension RxL | 86 | 9.67 | 0.02 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 4.724
Siemens Dimension RxL | 38 | 6.95 | 0.16 | 2.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 4.472
Siemens Dimension RxL | 38 | 10.32 | 0.04 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 4.94
Siemens Dimension RxL | 38 | 6.91 | 0.09 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 4.99
Siemens Dimension RxL | 23 | 8.71 | 0.08 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.372
Siemens Dimension RxL | 23 | 5.43 | 0.11 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 7.362
Siemens Dimension RxL | 25 | 5.94 | 0.1 | 3.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 7.548
Siemens Dimension RxL | 25 | 11.4 | -0.31 | 2.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 4.982
Siemens Dimension RxL | 25 | 9.35 | -0.18 | 2.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 3.936
Siemens Dimension RxL | 25 | 5.2 | 0.16 | 3.5 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 7.02
Siemens Dimension RxL | 25 | 7.24 | -0.14 | 2.2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 4.172
Siemens Dimension RxL | 24-38 | 11.23 | -0.46 | 3.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 6.204
Siemens Dimension RxL | 24-38 | 6.04 | 0.11 | 2.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 5.794
Siemens Dimension RxL | 24-38 | 5.28 | 0.11 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 6.382
Siemens Dimension RxL | 24-38 | 7.93 | -0.21 | 2.7 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 5.082
Siemens Dimension RxL | 24-38 | 9.04 | -0.26 | 1.6 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 2.876
Siemens Dimension RxL | 16-23 | 5.34 | 0.02 | 2.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 5.704
Siemens Dimension RxL | 16-23 | 9.16 | 0 | 1.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 3.724
Siemens Dimension RxL | 16-23 | 5.37 | 0.06 | 3.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 6.92
Siemens Dimension RxL | 16-23 | 11.89 | -0.14 | 2.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 3.976
Siemens Dimension RxL | 16-23 | 5.96 | 0.02 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 5.116
Siemens Dimension RxL | 16 | 5.31 | 0.04 | 2.4 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 4.744
Siemens Dimension RxL | 16 | 10.28 | -0.31 | 1.6 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 2.826
Siemens Dimension RxL | 16 | 6.19 | -0.01 | 1.4 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 2.734
Siemens Dimension RxL | 16 | 11.76 | -0.47 | 2.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 4.626
Siemens Dimension RxL | 16 | 7.47 | -0.04 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 4.272
Siemens Dimension RxL | 14-22 | 9.02 | -0.09 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 3.046
Siemens Dimension RxL | 14-22 | 6.05 | 0.04 | 1.7 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 3.372
Siemens Dimension RxL | 14-22 | 11.42 | -0.29 | 2.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 3.826
Siemens Dimension RxL | 14-22 | 5.13 | 0.11 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 5.598
Siemens Dimension RxL | 14-22 | 7.58 | 0 | 2.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 4.704
Siemens Dimension RxL | 13-18 | 6.56 | 0.15 | 3.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 7.402
Siemens Dimension RxL | 13-18 | 9.47 | -0.06 | 2.5 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 4.84
Siemens Dimension RxL | 13-18 | 5.42 | 0.08 | 3.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 7.332
Siemens Dimension RxL | 13-18 | 8.57 | 0.06 | 1.5 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 3
Siemens Dimension RxL | 13-18 | 7.26 | 0.01 | 2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 3.93
Siemens Dimension RxL | 14 | 7.22 | 0.07 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 3.402
Siemens Dimension RxL | 14 | 5.24 | 0.05 | 2.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 4.754
Siemens Dimension RxL | 14 | 8.25 | -0.17 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 2.378
Siemens Dimension RxL | 14 | 9.39 | -0.4 | 3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 5.48
Siemens Dimension RxL | 14 | 6 | -0.12 | 4.2 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 8.112
Siemens Dimension RxL | 12 | 6.46 | 0.15 | 3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 6.03
Siemens Dimension RxL | 12 | 8.92 | -0.19 | 2.8 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 5.298
Siemens Dimension RxL | 12 | 5.46 | 0.15 | 4.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 8.578
Siemens Dimension RxL | 12 | 8.31 | -0.04 | 4.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 7.996
Siemens Dimension RxL | 12 | 7.45 | -0.17 | 3.9 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 7.474
Siemens Dimension RxL | 11 | 9.18 | -0.19 | 2.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 5.102
Siemens Dimension RxL | 11 | 5.13 | 0.21 | 3.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 7.07
Siemens Dimension RxL | 11 | 6.23 | 0.19 | 2.2 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 4.502
Siemens Dimension RxL | 11 | 5.4 | 0.12 | 4.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 9.136
Siemens Dimension RxL | 11 | 8.19 | 0.11 | 2.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 5.402
Siemens Dimension RxL | 13 | 7.62 | 0.2 | 3.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 7.452
Siemens Dimension RxL | 13 | 5.25 | 0.05 | 5.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 10.242
Siemens Dimension RxL | 13 | 8.31 | 0 | 3.2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 6.272
Siemens Dimension RxL | 13 | 10.47 | 0.14 | 3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 6.02
Siemens Dimension RxL | 13 | 5.98 | 0.11 | 5.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 10.694
Siemens Dimension RxL | 15 | 7.29 | 0.04 | 4.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 8.468
Siemens Dimension RxL | 15 | 7.05 | 0.24 | 3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 6.12
Siemens Dimension RxL | 15 | 8.59 | -0.06 | 1.7 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 3.272
Siemens Dimension RxL | 15 | 9.44 | -0.06 | 2.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 4.84
Siemens Dimension RxL | 15 | 5.54 | 0.09 | 4.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 8.322
Siemens Dimension RxL new reagent | 34 | 6.32 | 0.02 | 4.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 8.252
Siemens Dimension RxL new reagent | 34 | 7.52 | -0.08 | 3.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 6.78
Siemens Dimension RxL new reagent | 34 | 9.2 | 0 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.488
Siemens Dimension RxL new reagent | 114 | 5.67 | 0.07 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 5.95
Siemens Dimension RxL new reagent | 114 | 9.62 | 0.22 | 2.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 5.512
Siemens Dimension RxL new reagent | 114 | 7.29 | 0.09 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 5.578
Siemens Dimension RxL new reagent | 131 | 5.7 | 0.3 | 3.4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 6.964
Siemens Dimension RxL new reagent | 131 | 8.18 | -0.12 | 3.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 6.936
Siemens Dimension RxL new reagent | 131 | 5.92 | 0.27 | 3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 6.15
Siemens Dimension RxL original reagent | 178 | 8.48 | -0.02 | 2.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.272
Siemens Dimension RxL original reagent | 178 | 5.69 | 0.29 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.582
Siemens Dimension RxL original reagent | 178 | 6.53 | 0.13 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 5.422
Siemens Dimension RxL original reagent | 76 | 5.72 | 0.12 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 6
Siemens Dimension RxL original reagent | 76 | 9.22 | -0.18 | 3.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 5.896
Siemens Dimension RxL original reagent | 76 | 7.07 | -0.13 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 4.77
Siemens Dimension RxL original reagent | 11 | 5.55 | 0.15 | 2.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 5.638
Siemens Dimension RxL original reagent | 11 | 8.09 | -0.21 | 2.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 4.886
Siemens Dimension RxL original reagent | 11 | 5.82 | 0.17 | 2.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 5.07
Siemens Dimension RxL original reagent | 76 | 6.42 | 0.12 | 2.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 5.02
Siemens Dimension RxL original reagent | 76 | 7.66 | 0.06 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.98
Siemens Dimension RxL original reagent | 76 | 9.26 | 0.06 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 4.764
Siemens Dimension series | 32 | 9.0399 | 5.2 | 8.5 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Siemens | 21.86
Siemens Dimension series | 33 | 5.9381 | 3.2 | 7.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Siemens | 17.116
Siemens Dimension series | 17 | 6.176 | 1.7 | 5.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Siemens | 11.696
Siemens Dimension series | 18 | 7.3655 | -0.9 | 4.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Siemens | 7.332
Siemens Dimension series | 47 | 6.0937 | 0.8 | 4.3 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Siemens | 9.228
Siemens Dimension series | 46 | 7.3655 | -0.9 | 3.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Siemens | 6.352
Siemens Dimension series | 18 | 6.4688 | 3 | 4.3 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Siemens | 11.428
Siemens Dimension series | 19 | 7.5119 | 0.6 | 5.7 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Siemens | 11.772
Siemens Dimension series | 56 | 6.3956 | 2.2 | 3.9 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Siemens | 9.844
Siemens Dimension series | 56 | 7.4844 | 0.3 | 4 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Siemens | 8.14
Siemens Dimension series | 23 | 7.6125 | 1.6 | 5.4 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Siemens | 12.184
Siemens Dimension series | 23 | 6.4139 | 4.4 | 4.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Siemens | 13.808
Siemens Dimension series | 93 | 7.3929 | -0.8 | 3.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Siemens | 6.648
Siemens Dimension series | 93 | 6.1943 | 2 | 5.3 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Siemens | 12.388
Siemens Dimension series | 102 | 8.3903 | -1.9 | 2.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Siemens | 3.784
Siemens Dimension series | 102 | 5.7459 | 1.1 | 4.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Siemens | 10.116
Siemens Dimension series | 33 | 5.8008 | 3.4 | 5.2 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Siemens | 13.592
Siemens Dimension series | 32 | 5.7642 | 3 | 5.6 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Siemens | 13.976
Siemens Dimension series | 109 | 5.7185 | 2.5 | 4.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 11.32
Siemens Dimension series | 109 | 5.7093 | 2.4 | 4.8 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Siemens | 11.808
Siemens Dimension Vista | 290 | 5.5 | 0.04 | 3.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 6.9
Siemens Dimension Vista | 290 | 5.84 | 0.18 | 3.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 6.648
Siemens Dimension Vista | 290 | 8.91 | -0.4 | 2.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 4.108
Siemens Dimension Vista | 290 | 5.3 | 0.02 | 3.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 6.684
Siemens Dimension Vista | 290 | 7.55 | 0.14 | 2.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 5.04
Siemens Dimension Vista | 22 | 4.9 |  | 4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 7.84
Siemens Dimension Vista | 22 | 8.6 |  | 2.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 4.508
Siemens Dimension Vista | 22 | 6.1 |  | 5.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 10.976
Siemens Dimension Vista | 36 | 6.84 | 0.24 | 6.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 12.98
Siemens Dimension Vista | 36 | 9.21 | -0.29 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 4.414
Siemens Dimension Vista | 36 | 7.45 | 0.05 | 4.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 8.674
Siemens Dimension Vista | 61 | 5.65 | -0.25 | 2.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 5.042
Siemens Dimension Vista | 61 | 9.23 | -0.57 | 1.7 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 2.762
Siemens Dimension Vista | 61 | 7.64 | 0.24 | 3.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 7.688
Siemens Dimension Vista | 113 | 5.12 | -0.08 | 3.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 6.192
Siemens Dimension Vista | 113 | 8.75 | 0.05 | 2.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 4.166
Siemens Dimension Vista | 113 | 6.21 | -0.09 | 3.3 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 6.378
Siemens Dimension Vista | 125 | 8.75 | 0.25 | 2.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 4.758
Siemens Dimension Vista | 125 | 5.39 | -0.01 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.478
Siemens Dimension Vista | 125 | 6.66 | 0.26 | 4.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 8.492
Siemens Dimension Vista | 204 | 7.34 | 0.23 | 2.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 5.718
Siemens Dimension Vista | 204 | 9.49 | 0.17 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 5.854
Siemens Dimension Vista | 204 | 6.25 | 0.18 | 3.3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 6.648
Siemens Dimension Vista | 236 | 8.18 | 0.13 | 3.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 6.794
Siemens Dimension Vista | 236 | 5.25 | -0.05 | 3.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 7.006
Siemens Dimension Vista | 236 | 6.21 | 0.07 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 5.95
Siemens Dimension Vista | 269 | 6.81 | 0.23 | 3.1 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 6.306
Siemens Dimension Vista | 269 | 8.4 | 0.01 | 2.5 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 4.91
Siemens Dimension Vista | 269 | 5.73 | 0.08 | 2.9 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 5.764
Siemens Dimension Vista | 270 | 6.84 | 0.35 | 3.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 6.818
Siemens Dimension Vista | 270 | 7.4 | 0.43 | 2.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 6.114
Siemens Dimension Vista | 270 | 9.83 | 0.18 | 3.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 7.236
Siemens Dimension Vista | 259 | 6.98 | 0.19 | 2.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 5.09
Siemens Dimension Vista | 259 | 10.16 | -0.12 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 4.976
Siemens Dimension Vista | 259 | 6.97 | 0.15 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 5.246
Siemens Dimension Vista | 232 | 8.63 | 0 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.488
Siemens Dimension Vista | 232 | 5.45 | 0.13 | 2.6 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 5.226
Siemens Dimension Vista | 251 | 5.95 | 0.11 | 2.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 5.598
Siemens Dimension Vista | 251 | 11.84 | 0.13 | 1.9 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 3.854
Siemens Dimension Vista | 251 | 9.34 | -0.19 | 2.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 4.122
Siemens Dimension Vista | 251 | 5.09 | 0.05 | 3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 5.93
Siemens Dimension Vista | 251 | 7.44 | 0.06 | 2.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 4.96
Siemens Dimension Vista | 248-283 | 11.49 | -0.2 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 5.876
Siemens Dimension Vista | 248-283 | 6.04 | 0.11 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 5.402
Siemens Dimension Vista | 248-283 | 5.22 | 0.05 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 6.322
Siemens Dimension Vista | 248-283 | 8.19 | 0.05 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 5.146
Siemens Dimension Vista | 248-283 | 9.12 | -0.18 | 2.4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 4.524
Siemens Dimension Vista | 260-285 | 5.4 | 0.08 | 2.7 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 5.372
Siemens Dimension Vista | 260-285 | 9.05 | -0.11 | 2.1 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 4.006
Siemens Dimension Vista | 260-285 | 5.4 | 0.09 | 2.7 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 5.382
Siemens Dimension Vista | 260-285 | 12.08 | 0.05 | 1.9 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 3.774
Siemens Dimension Vista | 260-285 | 6.01 | 0.07 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 5.166
Siemens Dimension Vista | 259 | 5.36 | 0.09 | 2.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 4.598
Siemens Dimension Vista | 259 | 10.37 | -0.22 | 1.8 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 3.308
Siemens Dimension Vista | 259 | 6.23 | 0.03 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 4.538
Siemens Dimension Vista | 259 | 12.18 | -0.05 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 3.87
Siemens Dimension Vista | 259 | 7.6 | 0.09 | 2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 4.01
Siemens Dimension Vista | 260-291 | 8.86 | -0.25 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 4.65
Siemens Dimension Vista | 260-291 | 6.06 | 0.05 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 5.538
Siemens Dimension Vista | 260-291 | 11.39 | -0.32 | 2.2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 3.992
Siemens Dimension Vista | 260-291 | 5.01 | -0.01 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 6.066
Siemens Dimension Vista | 260-291 | 7.66 | 0.08 | 2.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 5.372
Siemens Dimension Vista | 274-299 | 6.61 | 0.2 | 2.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 5.884
Siemens Dimension Vista | 274-299 | 9.35 | -0.18 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 5.112
Siemens Dimension Vista | 274-299 | 5.44 | 0.1 | 3.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 6.176
Siemens Dimension Vista | 274-299 | 8.61 | 0.1 | 2.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 5.588
Siemens Dimension Vista | 274-299 | 7.47 | 0.22 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 5.12
Siemens Dimension Vista | 303 | 7.42 | 0.27 | 2.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 5.366
Siemens Dimension Vista | 303 | 5.14 | -0.05 | 3.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 7.398
Siemens Dimension Vista | 303 | 8.48 | 0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 4.764
Siemens Dimension Vista | 303 | 9.51 | -0.28 | 2.4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 4.424
Siemens Dimension Vista | 303 | 6.23 | 0.11 | 2.8 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 5.598
Siemens Dimension Vista | 280 | 6.49 | 0.18 | 3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 6.06
Siemens Dimension Vista | 280 | 8.86 | -0.25 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 4.846
Siemens Dimension Vista | 280 | 5.36 | 0.05 | 2.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 5.734
Siemens Dimension Vista | 280 | 8.47 | 0.12 | 3 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 6
Siemens Dimension Vista | 280 | 7.63 | 0.01 | 3 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 5.89
Siemens Dimension Vista | 309 | 9.14 | -0.23 | 4.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 7.806
Siemens Dimension Vista | 309 | 4.98 | 0.06 | 3.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 7.116
Siemens Dimension Vista | 309 | 6.2 | 0.16 | 3.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 6.628
Siemens Dimension Vista | 309 | 5.38 | 0.1 | 3.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 6.372
Siemens Dimension Vista | 309 | 8.23 | 0.15 | 4.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 8.382
Siemens Dimension Vista | 276 | 6.58 | 0.17 | 2.6 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 5.266
Siemens Dimension Vista | 276 | 8.51 | -0.09 | 2.4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 4.614
Siemens Dimension Vista | 276 | 5.49 | 0.07 | 2.9 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 5.754
Siemens Dimension Vista | 276 | 7.54 | 0.16 | 2.3 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 4.668
Siemens Dimension Vista | 276 | 9.35 | -0.4 | 2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 3.52
Siemens Dimension Vista | 279 | 5.23 | -0.02 | 2.9 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 5.664
Siemens Dimension Vista | 279 | 6.49 | 0.08 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 5.372
Siemens Dimension Vista | 279 | 8.03 | -0.18 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 3.936
Siemens Dimension Vista | 279 | 5.47 | 0 | 2.6 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 5.096
Siemens Dimension Vista | 279 | 9.14 | -0.34 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 4.168
Siemens Dimension Vista | 243 | 8.03 | 0.06 | 3.4 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 6.724
Siemens Dimension Vista | 243 | 5.98 | 0.09 | 3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 5.97
Siemens Dimension Vista | 243 | 5.13 | 0 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 5.684
Siemens Dimension Vista | 243 | 7.41 | 0.01 | 3.7 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 7.262
Siemens Dimension Vista | 243 | 8.98 | -0.19 | 2.1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 3.926
Siemens Dimension Vista | 245 | 7.68 | -0.13 | 3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 5.75
Siemens Dimension Vista | 245 | 5.94 | 0.03 | 3.3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 6.498
Siemens Dimension Vista | 245 | 8.47 | -0.22 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 4.092
Siemens Dimension Vista | 245 | 5.9 | 0.04 | 3.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 6.508
Siemens Dimension Vista | 245 | 9.3 | -0.26 | 2.1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 3.856
Siemens Dimension Vista | 232 | 5.47 | -0.01 | 3.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 6.262
Siemens Dimension Vista | 232 | 8.07 | -0.19 | 3.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 6.082
Siemens Dimension Vista | 232 | 5.15 | 0 | 2.8 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 5.488
Siemens Dimension Vista | 232 | 9.86 | -0.19 | 2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 3.73
Siemens Dimension Vista | 232 | 5.95 | -0.04 | 3.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 7.016
Siemens Dimension Vista | 228 | 8.49 | -0.37 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 3.942
Siemens Dimension Vista | 228 | 6.6 | -0.24 | 1.7 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 3.092
Siemens Dimension Vista | 228 | 9.12 | -0.27 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 3.258
Siemens Dimension Vista | 228 | 5.99 | -0.14 | 1.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 2.8
Siemens Dimension Vista | 228 | 7.11 | -0.41 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 3.118
Siemens Dimension Vista | 216 | 5.69 | -0.11 | 2.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 4.006
Siemens Dimension Vista | 216 | 8.83 | -0.29 | 1.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 3.238
Siemens Dimension Vista | 216 | 5.32 | -0.15 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 3.574
Siemens Dimension Vista | 216 | 6.89 | -0.36 | 1.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 2.776
Siemens Dimension Vista | 216 | 8.93 | -0.33 | 2.1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 3.786
Siemens Dimension Vista | 207 | 5.13 | -0.08 | 1.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 3.644
Siemens Dimension Vista | 207 | 6.5 | -0.18 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 2.956
Siemens Dimension Vista | 207 | 7.77 |  | 2 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 3.92
Siemens Dimension Vista | 207 | 5.86 | -0.12 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 3.408
Siemens Dimension Vista | 207 | 9.99 | -0.13 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 3.202
Siemens Dimension Vista | 282 | 7.7 | 0.28 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 5.376
Siemens Dimension Vista | 282 | 5.25 | 0.05 | 2.9 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 5.734
Siemens Dimension Vista | 282 | 8.25 | -0.06 | 2.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 4.644
Siemens Dimension Vista | 282 | 9.97 | -0.36 | 2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 3.56
Siemens Dimension Vista | 282 | 6.04 | 0.17 | 2.8 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 5.658
Siemens Dimension Vista | 300 | 7.45 | 0.2 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 4.708
Siemens Dimension Vista | 300 | 7.01 | 0.2 | 2.5 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 5.1
Siemens Dimension Vista | 300 | 8.46 | -0.19 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 3.926
Siemens Dimension Vista | 300 | 9.28 | -0.22 | 2.1 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 3.896
Siemens Dimension Vista | 300 | 5.54 | 0.09 | 2.8 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 5.578
Siemens Dimension Vista new reagent | 102 | 6.33 | 0.03 | 3.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 6.302
Siemens Dimension Vista new reagent | 121 | 6.59 | 0.29 | 2.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 5.778
Siemens Dimension Vista new reagent | 102 | 7.48 | -0.12 | 2.8 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 5.368
Siemens Dimension Vista new reagent | 121 | 7.8 | 0.2 | 2.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 5.884
Siemens Dimension Vista new reagent | 102 | 9.09 | -0.11 | 3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.77
Siemens Dimension Vista new reagent | 121 | 9 | -0.2 | 2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 3.72
Siemens Dimension Vista new reagent | 181 | 5.63 | 0.03 | 4.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 8.85
Siemens Dimension Vista new reagent | 181 | 9.11 | -0.29 | 4 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 7.55
Siemens Dimension Vista new reagent | 181 | 7.33 | 0.13 | 4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 7.97
Siemens Dimension Vista new reagent | 192 | 5.72 | 0.32 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 8.16
Siemens Dimension Vista new reagent | 192 | 8.61 | 0.31 | 2.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 5.798
Siemens Dimension Vista new reagent | 192 | 5.95 | 0.3 | 3.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 6.964
Siemens Dimension Vista original reagent | 25 | 6.48 | 0.18 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 7.824
Siemens Dimension Vista original reagent | 25 | 8.03 | 0.43 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 4.35
Siemens Dimension Vista original reagent | 25 | 9.15 | -0.05 | 1.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.498
Siemens Dimension Vista original reagent | 105 | 8.75 | 0.25 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 4.562
Siemens Dimension Vista original reagent | 105 | 5.39 | -0.01 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 5.478
Siemens Dimension Vista original reagent | 105 | 6.64 | 0.24 | 4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 8.08
Siemens Dimension Xpand | 15 | 5.58 | 0.12 | 2.4 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 4.824
Siemens Dimension Xpand | 15 | 5.86 | 0.2 | 3.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 6.668
Siemens Dimension Xpand | 15 | 9 | -0.31 | 3.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 6.158
Siemens Dimension Xpand | 15 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 15 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 139 | 5.3 |  | 3.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 7.448
Siemens Dimension Xpand | 139 | 8.3 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 8.036
Siemens Dimension Xpand | 139 | 6 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 7.644
Siemens Dimension Xpand | 141 | 6.6 | 0 | 3.3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 6.468
Siemens Dimension Xpand | 141 | 9.37 | -0.13 | 4.9 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 9.474
Siemens Dimension Xpand | 141 | 7.23 | -0.17 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 7.082
Siemens Dimension Xpand | 126 | 5.95 | 0.05 | 3.3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 6.518
Siemens Dimension Xpand | 126 | 9.44 | -0.36 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 7.48
Siemens Dimension Xpand | 126 | 7.13 | -0.27 | 3.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 5.806
Siemens Dimension Xpand | 120 | 5.39 | 0.19 | 3.1 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 6.266
Siemens Dimension Xpand | 120 | 8.37 | -0.33 | 3.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 6.922
Siemens Dimension Xpand | 120 | 6.28 | -0.02 | 3.3 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 6.448
Siemens Dimension Xpand | 108 | 8.43 | -0.07 | 2.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.614
Siemens Dimension Xpand | 108 | 5.67 | 0.27 | 3.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 6.346
Siemens Dimension Xpand | 108 | 6.46 | 0.06 | 3.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 6.724
Siemens Dimension Xpand | 74 | 7.11 | 0 | 2.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 4.508
Siemens Dimension Xpand | 74 | 9.41 | 0.09 | 2.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 4.206
Siemens Dimension Xpand | 74 | 6.22 | 0.15 | 3.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 6.814
Siemens Dimension Xpand | 68 | 8.12 | 0.07 | 3.1 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 6.146
Siemens Dimension Xpand | 68 | 5.42 | 0.12 | 3.2 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 6.392
Siemens Dimension Xpand | 68 | 6.34 | 0.2 | 3.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 6.472
Siemens Dimension Xpand | 59 | 6.76 | 0.18 | 3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 6.06
Siemens Dimension Xpand | 59 | 8.32 | -0.07 | 3.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 6.006
Siemens Dimension Xpand | 59 | 5.81 | 0.16 | 3.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 6.824
Siemens Dimension Xpand | 65 | 6.82 | 0.33 | 3.1 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 6.406
Siemens Dimension Xpand | 65 | 7.23 | 0.26 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 6.336
Siemens Dimension Xpand | 65 | 9.69 | 0.04 | 1.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 3.568
Siemens Dimension Xpand | 47 | 6.94 | 0.15 | 2.8 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 5.638
Siemens Dimension Xpand | 47 | 10.28 | 0 | 3.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 6.86
Siemens Dimension Xpand | 47 | 6.93 | 0.11 | 3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 5.99
Siemens Dimension Xpand | 21 | 8.67 | 0.04 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 5.724
Siemens Dimension Xpand | 21 | 5.5 | 0.18 | 3.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 7.628
Siemens Dimension Xpand | 22 | 5.87 | 0.03 | 2.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 5.518
Siemens Dimension Xpand | 22 | 11.41 | -0.3 | 2.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 4.208
Siemens Dimension Xpand | 22 | 9.4 | -0.13 | 3.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 7.122
Siemens Dimension Xpand | 22 | 5.08 | 0.04 | 3.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 6.508
Siemens Dimension Xpand | 22 | 7.26 | -0.12 | 2.8 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 5.368
Siemens Dimension Xpand | 21-42 | 11.35 | -0.34 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 5.344
Siemens Dimension Xpand | 21-42 | 6.03 | 0.1 | 4.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 8.528
Siemens Dimension Xpand | 21-42 | 5.28 | 0.11 | 3.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 7.558
Siemens Dimension Xpand | 21-42 | 8 | -0.14 | 2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 3.78
Siemens Dimension Xpand | 21-42 | 9 | -0.3 | 2.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 4.6
Siemens Dimension Xpand | 25-47 | 5.37 | 0.05 | 3.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 6.91
Siemens Dimension Xpand | 25-47 | 9.11 | -0.05 | 2.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 5.634
Siemens Dimension Xpand | 25-47 | 5.38 | 0.07 | 3.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 6.146
Siemens Dimension Xpand | 25-47 | 11.94 | -0.09 | 2.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 4.81
Siemens Dimension Xpand | 25-47 | 5.99 | 0.05 | 2.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 5.342
Siemens Dimension Xpand | 19 | 5.3 | 0.03 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 5.714
Siemens Dimension Xpand | 19 | 10.2 | -0.39 | 3.6 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 6.666
Siemens Dimension Xpand | 19 | 6.13 | -0.07 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 4.438
Siemens Dimension Xpand | 19 | 11.69 | -0.54 | 3.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 5.536
Siemens Dimension Xpand | 19 | 7.42 | -0.09 | 3.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 6.378
Siemens Dimension Xpand | 17-36 | 9.04 | -0.07 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 4.83
Siemens Dimension Xpand | 17-36 | 5.98 | -0.03 | 4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 7.81
Siemens Dimension Xpand | 17-36 | 11.43 | -0.28 | 2.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 4.62
Siemens Dimension Xpand | 17-36 | 5 | -0.02 | 3.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 6.84
Siemens Dimension Xpand | 17-36 | 7.52 | -0.06 | 2.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 4.644
Siemens Dimension Xpand | 13-29 | 6.59 | 0.18 | 2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 4.1
Siemens Dimension Xpand | 13-29 | 9.6 | 0.07 | 2.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 4.774
Siemens Dimension Xpand | 13-29 | 5.43 | 0.09 | 2.8 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 5.578
Siemens Dimension Xpand | 13-29 | 8.73 | 0.22 | 1.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 2.376
Siemens Dimension Xpand | 13-29 | 7.35 | 0.1 | 1.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.648
Siemens Dimension Xpand | 22 | 7.28 | 0.13 | 2.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 5.226
Siemens Dimension Xpand | 22 | 5.3 | 0.11 | 3.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 7.362
Siemens Dimension Xpand | 22 | 8.21 | -0.21 | 3.5 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 6.65
Siemens Dimension Xpand | 22 | 9.53 | -0.26 | 2.7 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 5.032
Siemens Dimension Xpand | 22 | 6.2 | 0.08 | 4.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 8.704
Siemens Dimension Xpand | 20 | 9.18 | -0.19 | 1.6 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 2.946
Siemens Dimension Xpand | 20 | 5.07 | 0.15 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 6.618
Siemens Dimension Xpand | 20 | 6.2 | 0.16 | 2.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 5.648
Siemens Dimension Xpand | 20 | 5.4 | 0.12 | 4.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 8.156
Siemens Dimension Xpand | 20 | 8.05 | -0.03 | 3.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 6.438
Siemens Dimension Xpand | 14 | 5.21 | -0.04 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 5.252
Siemens Dimension Xpand | 14 | 6.57 | 0.16 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 4.276
Siemens Dimension Xpand | 14 | 8.06 | -0.15 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 3.574
Siemens Dimension Xpand | 14 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 14 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 | 7.98 | 0.01 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 3.93
Siemens Dimension Xpand | 11 | 6.02 | 0.13 | 1.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 3.07
Siemens Dimension Xpand | 11 | 5.26 | 0.13 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 5.814
Siemens Dimension Xpand | 11 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 10 | 5.54 | 0.06 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 3.392
Siemens Dimension Xpand | 10 | 8.06 | -0.2 | 2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 3.72
Siemens Dimension Xpand | 10 | 5.17 | 0.02 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 3.352
Siemens Dimension Xpand | 10 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 10 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 | 7.71 | 0.29 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 5.386
Siemens Dimension Xpand | 11 | 5.33 | 0.13 | 2.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 4.442
Siemens Dimension Xpand | 11 | 8.37 | 0.06 | 2.5 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 4.96
Siemens Dimension Xpand | 11 | 10.46 | 0.13 | 2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 4.05
Siemens Dimension Xpand | 11 | 6.05 | 0.18 | 3.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 6.844
Siemens Dimension Xpand | 26 | 7.27 | 0.02 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 5.508
Siemens Dimension Xpand | 26 | 6.9 | 0.09 | 3.1 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 6.166
Siemens Dimension Xpand | 26 | 8.49 | -0.16 | 3.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 5.916
Siemens Dimension Xpand | 26 | 9.43 | -0.07 | 3.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 6.594
Siemens Dimension Xpand | 26 | 5.56 | 0.11 | 3.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 6.382
Siemens Dimension Xpand new reagent | 33 | 6.39 | 0.09 | 3.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 6.754
Siemens Dimension Xpand new reagent | 33 | 7.56 | -0.04 | 3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 5.84
Siemens Dimension Xpand new reagent | 33 | 9.24 | 0.04 | 2.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 4.156
Siemens Dimension Xpand new reagent | 65 | 5.63 | 0.03 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 7.478
Siemens Dimension Xpand new reagent | 65 | 9.58 | 0.18 | 3.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 7.04
Siemens Dimension Xpand new reagent | 65 | 7.27 | 0.07 | 3.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 6.538
Siemens Dimension Xpand new reagent | 69 | 5.66 | 0.26 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 6.728
Siemens Dimension Xpand new reagent | 69 | 8.1 | -0.2 | 2.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 5.288
Siemens Dimension Xpand new reagent | 69 | 5.9 | 0.25 | 3.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 7.11
Siemens Dimension Xpand original reagent | 37 | 5.66 | 0.06 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 5.156
Siemens Dimension Xpand original reagent | 37 | 9.19 | -0.21 | 3.4 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 6.454
Siemens Dimension Xpand original reagent | 37 | 7.06 | -0.14 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 5.348
Siemens Dimension Xpand original reagent | 13 | 5.65 | 0.25 | 2.6 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 5.346
Siemens Dimension Xpand original reagent | 13 | 8.13 | -0.17 | 2.1 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 3.946
Siemens Dimension Xpand original reagent | 13 | 5.76 | 0.11 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 5.206
Siemens Dimension Xpand original reagent | 47 | 6.3 | 0 | 2.7 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 5.292
Siemens Dimension Xpand original reagent | 47 | 7.46 | -0.14 | 2.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 5.152
Siemens Dimension Xpand original reagent | 47 | 9.15 | -0.05 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.242
Siemens Dimension Xpand original reagent | 69 | 8.43 | -0.07 | 2.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 5.614
Siemens Dimension Xpand original reagent | 69 | 5.66 | 0.26 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 6.14
Siemens Dimension Xpand original reagent | 69 | 6.42 | 0.02 | 3.6 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 7.076
Siemens Other | 13 | 5.993 | -0.3 | 4.9 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Siemens | 9.304
Siemens Other | 14 | 7.4204 | -0.3 | 3.6 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Siemens | 6.756
Synchron | 210 | 5 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Synchron | 7.252
Synchron | 210 | 8.1 |  | 3.5 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Synchron | 6.86
Synchron | 210 | 5.8 |  | 3.5 | GH2-03 | 6 | 2009 | CAP | Fresh | Synchron | 6.86
Synchron | 269 | 6.43 | -0.17 | 3.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Synchron | 6.886
Synchron | 269 | 9.18 | -0.32 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Synchron | 6.344
Synchron | 269 | 7.06 | -0.34 | 3.6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Synchron | 6.716
Synchron | 241 | 5.82 | -0.08 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Synchron | 5.996
Synchron | 241 | 9.6 | -0.2 | 3.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Synchron | 6.268
Synchron | 241 | 7.11 | -0.29 | 3 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Synchron | 5.59
Synchron | 291 | 5.32 | 0.12 | 4.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Synchron | 9.136
Synchron | 291 | 8.51 | -0.19 | 4.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Synchron | 8.434
Synchron | 291 | 6.23 | -0.07 | 4.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Synchron | 7.966
Synchron | 263 | 8.52 | 0.02 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Synchron | 6.096
Synchron | 263 | 5.45 | 0.05 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Synchron | 7.694
Synchron | 263 | 6.38 | -0.02 | 3.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Synchron | 6.644
Synchron | 270 | 6.22 | -0.08 | 3.3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Synchron | 6.388
Synchron | 270 | 7.55 | -0.05 | 3.2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Synchron | 6.222
Synchron | 270 | 9.37 | 0.17 | 3.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Synchron | 6.834
Synchron | 300 | 5.61 | 0.01 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Synchron | 7.458
Synchron | 300 | 9.53 | 0.13 | 3.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Synchron | 7.774
Synchron | 300 | 7.11 | -0.09 | 3.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Synchron | 6.77
Synchron | 268 | 5.39 | -0.01 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Synchron | 6.458
Synchron | 268 | 8.24 | -0.06 | 3.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Synchron | 6.408
Synchron | 268 | 5.48 | -0.17 | 3.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Synchron | 5.906
Sysmex | 7 | 7.1002 | -4 | 5.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Sysmex | 7.368
Sysmex | 7 | 5.691 | -3.5 | 6.1 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Sysmex | 8.456
Sysmex bx-3010 | 8 | 7.9328 | -6.9 | 10.8 | Sample 1 | 8.5641 | 2019 | EURA1c | Frozen | Sysmex | 14.268
Sysmex bx-3010 | 8 | 5.261 | -4.2 | 13 | Sample 2 | 5.6453 | 2019 | EURA1c | Frozen | Sysmex | 21.28
Thermo Fisher Scientific/Konelab | 16 | 5.7459 | 2.8 | 5.5 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Thermo Scientific | 13.58
Thermo Fisher Scientific/Konelab | 16 | 5.7002 | 2.3 | 5.1 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Thermo Scientific | 12.296
Thermo Fisher Scientific/Konelab | 17 | 7.4661 | 0 | 5.5 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Thermo Scientific | 10.78
Thermo Fisher Scientific/Konelab | 17 | 6.112 | 1.1 | 6.5 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Thermo Scientific | 13.84
Tosoh A1c 2.2 Plus | 193 | 11.6 |  | 2.9 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 193 | 5.4 |  | 2.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 193 | 8.8 |  | 2.6 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Tosoh | 5.096
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 181 | 12.6 |  | 3.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 6.076
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 181 | 12.6 |  | 3.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 6.076
Tosoh A1c 2.2 Plus | 186 | 5.4 |  | 3.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Tosoh | 6.86
Tosoh A1c 2.2 Plus | 186 | 12.4 |  | 2.7 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Tosoh | 5.292
Tosoh A1c 2.2 Plus | 186 | 8 |  | 2.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Tosoh | 5.684
Tosoh A1c 2.2 Plus | 140 | 9.8 |  | 2.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Tosoh | 4.9
Tosoh A1c 2.2 Plus | 140 | 12 |  | 2.4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Tosoh | 4.704
Tosoh A1c 2.2 Plus | 140 | 6.5 |  | 3.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Tosoh | 6.272
Tosoh A1c 2.2 Plus | 130 | 13.9 |  | 3.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Tosoh | 7.644
Tosoh A1c 2.2 Plus | 130 | 5.5 |  | 4.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Tosoh | 8.82
Tosoh A1c 2.2 Plus | 130 | 9.4 |  | 3.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Tosoh | 6.664
Tosoh A1c 2.2 Plus | 113 | 12.3 |  | 3.3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Tosoh | 6.468
Tosoh A1c 2.2 Plus | 113 | 7.2 |  | 2.8 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Tosoh | 5.488
Tosoh A1c 2.2 Plus | 113 | 5.3 |  | 3.1 | GH2-06 | 5 | 2008 | CAP | Fresh | Tosoh | 6.076
Tosoh A1c 2.2 Plus | 91 | 5.3 |  | 3.5 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 6.86
Tosoh A1c 2.2 Plus | 91 | 8.7 |  | 2.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 5.096
Tosoh A1c 2.2 Plus | 91 | 6.3 |  | 2.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 5.488
Tosoh A1c 2.2 Plus | 71 | 6.87 | 0.27 | 3.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 6.346
Tosoh A1c 2.2 Plus | 71 | 9.63 | 0.13 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 5.618
Tosoh A1c 2.2 Plus | 71 | 7.73 | 0.33 | 3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 6.21
Tosoh A1c 2.2 Plus | 46 | 6.21 | 0.31 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 5.798
Tosoh A1c 2.2 Plus | 46 | 9.96 | 0.16 | 2.9 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 5.844
Tosoh A1c 2.2 Plus | 46 | 7.66 | 0.26 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 5.748
Tosoh A1c 2.2 Plus | 14 | 5.37 | 0.17 | 2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 4.09
Tosoh A1c 2.2 Plus | 14 | 8.88 | 0.18 | 2.9 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 5.864
Tosoh A1c 2.2 Plus | 14 | 6.55 | 0.25 | 2.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 5.738
Tosoh G11 Auto HPLC | 8 | 7.5759 | 1.2 | 4.4 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Tosoh | 9.824
Tosoh G11 Auto HPLC | 8 | 6.2584 | 2.7 | 6.8 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Tosoh | 16.028
Tosoh G11 Auto HPLC | 75 | 7.5668 | 1.1 | 1.8 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Tosoh | 4.628
Tosoh G11 Auto HPLC | 76 | 6.1394 | 1.4 | 2.1 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Tosoh | 5.516
Tosoh G11 Auto HPLC | 148 | 8.7654 | 2.2 | 2.7 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Tosoh | 7.492
Tosoh G11 Auto HPLC | 148 | 5.7368 | 1 | 3.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Tosoh | 7.664
Tosoh G11 Auto HPLC | 112 | 5.6361 | 1.6 | 5.1 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 11.596
Tosoh G11 Auto HPLC | 102 | 5.6544 | 1.8 | 5.1 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 11.796
Tosoh G11 Auto HPLC | 124 | 5.5812 | 1 | 2.3 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 5.508
Tosoh G11 Auto HPLC | 124 | 5.5904 | 1.1 | 2.4 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 5.804
Tosoh G11 Auto HPLC | 11 | 5.2 | -0.05 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Tosoh | 5.046
Tosoh G11 Auto HPLC | 11 | 6.49 | 0.08 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Tosoh | 4.196
Tosoh G11 Auto HPLC | 11 | 8.27 | 0.06 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Tosoh | 3.784
Tosoh G11 Auto HPLC | 11 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 11 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 12 | 5.55 | 0.07 | 1.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Tosoh | 2.422
Tosoh G11 Auto HPLC | 12 | 8.32 | 0.06 | 0.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Tosoh | 1.824
Tosoh G11 Auto HPLC | 12 | 5.14 | -0.01 | 1.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Tosoh | 2.538
Tosoh G11 Auto HPLC | 12 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 12 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 11 | 9.07 | 0.21 | 0.9 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Tosoh | 1.974
Tosoh G11 Auto HPLC | 11 | 6.95 | 0.11 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Tosoh | 3.05
Tosoh G11 Auto HPLC | 11 | 9.54 | 0.15 | 1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Tosoh | 2.11
Tosoh G11 Auto HPLC | 11 |  |  |  | GH5-04 | 6.13 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 11 |  |  |  | GH5-05 | 7.52 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 10 | 5.21 | 0 | 1.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Tosoh | 2.156
Tosoh G11 Auto HPLC | 10 | 6.75 | 0.07 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Tosoh | 2.814
Tosoh G11 Auto HPLC | 10 | 8.36 |  | 2.3 | GH5-13 |  | 2021 | CAP | Fresh | Tosoh | 4.508
Tosoh G11 Auto HPLC | 10 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Auto HPLC | 10 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 33 | 6.0296 | 0.1 | 4.8 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Tosoh | 9.508
Tosoh G7 Auto HPLC | 34 | 7.3655 | -0.9 | 4.7 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Tosoh | 8.312
Tosoh G7 Auto HPLC | 27 | 6.0754 | 0.6 | 5.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Tosoh | 10.596
Tosoh G7 Auto HPLC | 27 | 7.5942 | 1.6 | 6.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Tosoh | 13.752
Tosoh G7 Auto HPLC | 15 | 6.3498 | 1.7 | 5.5 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Tosoh | 12.48
Tosoh G7 Auto HPLC | 17 | 7.5119 | 0.6 | 4.5 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Tosoh | 9.42
Tosoh G7 Auto HPLC | 31 | 6.3224 | 1.4 | 4.5 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Tosoh | 10.22
Tosoh G7 Auto HPLC | 30 | 7.6034 | 1.6 | 3.2 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Tosoh | 7.872
Tosoh G7 Auto HPLC | 17 | 7.5668 | 1.1 | 4.2 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Tosoh | 9.332
Tosoh G7 Auto HPLC | 17 | 6.1211 | 1.2 | 5.2 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Tosoh | 11.392
Tosoh G7 Auto HPLC | 23 | 7.5942 | 1.4 | 5.2 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Tosoh | 11.592
Tosoh G7 Auto HPLC | 24 | 6.1669 | 1.7 | 4.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Tosoh | 10.324
Tosoh G7 Auto HPLC | 19 | 8.5367 | -0.3 | 6.6 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Tosoh | 12.636
Tosoh G7 Auto HPLC | 19 | 5.6453 | 0 | 7.1 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Tosoh | 13.916
Tosoh G7 Auto HPLC | 39 | 5.6453 | 1.7 | 4.4 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 10.324
Tosoh G7 Auto HPLC | 39 | 5.6361 | 1.6 | 4.5 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 10.42
Tosoh G7 Auto HPLC | 15 | 5.5355 | 0.5 | 8 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 16.18
Tosoh G7 Auto HPLC | 15 | 5.5263 | 0.4 | 7.6 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 15.296
Tosoh G7 Auto HPLC | 196 | 11.3 |  | 2.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Tosoh | 4.116
Tosoh G7 Auto HPLC | 196 | 5.2 |  | 2.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Tosoh | 4.508
Tosoh G7 Auto HPLC | 196 | 8.6 |  | 2.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Tosoh | 4.9
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 3.724
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 3.528
Tosoh G7 Auto HPLC | 200 | 12.3 |  | 1.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 2.94
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 3.724
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 3.528
Tosoh G7 Auto HPLC | 200 | 12.3 |  | 1.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 2.94
Tosoh G7 Auto HPLC | 261 | 5.4 |  | 2.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Tosoh | 4.704
Tosoh G7 Auto HPLC | 261 | 12.3 |  | 1.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Tosoh | 3.528
Tosoh G7 Auto HPLC | 261 | 7.9 |  | 2 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Tosoh | 3.92
Tosoh G7 Auto HPLC | 246 | 9.6 |  | 2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Tosoh | 3.92
Tosoh G7 Auto HPLC | 246 | 11.8 |  | 1.7 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Tosoh | 3.332
Tosoh G7 Auto HPLC | 246 | 6.4 |  | 2.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Tosoh | 4.312
Tosoh G7 Auto HPLC | 263 | 13.7 |  | 3.8 | GH2-01 | 13 | 2008 | CAP | Fresh | Tosoh | 7.448
Tosoh G7 Auto HPLC | 263 | 5.4 |  | 1.9 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Tosoh | 3.724
Tosoh G7 Auto HPLC | 263 | 9.2 |  | 2.6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Tosoh | 5.096
Tosoh G7 Auto HPLC | 294 | 12 |  | 2.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Tosoh | 4.116
Tosoh G7 Auto HPLC | 294 | 7 |  | 1.8 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Tosoh | 3.528
Tosoh G7 Auto HPLC | 294 | 5.2 |  | 1.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Tosoh | 3.724
Tosoh G7 Auto HPLC | 240 | 5.2 |  | 1.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 3.332
Tosoh G7 Auto HPLC | 240 | 8.6 |  | 1.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 2.744
Tosoh G7 Auto HPLC | 240 | 6.2 |  | 1.5 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 2.94
Tosoh G7 Auto HPLC | 248 | 6.8 | 0.2 | 1.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 3.924
Tosoh G7 Auto HPLC | 248 | 9.53 | 0.03 | 1.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 3.558
Tosoh G7 Auto HPLC | 248 | 7.67 | 0.27 | 2.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 4.778
Tosoh G7 Auto HPLC | 202 | 6.12 | 0.22 | 1.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 3.748
Tosoh G7 Auto HPLC | 202 | 9.8 | 0 | 1.4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 2.744
Tosoh G7 Auto HPLC | 202 | 7.56 | 0.16 | 1.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 3.1
Tosoh G7 Auto HPLC | 216 | 5.38 | 0.18 | 2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 4.1
Tosoh G7 Auto HPLC | 216 | 8.81 | 0.11 | 1.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 3.442
Tosoh G7 Auto HPLC | 216 | 6.48 | 0.18 | 1.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 3.316
Tosoh G7 Auto HPLC | 170 | 8.73 | 0.23 | 1.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Tosoh | 3.562
Tosoh G7 Auto HPLC | 170 | 5.56 | 0.16 | 1.6 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Tosoh | 3.296
Tosoh G7 Auto HPLC | 170 | 6.57 | 0.17 | 2.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Tosoh | 4.286
Tosoh G7 Auto HPLC | 181 | 6.62 | 0.32 | 2.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Tosoh | 4.632
Tosoh G7 Auto HPLC | 181 | 8.01 | 0.41 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Tosoh | 4.33
Tosoh G7 Auto HPLC | 181 | 9.56 | 0.36 | 2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Tosoh | 4.28
Tosoh G7 Auto HPLC | 184 | 5.73 | 0.13 | 2.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Tosoh | 4.246
Tosoh G7 Auto HPLC | 184 | 9.76 | 0.36 | 1.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Tosoh | 3.692
Tosoh G7 Auto HPLC | 184 | 7.58 | 0.38 | 1.6 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Tosoh | 3.516
Tosoh G7 Auto HPLC | 145 | 5.65 | 0.25 | 2.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Tosoh | 4.366
Tosoh G7 Auto HPLC | 145 | 8.61 | 0.31 | 1.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Tosoh | 3.838
Tosoh G7 Auto HPLC | 145 | 5.65 | 0 | 2.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Tosoh | 4.116
Tosoh G7 Auto HPLC | 116 | 7.38 | 0.27 | 1.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Tosoh | 3.602
Tosoh G7 Auto HPLC | 116 | 9.7 | 0.38 | 1.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Tosoh | 3.124
Tosoh G7 Auto HPLC | 116 | 6.34 | 0.27 | 1.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Tosoh | 3.994
Tosoh G7 Auto HPLC | 112 | 8.41 | 0.36 | 2.1 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Tosoh | 4.476
Tosoh G7 Auto HPLC | 112 | 5.53 | 0.23 | 2.4 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Tosoh | 4.934
Tosoh G7 Auto HPLC | 112 | 6.41 | 0.27 | 2.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Tosoh | 4.582
Tosoh G7 Auto HPLC | 88 | 6.77 | 0.19 | 1.9 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Tosoh | 3.914
Tosoh G7 Auto HPLC | 88 | 8.7 | 0.31 | 1.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Tosoh | 3.838
Tosoh G7 Auto HPLC | 88 | 5.81 | 0.16 | 2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Tosoh | 4.08
Tosoh G7 Auto HPLC | 102 | 6.81 | 0.32 | 1.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Tosoh | 4.044
Tosoh G7 Auto HPLC | 102 | 7.33 | 0.36 | 1.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Tosoh | 4.084
Tosoh G7 Auto HPLC | 102 | 10.06 | 0.41 | 1.9 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Tosoh | 4.134
Tosoh G7 Auto HPLC | 73 | 7.09 | 0.3 | 2.4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Tosoh | 5.004
Tosoh G7 Auto HPLC | 73 | 10.51 | 0.23 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Tosoh | 4.738
Tosoh G7 Auto HPLC | 73 | 7.09 | 0.27 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Tosoh | 5.366
Tosoh G7 Auto HPLC | 52 | 8.95 | 0.32 | 2.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Tosoh | 5.416
Tosoh G7 Auto HPLC | 52 | 5.48 | 0.16 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Tosoh | 5.452
Tosoh G7 Auto HPLC | 50 | 6.06 | 0.22 | 1.6 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Tosoh | 3.356
Tosoh G7 Auto HPLC | 50 | 11.95 | 0.24 | 1.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Tosoh | 3.572
Tosoh G7 Auto HPLC | 50 | 9.81 | 0.28 | 1.8 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Tosoh | 3.808
Tosoh G7 Auto HPLC | 50 | 5.2 | 0.16 | 1.6 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Tosoh | 3.296
Tosoh G7 Auto HPLC | 50 | 7.61 | 0.23 | 1.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Tosoh | 3.366
Tosoh G7 Auto HPLC | 45-60 | 11.98 | 0.29 | 2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Tosoh | 4.21
Tosoh G7 Auto HPLC | 45-60 | 6.15 | 0.22 | 2.1 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Tosoh | 4.336
Tosoh G7 Auto HPLC | 45-60 | 5.32 | 0.15 | 2.4 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Tosoh | 4.854
Tosoh G7 Auto HPLC | 45-60 | 8.41 | 0.27 | 1.9 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Tosoh | 3.994
Tosoh G7 Auto HPLC | 45-60 | 9.5 | 0.2 | 2.1 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Tosoh | 4.316
Tosoh G7 Auto HPLC | 36-46 | 5.43 | 0.11 | 1.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Tosoh | 3.834
Tosoh G7 Auto HPLC | 36-46 | 9.4 | 0.24 | 2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Tosoh | 4.16
Tosoh G7 Auto HPLC | 36-46 | 5.45 | 0.14 | 2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Tosoh | 4.06
Tosoh G7 Auto HPLC | 36-46 | 12.28 | 0.25 | 1.8 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Tosoh | 3.778
Tosoh G7 Auto HPLC | 36-46 | 6.06 | 0.12 | 2 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Tosoh | 4.04
Tosoh G7 Auto HPLC | 31 | 5.4 | 0.13 | 2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Tosoh | 4.05
Tosoh G7 Auto HPLC | 31 | 10.68 | 0.09 | 2.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Tosoh | 4.402
Tosoh G7 Auto HPLC | 31 | 6.34 | 0.14 | 1.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Tosoh | 3.472
Tosoh G7 Auto HPLC | 31 | 12.32 | 0.09 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Tosoh | 4.01
Tosoh G7 Auto HPLC | 31 | 7.67 | 0.16 | 2.1 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Tosoh | 4.276
Tosoh G7 Auto HPLC | 13-21 | 9.24 | 0.13 | 1.8 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Tosoh | 3.658
Tosoh G7 Auto HPLC | 13-21 | 6.16 | 0.15 | 2.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Tosoh | 5.246
Tosoh G7 Auto HPLC | 13-21 | 11.82 | 0.11 | 1.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Tosoh | 3.638
Tosoh G7 Auto HPLC | 13-21 | 5.16 | 0.14 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Tosoh | 5.628
Tosoh G7 Auto HPLC | 13-21 | 7.71 | 0.13 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Tosoh | 4.246
Tosoh G7 Auto HPLC | 0-10 | 6.57 | 0.16 | 1 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Tosoh | 2.12
Tosoh G7 Auto HPLC | 0-10 | 9.73 | 0.2 | 1.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Tosoh | 2.356
Tosoh G7 Auto HPLC | 0-10 | 5.42 | 0.08 | 1.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Tosoh | 3.02
Tosoh G7 Auto HPLC | 0-10 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 0-10 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 10 | 7.28 | 0.13 | 2.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Tosoh | 5.618
Tosoh G7 Auto HPLC | 10 | 5.28 | 0.09 | 1.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Tosoh | 3.422
Tosoh G7 Auto HPLC | 10 | 8.52 | 0.1 | 1.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Tosoh | 2.256
Tosoh G7 Auto HPLC | 10 |  |  |  | GH5-04 | 9.79 | 2018 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 10 |  |  |  | GH5-05 | 6.12 | 2018 | CAP | Fresh | Tosoh | 0
Tosoh G8 Auto HPLC | 85 | 5.9747 | -0.5 | 4.1 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Tosoh | 7.536
Tosoh G8 Auto HPLC | 83 | 7.3563 | -1 | 3.8 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Tosoh | 6.448
Tosoh G8 Auto HPLC | 234 | 6.0845 | 0.7 | 2.8 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Tosoh | 6.188
Tosoh G8 Auto HPLC | 234 | 7.5668 | 1.3 | 2.4 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Tosoh | 6.004
Tosoh G8 Auto HPLC | 79 | 6.2309 | 0.4 | 3.7 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Tosoh | 7.652
Tosoh G8 Auto HPLC | 79 | 7.4295 | -0.3 | 3.8 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Tosoh | 7.148
Tosoh G8 Auto HPLC | 281 | 6.3041 | 1.2 | 2.7 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Tosoh | 6.492
Tosoh G8 Auto HPLC | 281 | 7.5759 | 1.3 | 2.6 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Tosoh | 6.396
Tosoh G8 Auto HPLC | 98 | 7.4295 | -0.4 | 3.4 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Tosoh | 6.264
Tosoh G8 Auto HPLC | 97 | 6.1394 | 1.4 | 4.6 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Tosoh | 10.416
Tosoh G8 Auto HPLC | 351 | 7.6125 | 1.6 | 2.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Tosoh | 7.284
Tosoh G8 Auto HPLC | 349 | 6.2035 | 2.1 | 2.7 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Tosoh | 7.392
Tosoh G8 Auto HPLC | 354 | 8.802 | 2.6 | 2.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Tosoh | 7.304
Tosoh G8 Auto HPLC | 351 | 5.7734 | 1.4 | 2.9 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Tosoh | 7.084
Tosoh G8 Auto HPLC | 199 | 5.5995 | 1.2 | 3.5 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 8.06
Tosoh G8 Auto HPLC | 182 | 5.6087 | 1.3 | 3.7 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 8.552
Tosoh G8 Auto HPLC | 276 | 5.627 | 1.5 | 2.7 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 6.792
Tosoh G8 Auto HPLC | 274 | 5.6361 | 1.6 | 2.8 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 7.088
Tosoh G8 Auto HPLC | 41 | 5.2 |  | 1.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 3.528
Tosoh G8 Auto HPLC | 41 | 8.5 |  | 1.2 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 2.352
Tosoh G8 Auto HPLC | 41 | 6.2 |  | 1.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 2.548
Tosoh G8 Auto HPLC | 84 | 6.8 | 0.2 | 1.7 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 3.532
Tosoh G8 Auto HPLC | 84 | 9.52 | 0.02 | 1.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 2.96
Tosoh G8 Auto HPLC | 84 | 7.63 | 0.23 | 1.9 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 3.954
Tosoh G8 Auto HPLC | 116 | 6.1 | 0.2 | 1.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 3.924
Tosoh G8 Auto HPLC | 116 | 9.8 | 0 | 1.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 2.94
Tosoh G8 Auto HPLC | 116 | 7.56 | 0.16 | 1.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 3.1
Tosoh G8 Auto HPLC | 148 | 5.34 | 0.14 | 1.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 3.08
Tosoh G8 Auto HPLC | 148 | 8.77 | 0.07 | 1.6 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 3.206
Tosoh G8 Auto HPLC | 148 | 6.45 | 0.15 | 1.4 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 2.894
Tosoh G8 Auto HPLC | 186 | 8.69 | 0.19 | 1.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Tosoh | 3.522
Tosoh G8 Auto HPLC | 186 | 5.53 | 0.13 | 1.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Tosoh | 3.07
Tosoh G8 Auto HPLC | 186 | 6.56 | 0.16 | 1.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Tosoh | 2.904
Tosoh G8 Auto HPLC | 212 | 6.58 | 0.28 | 1.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Tosoh | 3.22
Tosoh G8 Auto HPLC | 212 | 7.99 | 0.39 | 1.4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Tosoh | 3.134
Tosoh G8 Auto HPLC | 212 | 9.54 | 0.34 | 1.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Tosoh | 3.084
Tosoh G8 Auto HPLC | 260 | 5.71 | 0.11 | 1.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Tosoh | 3.05
Tosoh G8 Auto HPLC | 260 | 9.75 | 0.35 | 1.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Tosoh | 2.702
Tosoh G8 Auto HPLC | 260 | 7.58 | 0.38 | 1.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Tosoh | 2.928
Tosoh G8 Auto HPLC | 284 | 5.62 | 0.22 | 1.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Tosoh | 3.16
Tosoh G8 Auto HPLC | 284 | 8.58 | 0.28 | 1.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Tosoh | 2.828
Tosoh G8 Auto HPLC | 284 | 5.64 | -0.01 | 1.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Tosoh | 3.518
Tosoh G8 Auto HPLC | 259 | 7.34 | 0.23 | 1.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Tosoh | 2.778
Tosoh G8 Auto HPLC | 259 | 9.67 | 0.35 | 1.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Tosoh | 2.702
Tosoh G8 Auto HPLC | 259 | 6.31 | 0.24 | 1.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Tosoh | 2.984
Tosoh G8 Auto HPLC | 280 | 8.41 | 0.36 | 1.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Tosoh | 3.3
Tosoh G8 Auto HPLC | 280 | 5.51 | 0.21 | 1.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Tosoh | 3.542
Tosoh G8 Auto HPLC | 280 | 6.4 | 0.26 | 1.7 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Tosoh | 3.592
Tosoh G8 Auto HPLC | 330 | 6.73 | 0.15 | 1.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Tosoh | 3.678
Tosoh G8 Auto HPLC | 330 | 8.68 | 0.29 | 1.6 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Tosoh | 3.426
Tosoh G8 Auto HPLC | 330 | 5.78 | 0.13 | 1.8 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Tosoh | 3.658
Tosoh G8 Auto HPLC | 327 | 6.77 | 0.28 | 1.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Tosoh | 3.808
Tosoh G8 Auto HPLC | 327 | 7.32 | 0.35 | 1.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Tosoh | 4.074
Tosoh G8 Auto HPLC | 327 | 10.05 | 0.4 | 1.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Tosoh | 3.928
Tosoh G8 Auto HPLC | 330 | 7.05 | 0.26 | 2.1 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Tosoh | 4.376
Tosoh G8 Auto HPLC | 330 | 10.5 | 0.22 | 1.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Tosoh | 3.944
Tosoh G8 Auto HPLC | 330 | 7.06 | 0.24 | 1.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Tosoh | 3.964
Tosoh G8 Auto HPLC | 287 | 8.93 | 0.3 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Tosoh | 4.024
Tosoh G8 Auto HPLC | 287 | 5.45 | 0.13 | 1.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Tosoh | 3.854
Tosoh G8 Auto HPLC | 285 | 6.03 | 0.19 | 1.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Tosoh | 3.522
Tosoh G8 Auto HPLC | 285 | 11.98 | 0.27 | 1.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Tosoh | 3.602
Tosoh G8 Auto HPLC | 285 | 9.81 | 0.28 | 1.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Tosoh | 3.612
Tosoh G8 Auto HPLC | 285 | 5.17 | 0.13 | 2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Tosoh | 4.05
Tosoh G8 Auto HPLC | 285 | 7.6 | 0.22 | 1.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Tosoh | 3.16
Tosoh G8 Auto HPLC | 298-344 | 11.88 | 0.19 | 1.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Tosoh | 3.326
Tosoh G8 Auto HPLC | 298-344 | 6.09 | 0.16 | 1.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Tosoh | 3.492
Tosoh G8 Auto HPLC | 298-344 | 5.26 | 0.09 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Tosoh | 3.422
Tosoh G8 Auto HPLC | 298-344 | 8.33 | 0.19 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Tosoh | 3.326
Tosoh G8 Auto HPLC | 298-344 | 9.41 | 0.11 | 1.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Tosoh | 3.05
Tosoh G8 Auto HPLC | 308-362 | 5.39 | 0.07 | 1.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Tosoh | 3.794
Tosoh G8 Auto HPLC | 308-362 | 9.38 | 0.22 | 1.5 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Tosoh | 3.16
Tosoh G8 Auto HPLC | 308-362 | 5.41 | 0.1 | 1.8 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Tosoh | 3.628
Tosoh G8 Auto HPLC | 308-362 | 12.17 | 0.14 | 1.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Tosoh | 3.08
Tosoh G8 Auto HPLC | 308-362 | 6.02 | 0.08 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Tosoh | 3.216
Tosoh G8 Auto HPLC | 313 | 5.34 | 0.07 | 1.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Tosoh | 3.206
Tosoh G8 Auto HPLC | 313 | 10.7 | 0.11 | 1.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Tosoh | 2.854
Tosoh G8 Auto HPLC | 313 | 6.29 | 0.09 | 1.4 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Tosoh | 2.834
Tosoh G8 Auto HPLC | 313 | 12.28 | 0.05 | 1.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Tosoh | 2.794
Tosoh G8 Auto HPLC | 313 | 7.65 | 0.14 | 1.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Tosoh | 3.08
Tosoh G8 Auto HPLC | 321-373 | 9.17 | 0.06 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Tosoh | 3.196
Tosoh G8 Auto HPLC | 321-373 | 6.07 | 0.06 | 1.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Tosoh | 3.196
Tosoh G8 Auto HPLC | 321-373 | 11.73 | 0.02 | 1.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Tosoh | 2.96
Tosoh G8 Auto HPLC | 321-373 | 5.1 | 0.08 | 1.7 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Tosoh | 3.412
Tosoh G8 Auto HPLC | 321-373 | 7.66 | 0.08 | 1.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Tosoh | 2.824
Tosoh G8 Auto HPLC | 335-389 | 6.53 | 0.12 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Tosoh | 3.256
Tosoh G8 Auto HPLC | 335-389 | 9.73 | 0.2 | 1.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Tosoh | 2.944
Tosoh G8 Auto HPLC | 335-389 | 5.4 | 0.06 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Tosoh | 3.392
Tosoh G8 Auto HPLC | 335-389 | 8.7 | 0.19 | 1.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Tosoh | 2.934
Tosoh G8 Auto HPLC | 335-389 | 7.38 | 0.13 | 1.4 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Tosoh | 2.874
Tosoh G8 Auto HPLC | 319 | 7.57 | 0.15 | 1.4 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Tosoh | 2.894
Tosoh G8 Auto HPLC | 319 | 5.38 | 0.18 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Tosoh | 3.512
Tosoh G8 Auto HPLC | 319 | 8.42 | 0.11 | 1.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Tosoh | 2.854
Tosoh G8 Auto HPLC | 319 | 10.51 | 0.18 | 1.4 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Tosoh | 2.924
Tosoh G8 Auto HPLC | 319 | 6.02 | 0.15 | 1.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Tosoh | 3.09
Tosoh G8 Auto HPLC | 378 | 7.35 | 0.1 | 1.4 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Tosoh | 2.844
Tosoh G8 Auto HPLC | 378 | 6.99 | 0.18 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Tosoh | 2.924
Tosoh G8 Auto HPLC | 378 | 8.72 | 0.07 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Tosoh | 2.618
Tosoh G8 Auto HPLC | 378 | 9.68 | 0.18 | 1.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Tosoh | 2.532
Tosoh G8 Auto HPLC | 378 | 5.54 | 0.09 | 1.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Tosoh | 2.834
Tosoh G8 Auto HPLC | 373 | 7.21 | 0.06 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Tosoh | 2.804
Tosoh G8 Auto HPLC | 373 | 5.29 | 0.1 | 1.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Tosoh | 2.844
Tosoh G8 Auto HPLC | 373 | 8.52 | 0.1 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Tosoh | 2.648
Tosoh G8 Auto HPLC | 373 | 9.92 | 0.13 | 1.1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Tosoh | 2.286
Tosoh G8 Auto HPLC | 373 | 6.23 | 0.11 | 1.3 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Tosoh | 2.658
Tosoh G8 Auto HPLC | 310 | 6.48 | 0.17 | 1.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Tosoh | 2.914
Tosoh G8 Auto HPLC | 310 | 9.27 | 0.16 | 1.2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Tosoh | 2.512
Tosoh G8 Auto HPLC | 310 | 5.37 | 0.06 | 1.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Tosoh | 2.608
Tosoh G8 Auto HPLC | 310 | 8.5 | 0.15 | 1.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Tosoh | 2.306
Tosoh G8 Auto HPLC | 310 | 7.76 | 0.14 | 1.2 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Tosoh | 2.492
Tosoh G8 Auto HPLC | 354 | 9.53 | 0.16 | 1.4 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Tosoh | 2.904
Tosoh G8 Auto HPLC | 354 | 5.02 | 0.1 | 1.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Tosoh | 3.236
Tosoh G8 Auto HPLC | 354 | 6.25 | 0.21 | 1.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Tosoh | 2.954
Tosoh G8 Auto HPLC | 354 | 5.47 | 0.19 | 1.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Tosoh | 2.934
Tosoh G8 Auto HPLC | 354 | 8.28 | 0.2 | 1.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Tosoh | 2.552
Tosoh G8 Auto HPLC | 347 | 5.6 | 0.14 | 1.4 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Tosoh | 2.884
Tosoh G8 Auto HPLC | 347 | 5.54 | -0.12 | 1.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Tosoh | 3.212
Tosoh G8 Auto HPLC | 347 | 9.55 | 0.24 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Tosoh | 2.788
Tosoh G8 Auto HPLC | 347 | 5.36 | 0.08 | 1.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Tosoh | 2.824
Tosoh G8 Auto HPLC | 347 | 7.6 | 0.19 | 1.3 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Tosoh | 2.738
Tosoh G8 Auto HPLC | 294 | 6.61 | 0.2 | 1.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Tosoh | 2.944
Tosoh G8 Auto HPLC | 294 | 8.93 | 0.33 | 1.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Tosoh | 2.878
Tosoh G8 Auto HPLC | 294 | 5.55 | 0.13 | 1.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Tosoh | 3.266
Tosoh G8 Auto HPLC | 294 | 7.62 | 0.24 | 1.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Tosoh | 2.984
Tosoh G8 Auto HPLC | 294 | 10.07 | 0.32 | 1.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Tosoh | 3.064
Tosoh G8 Auto HPLC | 337 | 5.3 | 0.05 | 1.9 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Tosoh | 3.774
Tosoh G8 Auto HPLC | 337 | 6.65 | 0.24 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Tosoh | 3.376
Tosoh G8 Auto HPLC | 337 | 8.46 | 0.25 | 1.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Tosoh | 2.994
Tosoh G8 Auto HPLC | 337 | 5.56 | 0.09 | 1.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Tosoh | 3.422
Tosoh G8 Auto HPLC | 337 | 9.79 | 0.31 | 1.4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Tosoh | 3.054
Tosoh G8 Auto HPLC | 308 | 8.2 | 0.23 | 1.5 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Tosoh | 3.17
Tosoh G8 Auto HPLC | 308 | 5.99 | 0.1 | 1.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Tosoh | 3.432
Tosoh G8 Auto HPLC | 308 | 5.18 | 0.05 | 2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Tosoh | 3.97
Tosoh G8 Auto HPLC | 308 | 7.58 | 0.18 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Tosoh | 3.904
Tosoh G8 Auto HPLC | 308 | 9.43 | 0.26 | 1.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Tosoh | 3.004
Tosoh G8 Auto HPLC | 271 | 7.99 | 0.18 | 1.6 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Tosoh | 3.316
Tosoh G8 Auto HPLC | 271 | 6.03 | 0.12 | 1.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Tosoh | 3.844
Tosoh G8 Auto HPLC | 271 | 8.96 | 0.27 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Tosoh | 3.602
Tosoh G8 Auto HPLC | 271 | 6.04 | 0.18 | 1.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Tosoh | 3.708
Tosoh G8 Auto HPLC | 271 | 9.84 | 0.28 | 1.5 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Tosoh | 3.22
Tosoh G8 Auto HPLC | 316 | 5.59 | 0.11 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Tosoh | 3.442
Tosoh G8 Auto HPLC | 316 | 8.46 | 0.2 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Tosoh | 2.944
Tosoh G8 Auto HPLC | 316 | 5.18 | 0.03 | 1.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Tosoh | 3.754
Tosoh G8 Auto HPLC | 316 | 10.39 | 0.34 | 1.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Tosoh | 2.888
Tosoh G8 Auto HPLC | 316 | 6.11 | 0.12 | 1.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Tosoh | 3.256
Tosoh G8 Auto HPLC | 311 | 9.18 | 0.32 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Tosoh | 3.456
Tosoh G8 Auto HPLC | 311 | 7.03 | 0.19 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Tosoh | 3.914
Tosoh G8 Auto HPLC | 311 | 9.65 | 0.26 | 1.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Tosoh | 3.2
Tosoh G8 Auto HPLC | 311 | 6.19 | 0.06 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Tosoh | 3.588
Tosoh G8 Auto HPLC | 311 | 7.69 | 0.17 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Tosoh | 3.306
Tosoh G8 Auto HPLC | 266 | 5.95 | 0.15 | 1.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Tosoh | 3.286
Tosoh G8 Auto HPLC | 266 | 9.38 | 0.26 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Tosoh | 3.396
Tosoh G8 Auto HPLC | 266 | 5.58 | 0.11 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Tosoh | 3.246
Tosoh G8 Auto HPLC | 266 | 7.37 | 0.12 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Tosoh | 3.06
Tosoh G8 Auto HPLC | 266 | 9.58 | 0.32 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Tosoh | 3.26
Tosoh G8 Auto HPLC | 304 | 5.25 | 0.04 | 1.6 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Tosoh | 3.176
Tosoh G8 Auto HPLC | 304 | 6.84 | 0.16 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Tosoh | 2.904
Tosoh G8 Auto HPLC | 304 | 8.52 |  | 2.3 | GH5-13 |  | 2021 | CAP | Fresh | Tosoh | 4.508
Tosoh G8 Auto HPLC | 304 | 6.09 | 0.11 | 1.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Tosoh | 2.854
Tosoh G8 Auto HPLC | 304 | 10.26 | 0.14 | 1.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Tosoh | 3.276
Tosoh GX | 22 | 8.8112 | 2.7 | 3.9 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Tosoh | 10.344
Tosoh GX | 22 | 5.7917 | 1.6 | 2.6 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Tosoh | 6.696
Tosoh GX | 20 | 5.5904 | 1.1 | 4.8 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 10.508
Tosoh GX | 20 | 5.5995 | 1.2 | 5.2 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Tosoh | 11.392
Tosoh GX | 22 | 5.5995 | 1.2 | 3.4 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 7.864
Tosoh GX | 22 | 5.6178 | 1.4 | 3.7 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 8.652
Tosoh not specified/other | 99 | 5.5995 | 1.2 | 3.1 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 7.276
Tosoh not specified/other | 98 | 5.6087 | 1.3 | 2.7 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Tosoh | 6.592
Tosoh not specified/other | 55 | 6.3133 | 1.3 | 3.5 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Tosoh | 8.16
Tosoh not specified/other | 55 | 7.4478 | -0.1 | 2.9 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Tosoh | 5.584
Tosoh not specified/other | 58 | 7.5027 | 0.4 | 2.9 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Tosoh | 6.084
Tosoh not specified/other | 58 | 6.1028 | 1 | 3.4 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Tosoh | 7.664
Trinity Biotech HPLC | 20 | 7.15 | 0.04 | 2.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Trinity Biotech | 4.94
Trinity Biotech HPLC | 20 | 9.23 | -0.09 | 2.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Trinity Biotech | 5.398
Trinity Biotech HPLC | 20 | 6.11 | 0.04 | 2.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Trinity Biotech | 4.352
Trinity Biotech HPLC | 18 | 7.96 | -0.09 | 1.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Trinity Biotech | 3.046
Trinity Biotech HPLC | 18 | 5.32 | 0.02 | 1.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Trinity Biotech | 2.96
Trinity Biotech HPLC | 18 | 6.14 | 0 | 1.7 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Trinity Biotech | 3.332
Trinity Biotech HPLC | 16 | 6.49 | -0.09 | 1.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Trinity Biotech | 3.438
Trinity Biotech HPLC | 16 | 8.38 | -0.01 | 3.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Trinity Biotech | 7.242
Trinity Biotech HPLC | 16 | 5.58 | -0.07 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Trinity Biotech | 5.026
Trinity Biotech HPLC | 20 | 6.39 | -0.1 | 1.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Trinity Biotech | 3.624
Trinity Biotech HPLC | 20 | 6.77 | -0.2 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Trinity Biotech | 4.112
Trinity Biotech HPLC | 20 | 9.51 | -0.14 | 3.1 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Trinity Biotech | 5.936
Trinity Biotech HPLC | 14 | 6.66 | -0.13 | 1.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Trinity Biotech | 2.81
Trinity Biotech HPLC | 14 | 10.26 | -0.02 | 1.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Trinity Biotech | 3.704
Trinity Biotech HPLC | 14 | 6.68 | -0.14 | 1.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Trinity Biotech | 3.584
Trinity Biotech HPLC | 13 | 8.5 | -0.13 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Trinity Biotech | 3.594
Trinity Biotech HPLC | 13 | 5.31 | -0.01 | 2 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Trinity Biotech | 3.91
Trinity Biotech HPLC | 14 | 5.73 | -0.11 | 1.9 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Trinity Biotech | 3.614
Trinity Biotech HPLC | 14 | 11.39 | -0.32 | 1.9 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Trinity Biotech | 3.404
Trinity Biotech HPLC | 14 | 9.39 | -0.14 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Trinity Biotech | 5.152
Trinity Biotech HPLC | 14 | 5.05 | 0.01 | 1.9 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Trinity Biotech | 3.734
Trinity Biotech HPLC | 14 | 7.24 | -0.14 | 2.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Trinity Biotech | 3.976
Trinity Biotech HPLC | 12-13 | 11.72 | 0.03 | 2.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Trinity Biotech | 5.126
Trinity Biotech HPLC | 12-13 | 5.94 | 0.01 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Trinity Biotech | 5.302
Trinity Biotech HPLC | 12-13 | 5.22 | 0.05 | 2.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Trinity Biotech | 5.538
Trinity Biotech HPLC | 12-13 | 8.22 | 0.08 | 2.3 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Trinity Biotech | 4.588
Trinity Biotech HPLC | 12-13 | 9.44 | 0.14 | 2.8 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Trinity Biotech | 5.628
Trinity Biotech HPLC | 9-10 | 5.42 | 0.1 | 2.7 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Trinity Biotech | 5.392
Trinity Biotech HPLC | 9-10 | 9.26 | 0.1 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Trinity Biotech | 5.588
Trinity Biotech HPLC | 9-10 | 5.37 | 0.06 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Trinity Biotech | 4.372
Trinity Biotech HPLC | 9-10 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Trinity Biotech | 0
Trinity Biotech HPLC | 9-10 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Trinity Biotech | 0
Trinity Biotech HPLC (Affinity) | 27 | 6.66 | 0.06 | 3.4 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Trinity Biotech | 6.724
Trinity Biotech HPLC (Affinity) | 27 | 9.46 | -0.04 | 3.9 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Trinity Biotech | 7.604
Trinity Biotech HPLC (Affinity) | 27 | 7.36 | -0.04 | 3.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Trinity Biotech | 6.624
Trinity Biotech HPLC (Affinity) | 18 | 5.85 | -0.05 | 3.3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Trinity Biotech | 6.418
Trinity Biotech HPLC (Affinity) | 18 | 9.69 | -0.11 | 1.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Trinity Biotech | 3.026
Trinity Biotech HPLC (Affinity) | 18 | 7.36 | -0.04 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Trinity Biotech | 6.82
Trinity Biotech HPLC (Affinity) | 20 | 5.33 | 0.13 | 2.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Trinity Biotech | 5.03
Trinity Biotech HPLC (Affinity) | 20 | 8.63 | -0.07 | 3.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Trinity Biotech | 6.006
Trinity Biotech HPLC (Affinity) | 20 | 6.38 | 0.08 | 2.4 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Trinity Biotech | 4.784
Trinity Biotech HPLC (Affinity) | 19 | 8.58 | 0.08 | 2.6 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Trinity Biotech | 5.176
Trinity Biotech HPLC (Affinity) | 19 | 5.57 | 0.17 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Trinity Biotech | 6.05
Trinity Biotech HPLC (Affinity) | 19 | 6.46 | 0.06 | 1.8 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Trinity Biotech | 3.588
Trinity Biotech HPLC (Affinity) | 20 | 6.41 | 0.11 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Trinity Biotech | 7.558
Trinity Biotech HPLC (Affinity) | 20 | 7.68 | 0.08 | 2.3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Trinity Biotech | 4.588
Trinity Biotech HPLC (Affinity) | 20 | 9.29 | 0.09 | 2.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Trinity Biotech | 4.206
Trinity Biotech HPLC (Affinity) | 31 | 5.76 | 0.16 | 3.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Trinity Biotech | 6.236
Trinity Biotech HPLC (Affinity) | 31 | 9.25 | -0.15 | 2.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Trinity Biotech | 4.75
Trinity Biotech HPLC (Affinity) | 31 | 7.31 | 0.11 | 2.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Trinity Biotech | 4.618
Trinity Biotech HPLC (Affinity) | 30 | 5.53 | 0.13 | 1.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Trinity Biotech | 3.658
Trinity Biotech HPLC (Affinity) | 30 | 8.28 | -0.02 | 2.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Trinity Biotech | 4.488
Trinity Biotech HPLC (Affinity) | 30 | 5.69 | 0.04 | 1.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Trinity Biotech | 3.568
Trinity Biotech Premier | 50 | 6.56 | -0.02 | 2.7 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Trinity Biotech | 5.272
Trinity Biotech Premier | 50 | 8.4 | 0.01 | 2.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Trinity Biotech | 5.302
Trinity Biotech Premier | 50 | 5.69 | 0.04 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Trinity Biotech | 5.136
Trinity Biotech Premier | 41 | 6.42 | -0.07 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Trinity Biotech | 4.83
Trinity Biotech Premier | 41 | 6.83 | -0.14 | 2.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Trinity Biotech | 4.564
Trinity Biotech Premier | 41 | 9.51 | -0.14 | 2.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Trinity Biotech | 5.152
Trinity Biotech Premier | 60 | 6.74 | -0.05 | 2.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Trinity Biotech | 4.458
Trinity Biotech Premier | 60 | 10.22 | -0.06 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Trinity Biotech | 4.448
Trinity Biotech Premier | 60 | 6.83 | 0.01 | 2.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Trinity Biotech | 4.518
Trinity Biotech Premier | 54 | 8.54 | -0.09 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Trinity Biotech | 3.634
Trinity Biotech Premier | 54 | 5.35 | 0.03 | 2.2 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Trinity Biotech | 4.342
Trinity Biotech Premier | 53 | 5.84 | 0 | 2.9 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Trinity Biotech | 5.684
Trinity Biotech Premier | 53 | 11.73 | 0.02 | 3.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Trinity Biotech | 6.488
Trinity Biotech Premier | 53 | 9.61 | 0.08 | 3 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Trinity Biotech | 5.96
Trinity Biotech Premier | 53 | 5.13 | 0.09 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Trinity Biotech | 5.382
Trinity Biotech Premier | 53 | 7.41 | 0.03 | 2.7 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Trinity Biotech | 5.322
Trinity Biotech Premier | 55-65 | 12 | 0.31 | 1.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Trinity Biotech | 3.25
Trinity Biotech Premier | 55-65 | 6.06 | 0.13 | 1.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Trinity Biotech | 3.854
Trinity Biotech Premier | 55-65 | 5.3 | 0.13 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Trinity Biotech | 3.462
Trinity Biotech Premier | 55-65 | 8.33 | 0.19 | 1.8 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Trinity Biotech | 3.718
Trinity Biotech Premier | 55-65 | 9.67 | 0.37 | 1.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Trinity Biotech | 3.702
Trinity Biotech Premier | 54-63 | 5.42 | 0.1 | 2.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Trinity Biotech | 4.608
Trinity Biotech Premier | 54-63 | 9.39 | 0.23 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Trinity Biotech | 4.542
Trinity Biotech Premier | 54-63 | 5.45 | 0.14 | 2.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Trinity Biotech | 4.648
Trinity Biotech Premier | 54-63 | 12.36 | 0.33 | 2.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Trinity Biotech | 4.446
Trinity Biotech Premier | 54-63 | 6.14 | 0.2 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Trinity Biotech | 4.708
Trinity Biotech Premier | 52 | 5.3 | 0.03 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Trinity Biotech | 5.126
Trinity Biotech Premier | 52 | 10.7 | 0.11 | 1.9 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Trinity Biotech | 3.834
Trinity Biotech Premier | 52 | 6.29 | 0.09 | 2.1 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Trinity Biotech | 4.206
Trinity Biotech Premier | 52 | 12.48 | 0.25 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Trinity Biotech | 4.17
Trinity Biotech Premier | 52 | 7.63 | 0.12 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Trinity Biotech | 4.432
Trinity Biotech Premier | 58-69 | 9.2 | 0.09 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Trinity Biotech | 4.402
Trinity Biotech Premier | 58-69 | 6.12 | 0.11 | 2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Trinity Biotech | 4.03
Trinity Biotech Premier | 58-69 | 11.99 | 0.28 | 2.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Trinity Biotech | 4.396
Trinity Biotech Premier | 58-69 | 5.07 | 0.05 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Trinity Biotech | 4.558
Trinity Biotech Premier | 58-69 | 7.68 | 0.1 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Trinity Biotech | 4.216
Trinity Biotech Premier | 67-75 | 6.36 | -0.05 | 1.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Trinity Biotech | 3.674
Trinity Biotech Premier | 67-75 | 9.53 | 0 | 2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Trinity Biotech | 3.92
Trinity Biotech Premier | 67-75 | 5.36 | 0.02 | 1.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Trinity Biotech | 3.744
Trinity Biotech Premier | 67-75 | 8.5 | -0.01 | 2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Trinity Biotech | 3.91
Trinity Biotech Premier | 67-75 | 7.39 | 0.14 | 2.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Trinity Biotech | 4.256
Trinity Biotech Premier | 85 | 7.35 | 0.2 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Trinity Biotech | 3.532
Trinity Biotech Premier | 85 | 5.41 | 0.22 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Trinity Biotech | 4.14
Trinity Biotech Premier | 85 | 8.55 | 0.13 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Trinity Biotech | 2.874
Trinity Biotech Premier | 85 | 9.96 | 0.17 | 1.5 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Trinity Biotech | 3.11
Trinity Biotech Premier | 85 | 6.33 | 0.21 | 1.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Trinity Biotech | 3.934
Trinity Biotech Premier | 70 | 6.43 | 0.12 | 1.6 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Trinity Biotech | 3.256
Trinity Biotech Premier | 70 | 9.29 | 0.18 | 1.3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Trinity Biotech | 2.728
Trinity Biotech Premier | 70 | 5.56 | 0.25 | 1.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Trinity Biotech | 3.778
Trinity Biotech Premier | 70 | 8.48 | 0.13 | 1.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Trinity Biotech | 3.462
Trinity Biotech Premier | 70 | 7.81 | 0.19 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Trinity Biotech | 3.13
Trinity Biotech Premier | 69 | 7.38 | -0.04 | 1.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Trinity Biotech | 3.684
Trinity Biotech Premier | 69 | 5.21 | 0.01 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Trinity Biotech | 3.93
Trinity Biotech Premier | 69 | 8.39 | 0.08 | 1.8 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Trinity Biotech | 3.608
Trinity Biotech Premier | 69 | 10.51 | 0.18 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Trinity Biotech | 3.904
Trinity Biotech Premier | 69 | 5.91 | 0.04 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Trinity Biotech | 3.764
Trinity Biotech Premier | 79 | 7.28 | 0.03 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Trinity Biotech | 4.538
Trinity Biotech Premier | 79 | 6.79 | -0.02 | 2.3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Trinity Biotech | 4.488
Trinity Biotech Premier | 79 | 8.92 | 0.27 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Trinity Biotech | 4.778
Trinity Biotech Premier | 79 | 9.58 | 0.08 | 2.1 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Trinity Biotech | 4.196
Trinity Biotech Premier | 79 | 5.55 | 0.1 | 2.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Trinity Biotech | 4.804
Trinity Biotech Premier Hb9210 | 11 | 7.17 | 0.06 | 2.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Trinity Biotech | 4.96
Trinity Biotech Premier Hb9210 | 11 | 9.26 | -0.06 | 2.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Trinity Biotech | 4.644
Trinity Biotech Premier Hb9210 | 11 | 6.17 | 0.1 | 2.5 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Trinity Biotech | 5
Trinity Biotech Premier Hb9210 | 29 | 7.96 | -0.09 | 2.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Trinity Biotech | 4.81
Trinity Biotech Premier Hb9210 | 29 | 5.36 | 0.06 | 2.9 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Trinity Biotech | 5.744
Trinity Biotech Premier Hb9210 | 29 | 6.18 | 0.04 | 2.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Trinity Biotech | 4.94
Trinity Biotech Premier Hb9210 | 17 | 7.3563 | -1.2 | 4.3 | Sample 1 | 7.4662 | 2018 | EURA1c | Frozen | Trinity Biotech | 7.228
Trinity Biotech Premier Hb9210 | 16 | 6.0113 | 0 | 3.9 | Sample 2 | 6.0113 | 2018 | EURA1c | Frozen | Trinity Biotech | 7.644
Trinity Biotech Premier Hb9210 | 38 | 7.6125 | 1.6 | 3 | Sample 1 | 7.4662 | 2018 | EURA1c | Fresh | Trinity Biotech | 7.48
Trinity Biotech Premier Hb9210 | 37 | 6.1028 | 1 | 4 | Sample 2 | 6.0113 | 2018 | EURA1c | Fresh | Trinity Biotech | 8.84
Trinity Biotech Premier Hb9210 | 29 | 8.6556 | 1 | 2.4 | Sample 1 | 8.5641 | 2019 | EURA1c | Fresh | Trinity Biotech | 5.704
Trinity Biotech Premier Hb9210 | 28 | 5.6727 | 0.3 | 5.4 | Sample 2 | 5.6453 | 2019 | EURA1c | Fresh | Trinity Biotech | 10.884
Trinity Biotech Premier Hb9210 | 27 | 5.5355 | 0.5 | 5.1 | Sample 1 | 5.5812 | 2020 | EURA1c | Frozen | Trinity Biotech | 10.496
Trinity Biotech Premier Hb9210 | 26 | 5.5355 | 0.5 | 4.4 | Sample 2 | 5.5812 | 2020 | EURA1c | Frozen | Trinity Biotech | 9.124
Trinity Biotech Premier Hb9210 | 22 | 5.5721 | 0.9 | 4.2 | Sample 1 | 5.5446 | 2020 | EURA1c | Fresh | Trinity Biotech | 9.132
Trinity Biotech Premier Hb9210 | 23 | 5.5812 | 1 | 6.6 | Sample 2 | 5.5446 | 2020 | EURA1c | Fresh | Trinity Biotech | 13.936
Trinity Biotech Premier Hb9210 | 16 | 5.9473 | -0.8 | 4.2 | Sample 1 | 6.0204 | 2016 | EURA1c | Frozen | Trinity Biotech | 7.432
Trinity Biotech Premier Hb9210 | 16 | 7.3655 | -0.9 | 3.2 | Sample 2 | 7.4478 | 2016 | EURA1c | Frozen | Trinity Biotech | 5.372
Trinity Biotech Premier Hb9210 | 27 | 6.0662 | 0.5 | 3.5 | Sample 1 | 6.0204 | 2016 | EURA1c | Fresh | Trinity Biotech | 7.36
Trinity Biotech Premier Hb9210 | 27 | 7.6125 | 1.8 | 4 | Sample 2 | 7.4478 | 2016 | EURA1c | Fresh | Trinity Biotech | 9.64
Trinity Biotech Premier Hb9210 | 20 | 6.2675 | 0.8 | 2.2 | Sample 1 | 6.1943 | 2017 | EURA1c | Frozen | Trinity Biotech | 5.112
Trinity Biotech Premier Hb9210 | 21 | 7.457 | 0 | 2.4 | Sample 2 | 7.457 | 2017 | EURA1c | Frozen | Trinity Biotech | 4.704
Trinity Biotech Premier Hb9210 | 29 | 6.295 | 1.1 | 3.7 | Sample 1 | 6.1943 | 2017 | EURA1c | Fresh | Trinity Biotech | 8.352
Trinity Biotech Premier Hb9210 | 29 | 7.5942 | 1.5 | 2.8 | Sample 2 | 7.457 | 2017 | EURA1c | Fresh | Trinity Biotech | 6.988
Trinity Biotech Premier Hb9210 HPLC | 87 | 9.45 | 0.08 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Trinity Biotech | 3.608
Trinity Biotech Premier Hb9210 HPLC | 87 | 5.12 | 0.2 | 1.8 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Trinity Biotech | 3.728
Trinity Biotech Premier Hb9210 HPLC | 87 | 6.11 | 0.07 | 1.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Trinity Biotech | 3.598
Trinity Biotech Premier Hb9210 HPLC | 87 | 5.38 | 0.1 | 1.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Trinity Biotech | 3.236
Trinity Biotech Premier Hb9210 HPLC | 87 | 8.2 | 0.12 | 1.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Trinity Biotech | 3.452
Trinity Biotech Premier Hb9210 HPLC | 83 | 8.79 | -0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Trinity Biotech | 3.262
Trinity Biotech Premier Hb9210 HPLC | 83 | 6.81 | -0.03 | 1.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Trinity Biotech | 3.498
Trinity Biotech Premier Hb9210 HPLC | 83 | 9.36 | -0.03 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Trinity Biotech | 3.498
Trinity Biotech Premier Hb9210 HPLC | 83 | 6.14 | 0.01 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Trinity Biotech | 3.538
Trinity Biotech Premier Hb9210 HPLC | 83 | 7.55 | 0.03 | 1.7 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Trinity Biotech | 3.362
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.43 | -0.03 | 2.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Trinity Biotech | 4.282
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.74 | 0.08 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Trinity Biotech | 5.568
Trinity Biotech Premier Hb9210 HPLC | 89 | 9.43 | 0.12 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Trinity Biotech | 4.04
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.35 | 0.07 | 2.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Trinity Biotech | 4.382
Trinity Biotech Premier Hb9210 HPLC | 89 | 7.45 | 0.04 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Trinity Biotech | 4.352
Trinity Biotech Premier Hb9210 HPLC | 80 | 6.39 | -0.02 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Trinity Biotech | 3.704
Trinity Biotech Premier Hb9210 HPLC | 80 | 8.73 | 0.13 | 1.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Trinity Biotech | 3.462
Trinity Biotech Premier Hb9210 HPLC | 80 | 5.51 | 0.09 | 2.1 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Trinity Biotech | 4.206
Trinity Biotech Premier Hb9210 HPLC | 80 | 7.5 | 0.12 | 1.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Trinity Biotech | 3.648
Trinity Biotech Premier Hb9210 HPLC | 80 | 9.93 | 0.18 | 1.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Trinity Biotech | 3.904
Trinity Biotech Premier Hb9210 HPLC | 86 | 5.37 | 0.12 | 2.1 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Trinity Biotech | 4.236
Trinity Biotech Premier Hb9210 HPLC | 86 | 6.46 | 0.05 | 1.8 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Trinity Biotech | 3.578
Trinity Biotech Premier Hb9210 HPLC | 86 | 8.41 | 0.2 | 1.6 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Trinity Biotech | 3.336
Trinity Biotech Premier Hb9210 HPLC | 86 | 5.61 | 0.14 | 2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Trinity Biotech | 4.06
Trinity Biotech Premier Hb9210 HPLC | 86 | 9.73 | 0.25 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Trinity Biotech | 3.582
Trinity Biotech Premier Hb9210 HPLC | 80 | 8.11 | 0.14 | 2.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Trinity Biotech | 4.256
Trinity Biotech Premier Hb9210 HPLC | 80 | 6.06 | 0.17 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Trinity Biotech | 4.678
Trinity Biotech Premier Hb9210 HPLC | 80 | 5.28 | 0.15 | 2.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Trinity Biotech | 4.462
Trinity Biotech Premier Hb9210 HPLC | 80 | 7.53 | 0.13 | 2.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Trinity Biotech | 5.03
Trinity Biotech Premier Hb9210 HPLC | 80 | 9.31 | 0.14 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Trinity Biotech | 4.648
Trinity Biotech Premier Hb9210 HPLC | 71 | 7.94 | 0.13 | 2.2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Trinity Biotech | 4.442
Trinity Biotech Premier Hb9210 HPLC | 71 | 6.03 | 0.12 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Trinity Biotech | 4.432
Trinity Biotech Premier Hb9210 HPLC | 71 | 8.84 | 0.15 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Trinity Biotech | 4.462
Trinity Biotech Premier Hb9210 HPLC | 71 | 5.97 | 0.11 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Trinity Biotech | 5.01
Trinity Biotech Premier Hb9210 HPLC | 71 | 9.81 | 0.25 | 2.1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Trinity Biotech | 4.366
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.47 | -0.01 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Trinity Biotech | 3.518
Trinity Biotech Premier Hb9210 HPLC | 79 | 8.2 | -0.06 | 1.8 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Trinity Biotech | 3.468
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.2 | 0.05 | 2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Trinity Biotech | 3.97
Trinity Biotech Premier Hb9210 HPLC | 79 | 9.94 | -0.11 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Trinity Biotech | 3.222
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.95 | -0.04 | 2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Trinity Biotech | 3.88
Trinity Biotech Premier Hb9210 HPLC | 73 | 5.78 | -0.02 | 2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Trinity Biotech | 3.9
Trinity Biotech Premier Hb9210 HPLC | 73 | 9.26 | 0.14 | 2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Trinity Biotech | 4.06
Trinity Biotech Premier Hb9210 HPLC | 73 | 5.54 | 0.07 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Trinity Biotech | 3.794
Trinity Biotech Premier Hb9210 HPLC | 73 | 7.38 | 0.13 | 2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Trinity Biotech | 4.05
Trinity Biotech Premier Hb9210 HPLC | 73 | 9.37 | 0.11 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Trinity Biotech | 4.03
Trinity Biotech Premier Hb9210 HPLC | 81 | 5.31 | 0.1 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Trinity Biotech | 4.412
Trinity Biotech Premier Hb9210 HPLC | 81 | 6.74 | 0.06 | 2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Trinity Biotech | 3.98
Trinity Biotech Premier Hb9210 HPLC | 81 | 8.5 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Trinity Biotech | 4.704
Trinity Biotech Premier Hb9210 HPLC | 81 | 6.04 | 0.06 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Trinity Biotech | 4.176
Trinity Biotech Premier Hb9210 HPLC | 81 | 10.3 | 0.18 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Trinity Biotech | 3.904
Vitros 5.1 FS Chem System | 26 | 11 |  | 4.5 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Vitros | 8.82
Vitros 5.1 FS Chem System | 26 | 5.7 |  | 3.7 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Vitros | 7.252
Vitros 5.1 FS Chem System | 26 | 8.2 |  | 3.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Vitros | 7.252
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 2.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Vitros | 5.292
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Vitros | 7.644
Vitros 5.1 FS Chem System | 39 | 11.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Vitros | 8.82
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 2.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Vitros | 5.292
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Vitros | 7.644
Vitros 5.1 FS Chem System | 39 | 11.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Vitros | 8.82
Vitros 5.1 FS Chem System | 66 | 5.7 |  | 3.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Vitros | 7.056
Vitros 5.1 FS Chem System | 66 | 12.3 |  | 4.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Vitros | 8.624
Vitros 5.1 FS Chem System | 66 | 7.7 |  | 4 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Vitros | 7.84
Vitros 5.1 FS Chem System | 83 | 8.9 |  | 4.9 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Vitros | 9.604
Vitros 5.1 FS Chem System | 83 | 11.2 |  | 5.3 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Vitros | 10.388
Vitros 5.1 FS Chem System | 83 | 6.1 |  | 3.8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Vitros | 7.448
Vitros 5.1 FS Chem System | 98 | 12.9 |  | 5.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Vitros | 11.564
Vitros 5.1 FS Chem System | 98 | 5.3 |  | 2.9 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Vitros | 5.684
Vitros 5.1 FS Chem System | 98 | 8.6 |  | 3.7 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Vitros | 7.252
Vitros 5.1 FS Chem System | 117 | 11.5 |  | 4.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Vitros | 8.82
Vitros 5.1 FS Chem System | 117 | 6.6 |  | 2.5 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Vitros | 4.9
Vitros 5.1 FS Chem System | 117 | 5.2 |  | 3.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Vitros | 7.056
Vitros 5.1 FS Chem System | 111 | 5.1 |  | 3.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Vitros | 6.664
Vitros 5.1 FS Chem System | 111 | 8.3 |  | 3.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Vitros | 7.448
Vitros 5.1 FS Chem System | 111 | 5.9 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Vitros | 7.644
Vitros 5.1 FS Chem System | 124 | 6.8 | 0.2 | 4.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Vitros | 8.236
Vitros 5.1 FS Chem System | 124 | 9.58 | 0.08 | 3.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Vitros | 7.332
Vitros 5.1 FS Chem System | 124 | 7.55 | 0.15 | 4.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Vitros | 8.382
Vitros 5.1 FS Chem System | 212 | 5.28 | -0.12 | 2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Vitros | 3.8
Vitros 5.1 FS Chem System | 212 | 8.1 | -0.2 | 2.4 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Vitros | 4.504
Vitros 5.1 FS Chem System | 212 | 5.41 | -0.24 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Vitros | 4.856
Vitros 5.1 FS Chem System | 132 | 6.27 | 0.37 | 3.4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Vitros | 7.034
Vitros 5.1 FS Chem System | 132 | 10.2 | 0.4 | 3.7 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Vitros | 7.652
Vitros 5.1 FS Chem System | 132 | 7.85 | 0.45 | 3.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Vitros | 7.702
Vitros 5.1 FS Chem System | 154 | 5.52 | 0.32 | 4.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Vitros | 9.336
Vitros 5.1 FS Chem System | 154 | 8.92 | 0.22 | 3.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Vitros | 6.884
Vitros 5.1 FS Chem System | 154 | 6.57 | 0.27 | 3.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Vitros | 7.522
Vitros 5.1 FS Chem System | 155 | 8.71 | 0.21 | 3.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Vitros | 7.07
Vitros 5.1 FS Chem System | 155 | 5.48 | 0.08 | 3.6 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Vitros | 7.136
Vitros 5.1 FS Chem System | 155 | 6.56 | 0.16 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Vitros | 6.628
Vitros 5.1 FS. 4600. 5600 | 188 | 9.38 | 0.01 | 2.9 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Vitros | 5.694
Vitros 5.1 FS. 4600. 5600 | 188 | 4.88 | -0.04 | 1.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Vitros | 3.292
Vitros 5.1 FS. 4600. 5600 | 188 | 5.9 | -0.14 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Vitros | 4.564
Vitros 5.1 FS. 4600. 5600 | 188 | 5.19 | -0.09 | 2.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Vitros | 4.026
Vitros 5.1 FS. 4600. 5600 | 188 | 8.04 | -0.04 | 2.8 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Vitros | 5.448
Vitros 5.1 FS. 4600. 5600 | 193 | 5.41 | -0.05 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Vitros | 4.85
Vitros 5.1 FS. 4600. 5600 | 193 | 5.59 | -0.07 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Vitros | 5.222
Vitros 5.1 FS. 4600. 5600 | 193 | 9.36 | 0.05 | 2.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Vitros | 5.342
Vitros 5.1 FS. 4600. 5600 | 193 | 5.17 | -0.11 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Vitros | 4.594
Vitros 5.1 FS. 4600. 5600 | 193 | 7.31 | -0.1 | 2.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Vitros | 5.584
Vitros 5.1 FS. 4600. 5600 | 163 | 6.32 | -0.09 | 2.7 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Vitros | 5.202
Vitros 5.1 FS. 4600. 5600 | 163 | 8.71 | 0.11 | 2.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Vitros | 5.402
Vitros 5.1 FS. 4600. 5600 | 163 | 5.31 | -0.11 | 2.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Vitros | 4.986
Vitros 5.1 FS. 4600. 5600 | 163 | 7.34 | -0.04 | 2.7 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Vitros | 5.252
Vitros 5.1 FS. 4600. 5600 | 163 | 9.94 | 0.19 | 2.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Vitros | 5.874
Vitros 5.1 FS. 4600. 5600 | 194 | 5.3 | 0.05 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Vitros | 5.93
Vitros 5.1 FS. 4600. 5600 | 194 | 6.54 | 0.13 | 3.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Vitros | 6.794
Vitros 5.1 FS. 4600. 5600 | 194 | 8.25 | 0.04 | 2.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Vitros | 5.528
Vitros 5.1 FS. 4600. 5600 | 194 | 5.52 | 0.05 | 3.3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Vitros | 6.518
Vitros 5.1 FS. 4600. 5600 | 194 | 9.46 | -0.02 | 3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Vitros | 5.86
Vitros 5.1 FS. 4600. 5600 | 184 | 8.1 | 0.13 | 2.9 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Vitros | 5.814
Vitros 5.1 FS. 4600. 5600 | 184 | 5.94 | 0.05 | 2.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Vitros | 4.95
Vitros 5.1 FS. 4600. 5600 | 184 | 5.07 | -0.06 | 2.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Vitros | 4.056
Vitros 5.1 FS. 4600. 5600 | 184 | 7.48 | 0.08 | 3.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Vitros | 6.156
Vitros 5.1 FS. 4600. 5600 | 184 | 9.02 | -0.15 | 3.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Vitros | 6.318
Vitros 5.1 FS. 4600. 5600 | 180 | 7.74 | -0.07 | 2.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Vitros | 5.418
Vitros 5.1 FS. 4600. 5600 | 180 | 5.88 | -0.03 | 2.4 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Vitros | 4.674
Vitros 5.1 FS. 4600. 5600 | 180 | 8.45 | -0.24 | 3.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Vitros | 5.836
Vitros 5.1 FS. 4600. 5600 | 180 | 5.82 | -0.04 | 2.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Vitros | 4.664
Vitros 5.1 FS. 4600. 5600 | 180 | 9.16 | -0.4 | 2.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Vitros | 5.284
Vitros 5.1 7600 Chem. Systems | 200 | 5.28 | -0.2 | 2.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Vitros | 4.308
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 200 | 8.2 | -0.06 | 3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Vitros | 5.82
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 200 | 4.94 | -0.21 | 2.2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Vitros | 4.102
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 200 | 9.77 | -0.28 | 2.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Vitros | 5.404
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 200 | 5.81 | -0.18 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Vitros | 4.72
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 183 | 5.75 | -0.05 | 2.5 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Vitros | 4.85
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 183 | 9.04 | -0.08 | 2.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Vitros | 5.408
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 183 | 5.41 | -0.06 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Vitros | 4.448
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 183 | 7.33 | 0.08 | 2.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Vitros | 5.176
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 183 | 9.26 | 0 | 2.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Vitros | 5.684
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 199 | 5.12 | -0.09 | 1.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Vitros | 3.634
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 199 | 6.67 | -0.01 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Vitros | 4.694
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 199 | 8.22 |  | 2.6 | GH5-13 |  | 2021 | CAP | Fresh | Vitros | 5.096
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 199 | 5.94 | -0.04 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Vitros | 4.076
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 199 | 10.15 | 0.03 | 3.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Vitros | 6.302
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 9.08 | 0.22 | 3.1 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Vitros | 6.296
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 6.86 | 0.02 | 2.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Vitros | 5.116
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 9.4 | 0.01 | 3.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Vitros | 6.674
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 6.14 | 0.01 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Vitros | 4.91
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 7.59 | 0.07 | 2.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Vitros | 5.558
0
0
0
0